7052 | 1220648-2025-30783 | 22633457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22633457&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634324 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the implant, the patient went into tachycardia arrhythmia requiring defibrillation. additionally, the patient had hypotensive episodes during support. furthermore, the patient had oozing at the access site. the patient's bandages were reinforced. support continued. it was also reported that the patient experienced thrombocytopenia, and blood products were given. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: access site bleed: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ thrombocytopenia: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation was completed. access site bleeding - major: patient requiring several units of blood for access site oozing, femstop used after 30 minutes of manual pressure during closure, no oozing appeared. unknown how the oozing occurred, so the cause of bleeding issue cannot be determined due to insufficient clinical details. after impella successfully placed and therapy initiated patient went in to tachycardia arrhythmia requiring defibrillation, patient became unstable requiring epi bolus and increase of dopamine. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. decision made to remove impella due to low platelet, the cause of the thrombocytopenia was not determined due to insufficient clinical details. lactic remains elevated but down. from attachment, unfortunately the pump was removed and disposed of due to the low platelet count and patients need for several units of blood. cardiologist was concerned for hemolysis (although urine remained clear yellow) and continual oozing. complaint device not returned for analyzing. data log reviewed, suctions occurred followed by mc spikes observed on 7/15 & 16 at p7, quick resolved impella positioning in aorta and unknown position alarms observed, no placement signal issues seen. the cause of hemolysis issue most likely biomaterial ingestion since mc spikes observed in logs review.
7053 | 1220648-2025-30784 | 22634648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22634648&pc=PYX | 07/15/2025 | 09/12/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675721 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a bipella support ongoing. the impella rp flex and impella 5.5 pump were placed for the late presenting myocardial infarction. the pump support was ongoing with the rp flex when pressure signal was irregular and there were signs of hemolysis. the heparin-induced thrombocytopenia panel also came back positive during the support so the anticoagulation was switched. the rp flex advanced at bedside by 1cm, and lactate dehydrogenase down trending now to 800, but medical doctor electing to remove device and replace with protek duo. | [5 empty] | The investigation of hemolysis, thrombocytopenia, and optical signal issue has been completed. the device was returned for investigation. placement signal issue: no evidence of physical damage was noted on the returned product other than a kink at the king protector. the pump underwent electrical testing, and all motor current lines were found within specification limits. one of the dp sensor electrical lines (v-/v+) was found to be open during electrical resistance testing. further investigation into the dp sensor line revealed an open connection from the blue handle down to the dp sensor/catheter. the blue handle was open, and the catheter was sliced at the kinked catheter location; however, no damage was found to the dp sensor electrical line. the exact exit location could not be investigated due to the delicacy of the dp sensor electrical line. data shows the placement signal dropped to zero on 5-july, followed by the default value on pump flow. no other abnormalities related to the optical signal issue were found. no placement signal not reliable (psnr) alarm was recorded during this event. later, during the case run, data log recorded several placement signal recovery trends with "imptevt plug connect" alarms, which may indicate disconnection between the pump and console. no abnormalities related to motor current abnormalities were noted during the placement signal issue. the reported failure mode of optical signal issue was changed to placement signal issue as an investigated failure mode based on the returned product investigation. hemolysis: data log shows a motor current spike during the reported hemolysis issue on 15-july with some saturated and damaged trend on it, which later resolved. pump flow could not be considered because it was showing a default value due to the dp sensor issue. clinically, hemolysis was confirmed with ldh values greater than 1000 on 15-july with adequate preload. ldh values trended downward; however, the doctor decided to remove the pump. no abnormalities related to hemolysis were reported on the returned product. the cause of the hemolysis issue was most likely biomaterial, based on data log review. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2025-30784. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7054 | 1220648-2025-30785 | 22634766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22634766&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696407 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant placed an impella cp for the support in a patient with acute myocardial infarction/cardiogenic shock. the pump was placed by the femoral access and done via single access with a 7 french companion. the pump was explanted the same day as implanted with positional issues and a gastrointestinal bleed. the impella cp was escalated to an impella 5.5. the patients outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7055 | 1220648-2025-30786 | 22635529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22635529&pc=OZD | 07/18/2025 | 09/15/2025 |  | 07/28/2025 |  |  | 07/18/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603965 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 99 KG |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella 5.5 support. during the automated impella controller (aic) to aci transfer, the power cable was not completely plugged in. the alarm was disabled and support continued without interruption. the staff was advised on how to monitor support without the placement signal. the issue was resolved. no patient harm was reported. | It was further reported that the patient expired 3 days after explant. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue been completed. no product was returned for evaluation. data logs were reviewed and lt of second console using during case shows a sudden change in optical parameters consistent with an air gap at the pump impella plug with psnr alarm triggered. optical parameters and psnr alarm resolved after approximately 5 hours. clinical details noted a psnr alarm was noted and during aic to aic transfer, power cable was not completely plugged in. issue resolved once plug was inserted fully. the cause of the optical signal issue was most likely due to an air gap at the impella plug per clinical details and data log analysis. b1 revised to include adverse event based on the additional information received. b5 updated based on additional information received. d6b explant date added. h1 revised based on the additional information received. h6 revised based on the additional information received.
7056 | 1220648-2025-30787 | 22635576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22635576&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690950 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Pulmonary Hypertension | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient arrived status post st-segment elevation myocardial infarction percutaneous coronary intervention to left anterior descending artery with cp. cp removed however the patient required increase support with up trending pulmonary hypertension pressures and drop in mixed venous oxygen saturation, intra-aortic balloon pump (iabp)placed. overnight patient has required increase support pressor support. decision for 5.5 placement and iabp removal. the 5.5 was successfully placed. during support the patient had a run of ventricular tachycardia and was requiring higher doses of milrinone & levophed. in addition the impella measuring deep at 8cm, plan was to reposition with surgery team. the patient was resting on impella support as bridge to left ventricular assist device. | [5 empty] | D6: there is no explant date at this time. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7057 | 1220648-2025-30739 | 22635835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22635835&pc=OZD | 07/13/2025 | 09/08/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084302 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 108 KG | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Application Program Problem | Unintended Application Program Shut Down |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the console had an automated impella controller failure that occurred, and metrics went blank, in addition to it shutting down temporarily. the automated impella controller was switched out and the patient remained stable. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: on 7/2, 5.5 pump (b)(6) was connected to and ran continuously until 12:37 on 7/13 when the console unexpected rebooted. hemodynamic support was returned 29 seconds later at which time the user received an unexpected controller shutdown alarm. there was no evidence of a controller failure before the reboot so presumably that was referencing the unexpected shutdown alarm.device analysis: the console was returned for investigation but the failure mode was not reproduced despite replicating the console settings from the time of the complaint.root cause: the cause of the unexpected reboot could not be determined because the failure mode was not reproduced during testing. h8 usage of device was erroneously selected on the initial manufacturer device report.
7058 | 1220648-2025-30791 | 22635901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22635901&pc=OZD | 07/17/2025 | 09/16/2025 |  | 07/28/2025 |  |  | 07/18/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female | 113 KG |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a hematoma at the access site and had pump suction alarms. anticoagulation was withheld and two units of blood products were given to the patient. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleeding and low pump flow has been completed. pump was not returned for investigation. as per the clinical report, a steep insertion angle was reported during the case. the cause of the access site bleeding issue was most likely use related to steep insert angle. according to the clinical report, the pump had suction alarms while running on a higher p-level, which resolved at p6. the patient was also given 2 units of blood. data log shows no motor current spike or positioning issue during the reported suction alarm while running on a higher p-level during case start. suction was seen stabilizing later during the case while running on p-6. the placement signal trend was also trending during the suction event. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details.
7059 | 1220648-2025-30789 | 22636013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22636013&pc=OZD | 07/14/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/18/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after successful implantation of an impella cp, proper flow could not be established. the pump was removed and replaced with a new one. no patient harm was reported. | [5 empty] | Patient information, device information, implant and explant dates and initial reporters name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7060 | 1220648-2025-30793 | 22636072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22636072&pc=OZD | 07/17/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/18/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655224 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while repositioning an implanted impella 5.5, the impella 5.5 came out of position and the catheter was kinked. the impella 5.5 was replaced and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7061 | 1220648-2025-30790 | 22636084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22636084&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682705 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had hemolysis despite p level reduction. the pump inlet was posterior and the pigtail was caught in papillary muscle. the doctor did not want to reposition. volume was given and the hemolysis cleared up. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation was completed. hemolysis: complaint device not returned for evaluation. data log reviewed, no mc spike, no position issue, no placement signal issue or suction alarms observed. no indication for any issue observed. clinical stated, persistent hemolysis despite p- level reduction, physician aware and decided not to attempt repositioning despite inlet being posterior and pigtail caught in papillary muscle. hemolysis cleared up after volume resuscitation overnight. the cause of hemolysis issue most likely was patient condition related since it cleared up after volume resuscitation despite positioning issue. positioning issues: inlet being posterior and pigtail caught in papillary muscle, unknown how the issue occurred. the cause of positioning issue cannot be determined due insufficient clinical details. b1 device problem was inadvertently omitted on the initial manufacturer device report number. h6 medical device problem code was updated as erroneously entered on the initial manufacturer device report number.
7062 | 1220648-2025-30742 | 22636173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22636173&pc=OZD | 07/15/2025 | 09/10/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672952 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 68 YR |  | Male | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a high purge pressure alarm appeared during impella 5.5 support. tissue plasminogen activator was run through the purge and the pressures stabilized without issue. support continued uninterrupted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation has been completed. no product was returned. clinical details report that a high purge pressure alarm triggered on (b)(6) 2025. there were no kinks noted in the purge line. the patient's ptts were subtherapeutic, and administering tpa resolved the complication. data logs were reviewed and the alarm was confirmed. roughly 1 hour after a routine bag change, purge pressure began to gradually rise for 2 hours without a motor current spike. the pressure remained elevated for 45 minutes. during the event, the purge cassette was switched out, and 4 minutes later, purge pressure gradually trended back down into acceptable range over a 10 minute period. there were no further purge pressure issues for the remainder of the case. product was not returned for investigation. based on the gradual rise in purge pressure without a motor current spike and gradual resolution after tpa administration, the cause of the high purge pressure was most likely biomaterial buildup.
7063 | 1220648-2025-30792 | 22636299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22636299&pc=OZD | 07/02/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/02/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644204 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp then followed by an impella rp flex device for mechanical circulatory support (mcs). the patient experienced limb ischemia, bleeding and subsequently expired. the patient presented with chest pain and was emergently taken to the catheterization lab. the patient subsequently coded and was shocked eight times before proceeding with the impella cp device inserted via the right femoral artery during code blue procedure. flows were initiated and eventually the patient stabilized with return of spontaneous circulation. the culprit was a 99% occluded proximal right coronary artery which was treated with angioplasty and drug-eluting stenting. the disease to the left anterior descending and diagonal arteries were planned to be treated on a different day. a right heart catheterization was performed, and the physician proceeded with an impella rp placement via the right femoral vein into left pulmonary artery without issue. right-sided support was started with flows of 2.5-3.0 l/min achieved. the patient was "extremely" fluid responsive and no marked improvement with rp flex placement. echocardiogram in the procedure area showed left ventricular function of approximately 55% and right ventricular function slightly diminished, however, not dilated. the patient was sent to unit on bipella (impella cp and rp flex) support. approximately one day into bipella mcs, systemic heparin drip was ordered with goal of activated clotting time above 200 seconds. the patient had decreased pulses in both extremities with more noticeable symptoms on the impella side. a fem-fem bypass was performed to manage the ischemia. the patient continued to have worsening hemodynamics. in addition, abdominal bleeding was observed, for which multiple blood products were given. a bedside abdominal tap revealed blood in abdomen that was firm. the patient continued to deteriorate from the underlying condition and the decision was made to withdraw care. | [5 empty] | There is not enough evidence to exclude the impella cp as an associated factor in the patient's demise outcome given the device related limb ischemia event requiring surgical intervention in addition to the heparin for impella support and its unknown impact to the patient's abdominal bleed. this is one of two devices associated with the event. refer to medical device report for the impella rp flex sn (b)(6), date of event 03-jul-2025 and patient identifier ending in (b)(6). the impella cp device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7064 | 1220648-2025-30795 | 22636386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22636386&pc=OZD | 07/17/2025 | 09/25/2025 |  | 07/28/2025 |  |  | 07/17/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637508 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 85 KG |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump stopped unexpectedly during patient support. the automated impella controller was swapped, but attempts to restart the pump continued to fail. the pump was subsequently explanted as a result of the noted failure. there was no reported patient harm. | [5 empty] | The investigation of pump stop has been completed. based on the signatures seen in data logs and open motor cable lines on returned product, the cause of the pump stop was determined to be open electrical lines. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7065 | 1220648-2025-30751 | 22637059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22637059&pc=OZD | 07/11/2025 | 09/05/2025 |  | 07/28/2025 |  |  | 07/11/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658472 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and start of the device, the implanted impella cp pump was unable to achieve flows higher than 1 l/m. the pump was repositioned; however, low pump flow remained. consequently, the impella cp device replaced as a result of the noted low pump flow and the pci was completed. there was no harm to the patient as a result of the event. | [5 empty] | The investigation of low pump flow has been completed. data review: data logs confirmed low pump flow when pump started and remained to the rest of the case. pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of low flow was unable to be determined as limited clinical details were provided and no problem was found on the pump. | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed.
7066 | 1220648-2025-30757 | 22637786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22637786&pc=OZD | 07/13/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/13/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026641276 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, high purge pressure occurred with "purge system blocked" alarms. no kinks were observed. purge flow and pressure changed. tissue plasma activator was added to the purge. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7067 | 1220648-2025-30741 | 22639264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22639264&pc=OZD | 07/09/2025 | 10/20/2025 |  | 07/28/2025 |  |  | 07/14/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528356 |  | 0048-0024 JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed and when performing the percutaneous coronary intervention, the position waveform disappeared before leaving the catheter room. the placement signal not reliable alarm sounded. the position waveform disappeared for several hours after that, but the waveform reappeared 5.6 hours later and was displayed without any problems after. there was no harm to the patient. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. clinical details report that a placement signal not reliable (psnr) alarm triggered during the case for roughly 5 hours, but resolved. there was no harm to the patient as a result, and neither the pump nor console were replaced due to the complication. data logs were reviewed, and the alarm was confirmed. from the moment the pump was plugged in (before in the patient's body), signal-to-noise ratio (snr0 was never greater than 1500. prior to the pump even being started, snr started to trend down until it was so low that the ps became unreliable and the alarm triggered roughly 20 minutes into the case. during this time, contrast increased, but gain, light, and lamp were relatively stable. this suggests that there was no lack of light getting to the optical sensor or being reflected back. over time, snr gradually improved, making the placement signal (ps) calculation reliable again, which explains why the alarm resolved. imc logs were also reviewed, and it was confirmed that the optical sensor g0 calibration was done without issue. returned product was investigated and there were no abnormalities. all 5s parameters were within specification, and the pump passed laser continuity testing. the pump was then tested in the uson, being pressurized to 50, 100, 150, and 200 mmhg. during this time, ps readings were accurate, all 5s parameters were within specification, and snr never dropped below ~3500. based on the fact that snr was very low the moment that the pump was connected while other optical signal metrics were stable and there were no issues with the returned pump, there was determined to be no problem found with the pump. upon review of data logs and the returned pump, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated. b7 other relevant history, including preexisting medical conditions updated. d4 model number updated. d10 concomitant medical products and therapy dates updated.
7068 | 1220648-2025-30788 | 22639266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22639266&pc=OZD | 07/15/2025 | 07/28/2025 |  | 07/28/2025 |  |  | 07/15/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630577 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had ventricular tachycardia. the pump was pulled back 1 cm and no issues were reported after. support continued. additionally, the anticontamination sleeve broke at the hub. support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7069 | 1220648-2025-30794 | 22639346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22639346&pc=OZD | 07/17/2025 | 09/25/2025 |  | 07/28/2025 |  |  | 07/17/2025 | 07/28/2025 |  |  |  |  | 07/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657655 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed as care escalation from prior cp pump (sn (b)(6)) for the acute myocardial infarction and cardiogenic patient. the pump was supporting for 8 days when the pump stopped. the team attempted to restart and troubleshoot but, pump was not restarted and was explanted in the operating room. the patient survived and noted to be stable. | [5 empty] | The investigation of pump stop has been completed. the impella device was not returned for evaluation. data log review showed pump is attempted to restart 3x but unsuccessful. no motor current (mc) spike prior to pump stop with alarm 119, impella stopped. mc spikes on subsequent restart attempts with alarm 115, impella stopped (rpm deviations). mc spikes to 32767ma when restarts are attempted. the cause of the pump stop was most likely an open electrical line due to the pump stop occurring with no purge issues, no mc spike, triggered both alarm #115 and #119, and upon restart attempts the motor current spikes to 32767ma. however, the cause of the open electrical is not known. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smart assist) section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7070 | 1220648-2025-30737 | 22621565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22621565&pc=OZD | 07/12/2025 | 07/25/2025 |  | 07/26/2025 |  |  | 07/12/2025 | 07/26/2025 |  |  |  |  | 07/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695337 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a hematoma developed at the patient's access site during impella 5.5 support. the patient was taken to the cardiovascular operating room for an evacuation to treat the condition. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7071 | 1220648-2025-30706 | 22607876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22607876&pc=OZD | 07/11/2025 | 09/10/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590018 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. the long peel-away sheath was placed in right femoral artery and a kink was noted in middle of sheath which was contributed to the calcified femoral artery. when inserting the impella cp device through the sheath, the physician encountered difficulty advancing past kink. the physician was able to advance pigtail, inlet and cannula past kink; however, the motor portion of the device could not be advanced. the physician, therefore, decided to remove device, place a short sheath and try with a buddy wire. while removing impella cp device, the inlet of the device became stuck on kink and unraveled or unwound the cannula. the impella cp device and the sheath were successfully removed as a unit without harm to the patient. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation was completed. clinical details state that the 14frx25cm sheath kinked in rfa due to calcified artery (fluoro image provided by field rep showing calcification). when inserting impella through sheath, md had difficulty advancing past kink but was able advance pigtail, inlet and cannula past kink. the motor housing while removing the pump in order to exchange sheaths, the inlet got stuck on the kink and unraveled the cannula. sheath and pump were removed together. the product was returned and evaluated. the sheath was kinked in multiple spots in the sheath body. the pigtail and inlet remained just outside the sheath tip with the cannula torn/coil unraveled and laying inside the sheath body. the end of the coil was still attached to the cage on the pigtail end. the motor housing end of the cannula was sitting outside the sheath hub. difficulty inserting/removing pump through introducer: the cause of the difficulty inserting/removing pump through the introducer was most likely a sheath kink as seen on the returned product. introducer damage - mechanical: the cause of the introducer damage was most likely a sheath kink as seen on the returned product due to the patient's calcified vessel per clinical details. cannula/pigtail detachment - peripheral vessel: the cause of the cannula/pigtail detachment (peripheral vessel) was most likely device interaction since clinical details note that the inlet got stuck on the sheath kink and the cannula unraveled. product damage (sheath kink): the cause of the sheath kink was most likely patient condition related since clinical details note calcified rfa. ¿device interaction or damage by another device: the cause of the device interaction was most likely patient condition related since clinical details noted a calcified rfa which likely caused the sheath kink which then likely caused the difficulty removing the pump from the sheath and ultimate detachment.
7072 | 1220648-2025-30707 | 22607955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22607955&pc=OZD | 07/11/2025 | 09/08/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526147 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient thrashed about while in route to the hospital and caused the impella cp pump to pull into the aorta. the movement and dislodgement additionally resulted in the formation of a hematoma in the patient's right femoral artery. no intervention was required in response to the hematoma. the staff opted to remove the pump as a result of the positioning issue and the patient was escalated to impella 5.5 support. patient survived. | [5 empty] | The investigation of positioning issues and has been completed. the impella device was not returned for evaluation. positioning issue: the cause of the positioning issue was most likely patient movement related as the patient thrashed causing pump to be pulled into the aorta. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7073 | 1220648-2025-30708 | 22607976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22607976&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Retroperitoneal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella cp support developed bleeding at the impella access site with hematoma present. dressing removed, repo sheath pulled back, angle changed, and perclose tightened. as a result of groin site status, the impella was explanted. the next day, the groin continued oozing with continued small hematoma. the patient was taken to cat scan where a retroperitoneal (rp) bleed was identified. subsequently the patient was taken to the interventional radiology (ir) for treatment of rp bleed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
7074 | 1220648-2025-30709 | 22608010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608010&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658591 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 5.5 pump was placed to escalate care from the cp and extracorporeal membrane oxygenation for the acute myocardial infarction complicated by cardiogenic shock. the pump was placed and the site began to bleed after placement at the axillary site. the bleed was treated by delivering factor vii, protamine and blood products. the pump remains on support despite the harm. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7075 | 1220648-2025-30710 | 22608089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608089&pc=OZD | 07/09/2025 | 09/04/2025 |  | 07/25/2025 |  |  | 07/09/2025 | 07/25/2025 |  |  | 09/30/2025 |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female | 67 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had impella cp support placed for acute myocardial infarction complicated by cardiogenic shock. the sheath was retained due to bleeding. it was not peeled away and was stitched to the site. the patient was transferred to the tertiary center with the sheath still retained. additionally, there were other harms/injuries during the support. there was ventricular tachycardia and ventricular fibrillation due to the right coronary artery. per the medical doctor they were not impella related. the patient's urine was red after coding. an echocardiogram showed the impella was a little shallow but there were no plans to reposition the pump. the medical doctor does not think that the red wrine is related to the impella. the patient survived the procedure and is in critical condition. additional information was received that the patient had atrial fibrillation and not ventricular tachycardia. | [5 empty] | The investigation of access site bleeding, hemolysis, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding - major: the cause of the bleeding issue most likely was use related due to improper technique since the physician decided to keep peel away sheath in place and stitched around the peel away using a figure of 8 suture due to significant bleeding. hemolysis: data log reviewed: no motor current (mc) spike outside of p level change observed. suction occurred. quick resolved impella in ventricle alarms seen. placement signal not reliable occurred. no indication observed to confirm the cause of the reported hemolysis issue. red urine symptoms resolved without impella repositioning or blood administrated since the issue solved. the cause of hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7076 | 1220648-2025-30711 | 22608091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608091&pc=OZD | 07/11/2025 | 08/27/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the patient was shocked three times for ventricular fibrillation. additionally, there were issues with the impella's position. the impella was in the papillary and while attempting to reposition, the impella was accidentally pulled back. attempts to reposition were unsuccessful. the pump was explanted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. positioning issues: unable to recross the valve so pump explanted. the cause of the positioning issue was most likely use related since the physician accidentally pulled the pump back during repositioning and was unable to recross. vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. h6 health effect - impact code 4647 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30711.
7077 | 1220648-2025-30712 | 22608187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608187&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/25/2025 |  |  | 06/27/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653485 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella 5.5 device to a patient undergoing a high-risk percutaneous coronary intervention (pci) was unsuccessful. the device was unable to pass through the axillary artery, proceed down to the left ventricle after attempts and different techniques. consequently, the impella 5.5 implant was aborted, and the physician attempt with an impella cp which was successfully placed. after successful insertion of the impella cp, the physician proceeded with pci of the left main internal mammary artery to the left anterior descending artery with percutaneous coronary angioplasty and stents (x2) placement. the patient was transferred to the unit noted to be in stable condition. however, two days later, the patient experienced a ventricular tachycardia event that resolved. the patient would eventually expire three days later; care was withdrawn. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since pt¿s anatomy would not allow impella to proceed down to the left ventricle. | This is one of two devices associated with the patient's implant experience. refer to the medical device report for the ventricular tachycardia event involving the impella cp device serial number (b)(6)_date of event 29-jun-2025 with patient identifier ending in 3790. the impella 5.5 device was not received from the customer; report indicates the device was discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7078 | 1220648-2025-30713 | 22608383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608383&pc=OZD | 07/09/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/09/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658826 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 107 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient developed a vancomycin-resistant enterococci infection and yeast infection in the patient's blood. providers stated this is from the sternotomy site. due to infection, the plan is to wean the patient from impella support. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings, ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7079 | 1220648-2025-30715 | 22608602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608602&pc=OZD | 07/11/2025 | 12/03/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1751191 |  | 1000432 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella ecp was about to be placed for use with a console when there was an observation made of a battery failure. the impella ecp was used on a new, second console without harm or injury. the console malfunction/battery failure is reportable despite the impella ecp software and pump not being fully marketed for release. | [5 empty] | The investigation for the reported issue has been completed. data review: console logs were consistent with what was reported in the complaint. an instance of battery failure (transport connection blown/missing) alarm [event set #360] was observed in the imc logs. device analysis: console was tested after arriving in field service. the reported battery failure was able to be reproduced while testing this console on an engineering lab bench. batttest results revealed that both of the batteries in the console were fully depleted. the battery failure alarm resolved after swapping out both batteries with known good ones. conclusion: the root cause of the battery failure was identified as depleted batteries resulting in battery lockout. h6 investigation type, findings and conclusions. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7080 | 1220648-2025-30717 | 22608787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608787&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648636 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support developed access site bleeding. manual pressure was held but bleeding would restart after pressure was removed. the decision was made to attempt fem-stop which increased the bleeding. the physician removed the impella cp due to the bleeding. the patient survived and is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7081 | 1220648-2025-30718 | 22608923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608923&pc=OZD | 07/14/2025 | 12/03/2025 |  | 07/25/2025 |  |  | 07/17/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417541 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a controller error alarm appeared on the aic during patient support. at the time of the alarm, the placement signal suddenly disappeared. the aic was replaced in response to the failure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. see the attached image rt_log.png. device analysis: the console (b)(4) was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the optical bench was evaluated for 5s parameters and the analog signals were checked for distortion, but no issues were found. the aic was tested with known good pump and purge cassette and no issues were found. root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data) and the optical signal issue was due to a firmware issue (optical bench fw anomaly). h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation
7082 | 1220648-2025-30719 | 22608973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22608973&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663884 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | High Blood Pressure/ Hypertension | Cardiac Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the decision was made to place an impella cp emergently. the impella cp was inserted and initiated. they proceeded with primary percutaneous coronary intervention (pci). however, it was reported that during pci, there was a perforation requiring covered stent and/or balloon tamponade. it was noted that the wire trapped under the stent in the left anterior descending artery (lad). upon removal, lad artery reoccluded and ostial left main dissection, multiple losses of rhythm and cardiopulmonary resuscitation required. the patient was sent to the intensive care unit where they reported a growing hematoma at groin site. a fem-stop was placed and hypertension was managed with nitroglycerin get. it was noted that the patient was well supported so the impella cp was successfully weaned and explanted. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7083 | 1220648-2025-30720 | 22609330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22609330&pc=OZD | 07/11/2025 | 09/08/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682104 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 85 KG |  | White | Death | Yes | No | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and after the implant, distal pulses were unable to be found. a femorofemoral bypass was placed and pulses were still unable to be found. support was continued. additionally, the patient had red urine with hemolyzing labs. the p level was lowered. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. hemolysis: data log reviewed, resolved impella position in aorta alarm observed at p9 after 50 mins of support followed by motor current (mc) spike. following the mc spike, motor current level was high for the respective p-level, a rise in purge pressure, and a rise in pump flow were observed. flows were saturated for the respective p-level. also reversable flow alarm and resolved impella positioning in ventricle, unknown impella position alarms and placement signal low observed. no suction alarm seen. clinical stated, uop greater than 30 ml/hr and red, p level dropped from p-9 to p-6. last lab tests resulted lactate still hemolyzing. patient stopped making urine around since 2:00am of 7/12. the cause of the hemolysis issue most like was biomaterial ingestion since mc spike clearly observed in the reviewed logs after support initiated. saturated flow: failure mode (fm) saturated flow was added since high flow for the respective p-level was observed following the motor current spike. the cause of the saturated flow was most likely biomaterial per data log review since the saturated flow began after a mc spike. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. fm thrombus was added since a biomaterial ingestion event was observed in the data logs.
7084 | 1220648-2025-30721 | 22610012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610012&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658784 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 77 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that there was an attempt to place an impella 5.5 pump to support the patient. the team troubleshot by using lubricating but all advancement failed and the impella 5.5 was aborted, and the impella cp was instead placed. the impella cp allowed for the high risk surgery. the patient had been admitted in to have a mitral valve repair. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7085 | 1220648-2025-30722 | 22610174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610174&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600682 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient had hematuria. echocardiogram was performed and the pump was repositioned but the hematuria persisted. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7086 | 1220648-2025-30723 | 22610270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610270&pc=OZD | 06/30/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661902 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 supporting for a patient admitted in for a high-risk surgery. the support allowed for the procedure and was then explanted after 30 hours. later an impact study documentation noted there had been an adverse event of sustained ventricular arrhythmia with "probable" relationship to the use of the pump. the patient was treated by "obtain potassium for multiple ventricular runs". the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7087 | 1220648-2025-30724 | 22610296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610296&pc=OZD | 06/29/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 06/29/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664327 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of impella 5.5 device to a patient undergoing a high-risk percutaneous coronary intervention (pci) was unsuccessful. the device was unable to pass through the axillary artery, proceed down to the left ventricle after attempts and different techniques. consequently, the impella 5.5 implant was aborted, and the physician attempt with an impella cp which was successfully placed. after successful insertion of the impella cp, the physician proceeded with pci of the left main internal mammary artery to the left anterior descending artery with percutaneous coronary angioplasty and stents (x2) placement. the patient was transferred to the unit intubated and sedated noted to be in stable condition. two days after start of the impella cp mechanical circulatory support, the patient experienced ventricular tachycardia but resolved. the next day, the patient was taken for a right heart catheterization and new pulmonary artery catheter placement. the plan was to wean and explant the impella cp. however, three days later, the patient remained intubated and sedated, family decided to withdraw care, and the patient would subsequently expire. | [5 empty] | Although care was withdrawn, there is not enough evidence to exclude the impella cp device as an associated factor given the arrhythmic event that occurred three days prior to the patient¿s demise. the untoward impact, if any, is unknown. this is one of two devices related to the patient's implant experience. refer to manufacturing report number 1220648-2025-30712 for the report on the impella 5.5 device that was unable to be delivered. the impella cp device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7088 | 1220648-2025-30725 | 22610328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610328&pc=OZD | 07/06/2025 | 09/02/2025 |  | 07/25/2025 |  |  | 07/07/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1539910 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the console got wet during a flood. the customer is requesting the console get checked out to ensure it was not damaged. the event is not related to a clinical procedure therefore no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. data analysis: the logs show that the last time the console was used was during preventive maintenance on 2/5 as confirmed by the field service report. no logs from the complaint date. device analysis: the console was tested thoroughly to confirm full functionality and passed. there was no evidence of any damage due to the flooding from the complaint. root cause: there was no issue found. the device operated to specification during testing.
7089 | 1220648-2025-30726 | 22610340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610340&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026694525 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump supporting for a patient admitted in with cardiomyopathy and a planned heart transplant. the impella 5.5 was bridging to the transplant. the pump was placed when the intra-aortic balloon pump was explanted. on the second day of support, heparin-induced thrombocytopenia panel came back positive. the patient remains on support. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7090 | 1220648-2025-30727 | 22610471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610471&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026696881 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 98 KG | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the day after support was started, the patient began to develop intermittent arrythmia and ventricular tachycardia. the conditions were monitored as support remained ongoing. roughly fourteen days into support, the pump stopped unexpectedly. efforts to restart were unsuccessful and the impella cp was removed in catheterization lab. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
7091 | 1220648-2025-30728 | 22610580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610580&pc=OZD | 07/11/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600866 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the patient had a large hematoma at the access site. the impella cp was removed and the hematoma was repaired. the nurse stated the patient was stable post removal of the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
7092 | 1220648-2025-30729 | 22610781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610781&pc=OZD | 07/12/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638896 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp started normally and then the pump flow decreased to less than one liter per minute the access sheath was flushed with a 20 milliliter syringe prior to insertion in normal fashion. the impella cp was then removed and replaced with a new impella cp. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7093 | 1220648-2025-30730 | 22610951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22610951&pc=OZD | 07/13/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/14/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701624 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, the patient developed a groin bleed. three units of blood were given, femostop was placed, and anticoagulation was held to treat the bleed. impella support continued uninterrupted. the patient was reported to be in stable condition following the event. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7094 | 1220648-2025-30731 | 22611012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22611012&pc=OZD | 07/12/2025 | 09/25/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697488 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 67 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Atrial Fibrillation | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was admitted as st-segment elevation myocardial infarction and an impella cp was placed. while on the unit, the patient continued to decompensate requiring additional pressors. the patient was taken back to the catheterization lab for diagnostic left heart catheterization and possible impella cp. the decision was made to place impella cp for additional support. levophed was given 33 mcg/kg/min, vasopressin: 0.04 u/min, epinephrine: 0.8 mcg/kg/min, and heparin: 800 u/hr. the 6fr sheath was exchanged for impella 14 fr sheath after serial dilation. the impella cp was advanced otw into left ventricle without complications and started in auto to 3.1 l/min flow. the patient was in atrial fibrillation and cardioversion was performed twice to normal sinus rhythm. impella peel away sheath was removed, repositioning sheath was advanced. suction alarms sounded, pulsatility lost, and the patient's pressure dropped. angiogram revealed the impella kinked at groin and in the left ventricle. impella was pulled back and straightened without difficulty but remained in suction. repositioning was attempted several times. an echocardiogram showed adequate placement with no sign of effusion or any other problems causing suction. however, the decision was made to replace impella due to possible kink/damage from exchanging the repositioning sheath. impella cp was removed, and a new impella 14 fr sheath was advanced into left femoral artery to advance the impella cp replacement pump into left ventricle without any complications. the impella cp replacement pump was slowly ramped up to p-4 without any complications. however, there was suction alarm at performance level p-5 was noted. echocardiogram showed a small left ventricle and dilated right ventricle. normal saline bolus for volume was started. the patient went asystole and cardiopulmonary resuscitation was started. however, despite all efforts, the patient was unable to survive this event. | The complainant had a impella cp pump placed to support the patient with a st elevation myocardial infarction. the impella cp was inserted without any issues. the physician removed peel away sheath and advanced repo sheath. he felt resistance and immediate suction alarms. angio revealed impella had kinked and doubled on itself in the iliac area above the repo sheath and had advanced fully into the lv folded on itself. he was able to easily pull the impella back and straighten it out. the suction alarm continued despite multiple repositioning attempts as well as viewing position with echocardiogram. decision was made to replace impella with a new impella. |  |  |  |  | Medical safety reviewed the event and determined the following: regarding the impella cp, there is no reason to believe the arrhythmic event is related to the impella cp. the underlying condition of the patient appears to have contributed to this. however, dissociation cannot be made, and the arrhythmic event is being reported as a serious injury. there is an identifiable connection, as the impella cp device was present at the time of the arrhythmic event. during an unsuccessful attempt to place the first impella cp, the device kinked folded on itself, and patient had an episode of atrial fibrillation. another cp was successfully implanted. regarding the replacement or second impella cp, there was suction event; the patient may have needed more support per the note. the patient went into asystole, and this, however, is unrelated to the impellas. conservatively the report is being submitted as death for the (second) impella cp. although, the patient's underlying condition is likely the reason for the patient going into asystole requiring cardiopulmonary resuscitation. the investigation of low pump flow, product damage (catheter kink), product damage (cannula kink), and vf/vt/af/at has been completed. low pump flow: the cause of low pump flow was most likely cannula kink due to patient having small left ventricle (lv). product damage(catheter kink): the cause of product damage(catheter kink) was most likely patient condition related since the patient had peripheral vascular disease (pvd). product damage(cannula kink): the cause of product damage(cannula kink) was most likely patient condition related since the patient had small lv. vf/vt/af/at: the root cause of the vf/vt/af/at was unable to be determined due to insufficient clinical details. this impella cp report is one of three devices associated with the event. refer to the related manufacturer report numbers: 1220648-2025-46809 rp flex. 1220648-2025-46810 second impella cp which is associated with the demise outcome. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30731. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30731. b5 patient outcome and additional failure mode was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30731. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-30731. h6 health effect - clinical code 4582 was erroneously selected on the initial manufacturer device report number 1220648-2025-30731. h6 health effect - impact code 4647 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30731. h10 missing mso statement. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7095 | 1220648-2025-30732 | 22611523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22611523&pc=OZD | 07/12/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670060 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support. it was reported that ganglion block was successfully performed on a patient overnight following 12 shocks for ventricular tachycardia. it was further reported that the patients impella leg was cool and mottled. the patient is off of inotropes and vasopressors. pulsatility was improving slightly. support was continued until weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7096 | 1220648-2025-30733 | 22612278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22612278&pc=OZD | 07/12/2025 | 09/03/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681680 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in with acute myocardial infarction and cardiogenic shock. during support, it was reported that the access site of the pump developed a bleed. the bleed was treated by a manual hold, re-sutured, and was given blood products. after two days, the pump was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since significant bleeding occurred after patient's forward tension sutures became misplaced and hemostasis was achieved by re-suturing.
7097 | 1220648-2025-30734 | 22612526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22612526&pc=OZD | 07/12/2025 | 09/08/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689379 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and during the implant, the impella 5.5 was attempted but was unsuccessful at getting through difficult anatomy and severe aortic calcification. the implant was aborted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of difficult anatomy and severe aortic calcification. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
7098 | 1220648-2025-30735 | 22612825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22612825&pc=OZD | 07/12/2025 | 07/25/2025 |  | 07/25/2025 |  |  | 07/12/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657875 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp side appeared to be ischemic in a patient. the patient has received a total of four units of blood overnight. the patients legs have improved and support was continued. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7099 | 1220648-2025-30736 | 22612920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22612920&pc=OZD | 07/11/2025 | 09/11/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679372 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 116 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient on impella 5.5 and extracorporeal membrane oxygenation support. during support, it was reported that the patient tested positive for heparin-induced thrombocytopenia and anticoagulation was changed. additionally, the low placement signal alarm continued intermittently. the alarm was silenced at this time and support was continued. | [5 empty] | The investigation of optical signal issue and thrombocytopenia has been completed. the impella device was not returned for evaluation. optical signal issue: data logs show no hard failure of the optical sensor. "placement signal low" alarms noted at beginning of case. additionally some "suction" alarms present at the same time. persistent "impella position unknown" alarms triggered on first day of support. clinical details noted that pump was in good position with "low placement signal" alarms at beginning of pump insertion. the cause of the optical signal issue could not be definitively determined as no product was returned and no hard failures of the optical sensor were observed on the data logs, additionally, limited clinical details were provided. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
7100 | 1220648-2025-30714 | 22620769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22620769&pc=PYX | 07/10/2025 | 09/03/2025 |  | 07/25/2025 |  |  | 07/11/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675626 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella rp flex had a placement signal not reliable alarm. the impella rp flex continued to support the patient. there was no harm to the patient and the patient survived the procedure. | [5 empty] | The investigation has been completed. data log review shown ps railing to 3000 with no impact to 5s parameters. alarm #1052 for psnr (epm sensor failure) was triggered at that time. cvp increased to 300. the pump was returned and inspected. upon connection with the aic, pump flow calculation disabled warning appears. a kink in the catheter was discovered where the catheter attaches to the motor housing. however, this kink appeared to be derived from how the pump was packaged in its return journey post use back to abiomed. electrical testing showed abnormal sensor s+/s- at 3452 ohms (below the expected 4800-6100ohms) and v+/v- at 3031 ohms (below the expected 4800-6100ohms). when the pump was attempted to be started, it immediately stopped with mc high alarms. psnr alarm 1052 also triggered. the cause of the placement signal issue was most likely an electrical short due to low electric readings found on the returned pump. however, the cause of the electric short is not known. fm optical signal issue was changed to fm placement signal issue due to the investigation of the returned pump showing dp electrical short with no impact on the optical sensor. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7101 | 1220648-2025-30716 | 22620806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22620806&pc=OZD | 07/11/2025 | 09/24/2025 |  | 07/25/2025 |  |  | 07/18/2025 | 07/25/2025 |  |  |  |  | 07/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675418 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Complainant in japan reported the patient was on impella cp support and during support, there was a pump stop. the pump restarted after the procedure was performed and there was no problems with purge pressure and no thrombus was found. there was no patient harm reported. |  |  |  |  | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. the cause of the temporary pump stop issue was not determined since the product was not returned and sufficient clinical information, including data log, was not provided. b1 adverse event added based on the reported patient outcome. b2 death and date of death added based on the reported patient outcome. b5 updated to reflect the patient outcome. d6b explant date added. h1 type of reportable event revised to death. h6 health effect - clinical code 4580 and health effect - impact code 1802 added to reflect patient outcome. | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7102 | 1220648-2025-30655 | 22585707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22585707&pc=OZD | 07/06/2025 | 09/03/2025 |  | 07/24/2025 |  |  | 07/06/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670897 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that a patient on day two of support developed dark red urine. no other information on resolution of the issue was provided. the patient was successfully weaned off from the device two days later and the patient was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30655.
7103 | 1220648-2025-30656 | 22585840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22585840&pc=PYX | 07/07/2025 | 08/29/2025 |  | 07/24/2025 |  |  | 07/07/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1184600 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female | 102 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while setting up for a new case, an automated impella controller (aic) failed to recognize an installed purge disc. a case start was attempted multiple times, but the issue persisted and the decision was made to replace the console. despite a successful subsequent implant / start-up of support, the patient continued to deteriorate from their underlying condition and passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. imc logs revealed that the console issued purge disc not detected alarm (event set #402) several times. the console was returned and examined, a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly.
7104 | 1220648-2025-30657 | 22585885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22585885&pc=OZD | 07/08/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/08/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608655 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient suffered a gastrointestinal bleed during impella 5.5 support. heparin was stopped and blood products were given to manage the condition. support continued uninterrupted. two days later, a ct scan revealed that the patient had multiple brain bleeds. the patient expired on support two days later. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
7105 | 1220648-2025-30658 | 22585897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22585897&pc=OZD | 07/07/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/07/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658311 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured an unstable rhythm/ventricular tachycardia during impella 5.5 support. the patient was shocked and given a bolus of lidocaine to help stabilize the rhythm. as support continued, the staff expressed concerns for a potential gastrointestinal bleed due to a drop in platelet levels. nine units of blood products were given and multiple washouts were performed to manage the potential bleeds. however, eventually care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7106 | 1220648-2025-30660 | 22585969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22585969&pc=OZD | 07/07/2025 | 10/20/2025 |  | 07/24/2025 |  |  | 07/08/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139181 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a controller error alarm appeared during patient support. the placement signal disappeared shortly after. support continued and the signal reappeared roughly fifteen minutes later. there was no noted patient injury. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: two consoles are used with pump sn: (b)(6). (b)(6) is the first, and (b)(6) is the second. (note: neither of the consoles has been returned. no imc logs were available for either console on the complaint date, so data analysis was performed using lt and rt logs.) according to the lt log, (b)(6) does not have the 'controller error' alarm. according to the rt log, on the complaint date, (b)(6) 2025, the second console, (b)(6), displayed ¿controller error' event #1045, which was resolved automatically in 12 minutes. all signals received from the optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to a crc error. note: the 'controller error' was misreported against (b)(6); originally, it was associated with (b)(6). device analysis: not applicable, as (b)(6) is not required for the investigation. root cause: the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7107 | 1220648-2025-30659 | 22586066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586066&pc=OZD | 07/08/2025 | 09/10/2025 |  | 07/24/2025 |  |  | 07/08/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636888 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, there was a small leak from the 23fr peel away sheath. the doctor was not concerned about leak and continued to use the sheath for the procedure. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. introducer damage - mechanical: no images of product were returned for evaluation. clinical details noted a slow leak from the neck of the 23fr axillary introducer. no additional details were provided. the cause of the introducer damage could not be definitively determined as no product was returned and limited clinical details were provided. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30659. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30659. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30659. h1 type of reportable event was erroneously selected on the initial manufacturer device report number 1220648-2025-30659. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30659. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30659.
7108 | 1220648-2025-30661 | 22586069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586069&pc=OZD | 07/03/2025 | 07/29/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675952 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impact study reported that the patient on impella 5.5 support experienced a generalized tonic-clonic seizure, witnessed by nursing staff. oxygen was administered via ambu bag with 100% fio2 and assisted ventilation due to respiratory insufficiency, although spontaneous breathing was maintained. seizure activity was successfully terminated after administering 5 mg iv midazolam; however, partial respiratory insufficiency persisted. the anesthesia team was informed, and reintubation was performed successfully, with a pogo score of 100% and cormack-lehane grade i view. there were no signs of aspiration. during the seizure, catecholamine requirements increased sharply due to transient hypercapnia. the impella device was set at p8 but triggered ventricular tachycardia, likely caused by a suction event. after reducing the pump speed to p6 and administering fluids, the arrhythmia self-resolved. position control confirmed the inlet was more than 4 cm into the ventricle, with no evidence of dislocation on apical imaging. cranial ct and ct angiography were performed without complications during transport. imaging showed no acute intracranial pathology. | The user facility reported that the patient with impella 5.5 support had aneurysm spurium. primary immunodeficiency relatedness assessment with ¿probable¿ related to the impella device and the impella procedure. recovered/resolved the same day. patient survived. |  |  |  |  | B5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30661. g2 report source; study was missing from manufacturer device report 1220648-2025-30661. h6 health effect - clinical code was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30661. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7109 | 1220648-2025-30662 | 22586147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586147&pc=OZD | 07/09/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674376 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 86 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Pericardial Effusion | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp flex and intra-aortic balloon pump (iabp) patients condition deteriorating so the team made the decision to replace the iabp to an impella cp. the patient was brought to the cathrterization lab for the exchange. upon advancing the impella into the ascending aorta and prior to crossing the aortic valve, the patient became unstable, mean artierial pressure dropped into the 40's and the patient became bradycardic. multiple rounds of epinephrine and bicarbonate were given. during the loss in blood pressure, the medical doctor (md) attempted to advance the impella into the left ventricle (lv) to initiate support. the patient¿s descending and ascending aorta were noted to be very torturous and the md was unable to advanced the impella into the lv far enough to initiate support due to the impella catheter buckling in the descending aorta. the patient continued to be unstable. the decision was made to use a single access to place the 7 french sheath and to use a jr4 catheter with a stiff wire to help straighten the tortuous anatomy. when the jr4 and wire were in the ascending aorta, the md was able to easily advance the impella into the lv without resistance or difficulty. the jr4 and wires were removed and support was initiated on auto with flows 0.2-0.4 liters per minute with map in the 40s. multiple rounds of epinephrine and bicarbonate were continued to be given. continuous suction alarms occurring on auto and the patient was taken out of auto mode to p8. there continued to be suction alarms so the p-level was then reduced to p6. suction was intermittently resolved until the patient's heart rate and mean arterial pressure dropped again. they wer unable to stabilize the patient with continuous suction alarms on the impella cp. the md requested a stat echocardiogram to assess the impella and for possible effusion. upon arrival, the echocardiogram revealed the pigtail of the impella cp through the left ventricle wall and a large pericardial effusion was present. the decision was made to place a pericardial drain. initially, the md drained 200 cubic centimeters of blood and then hooked to a vacutainer. the vacutainer continued to fill up and the decision was made to hook a pericardial drain via a 3-way stopcock directly to the patient¿s central line to be able to give back blood the md was pulling out of drain. emergent blood was ordered and also given. the effusion would not stop draining and after pulling and giving back over a liter of blood, it was determined by the team that it was not going to resolve. the team spoke with family and the decision was made to not escalate care from that point on. the patient was made do not attempt resuscitation and was taken back to intensive care unit (icu) with the catheterization lab team and icu nurses. upon arrival to unit, the patient expired on impella rp flex and impella cp support shortly after. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system with smartassist for use during cardiogenic shock section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7110 | 1220648-2025-30663 | 22586149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586149&pc=OZD | 07/09/2025 | 09/04/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683686 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, bleeding was observed at the access site. steep angle, forward tension suturing, and manual pressure helped control the bleed. a femstop was placed at the site. support continued. additionally, the patient had several instances of hypotension and ventricular tachycardia/ventricular fibrillation after the implant. the patient was defibrillated. furthermore, the patient had pink urine after the implant, creatine continued to rise and continuous renal replacement therapy was started. there was significant acute kidney injury noted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Access site bleeding: the root cause of access site bleeding was most likely use issue since hemostasis was achieved by steep angle and forward tension suturing. vt/vf: the root cause of the hypotension/ventricular tachycardia/ventricular fibrillation was unable to be determined due to insufficient clinical details. renal failure: the cause of the renal failure was not determined due to insufficient clinical details. hemolysis: after reviewing the data logs, sudden spike in motor current and purge pressure was observed along with suction and impella in ventricle alarm. the root cause of hemolysis was most likely incorrect positioning since repositioning impella successfully the following morning resolved the issue. h6 health effect - impact code 4641 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30663.
7111 | 1220648-2025-30664 | 22586381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586381&pc=OZD | 07/12/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/12/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688229 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp to support a patient admitted in with medical history, arrhythmia, and wheelchair dependent. the vascular team was consulted due to the bilateral lower extremity necrosis. the pump was placed but there was hemolysis signs observed. the hematuria and lack of urine production prompted echocardiogram imaging for the hemolysis. continuous renal replacement therapy was begun. the pump remains on for support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7112 | 1220648-2025-30665 | 22586471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586471&pc=OZD | 07/12/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/12/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688143 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp via the right femoral artery for mechanical circulatory support (mcs). the patient was transferred to the unit remaining on impella mcs. approximately four days after the start of the impella mcs, it was noted that the patient had developed a hematoma on the right flank and scrotum after receiving cardiopulmonary resuscitation and the device association is unknown. blood products were administered. the patient would continue on support for an additional day before being successfully weaned and device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7113 | 1220648-2025-30666 | 22586650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586650&pc=OZD | 07/08/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/08/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support experienced a subarachnoid hemorrhage. anticoagulation was withheld and the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7114 | 1220648-2025-30667 | 22586770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22586770&pc=OZD | 07/08/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/08/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674670 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced limb ischemia. the team placed a reperfusion sheath to restore flow and pulses. the pump was supporting for 46 hours when care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7115 | 1220648-2025-30668 | 22587219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587219&pc=OZD | 07/14/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585171 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 86 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, a thrombus was noted in the left atrium and the left ventricle. support was continued. additionally, oozing was noted. the location of the oozing is unknown and the intervention for the oozing is unknown. the patient expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7116 | 1220648-2025-30669 | 22587453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587453&pc=OZD | 07/08/2025 | 10/07/2025 |  | 07/24/2025 |  |  | 07/08/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617091 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Care was eventually withdrawn, and the patient expired. | It was reported that a patient was implanted with an impella 5.5 and had to be cardioverted due to electrophysiological issues. placement signal issues occurred due to the patient getting defibrillated. "position unknown" and "impella in aorta" alarms appeared. impella 5.5 support continued. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. the data log was not provided and could not be retrieved from the remote server. the cause of the optical signal issue was not determined since the product was not returned and sufficient clinical information was not provided. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b2 updated to include death. b5 updated to include death.h1 updated to death. h6 updated to include death. d6b updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7117 | 1220648-2025-30670 | 22587592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587592&pc=OZD | 07/09/2025 | 09/05/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient had a cp pump being prepped for the high-risk percutaneous coronary intervention. the pump sensor failed and the pump delivery was aborted. another pump was used. no harm or injury from delay in support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the pump was not returned for investigation. data log shows multiple pump connect activities during the case start. on the last pump plug connection, the data log shows psnr and 5s optical signal trends indicating an air gap trend. the pump was replaced with a new one due to the optical issue. the cause of the optical signal failure was most likely an air gap, based on data log review; however, the cause of the air gap was unknown without returned product investigation or sufficient clinical details.
7118 | 1220648-2025-30671 | 22587618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587618&pc=OZD | 07/14/2025 | 09/03/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690265 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 74 KG | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Perforation of Vessels | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had been attempted to be implanted with an impella 5.5 and during the implant, the pump was able to advance into the ascending aorta but would not dive into the left ventricle. the pump was retracted and noticeable tissue in the inlet was advised to avoid axillary insertion and replacement of the impella 5.5. vascular endovascular repair was performed from the axillary pocket and blood products were given. the patients outcome is stable. | [5 empty] | The investigation for delivery issues, vascular damage, and thrombus has been completed. the device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the pump could not be inserted due to patient having small vessel size. the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7119 | 1220648-2025-30672 | 22587636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587636&pc=OZD | 07/09/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586042 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, there was bleeding around the access site. manual pressure was held and another perclose opened but unable to close hole. the doctor gained contralateral access and went up and over for balloon tamponade. decision made for vascular surgery to do a cut down and repair artery. after review, the doctor recognized stick was too close to bifurcation for percloses to be effective. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7120 | 1220648-2025-30673 | 22587759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587759&pc=OZD | 07/09/2025 | 10/30/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689366 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient had a 5.5 pump placed to support for care escalation from an intra-aortic balloon pump and awaiting transplant. the pump was supporting when extracorporeal membrane oxygenation was added and the patient suffered from new ventricular tachycardia and ventricular tachycardia storm. the pump remains on despite the harm while awaiting transplant and getting icd shocks. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7121 | 1220648-2025-30674 | 22587872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22587872&pc=OZD | 07/09/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697773 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the impella cp was placed but the 14x25cm peel away sheath was causing ischemia and was occlusive. peel away taken out and repo placed. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7122 | 1220648-2025-30675 | 22588072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588072&pc=OZD | 07/06/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664157 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured atrial fibrillation with rigth ventricular response (rvr) with a "possible" relationship to the impella device used: on (b)(6) 2025, the patient went into atrial fibrillation with rvr with accompanying low blood pressure. norepinephrine and vasopressin increased to max doses. amiodarone bolus 150 mg was given two times. the patient cardioverted x5 times and converted to normal sinus rhythm (nsr). magnesium was administered. the patient went back into afib with bps dropping again. ibutelide was given with change of heart rhythm to sinus brady in 50¿s. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7123 | 1220648-2025-30676 | 22588140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588140&pc=PYX | 04/30/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/09/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | On 2025-07-09 abiomed japan forwarded to complaints, publication for complaint review. entitled: "a staged approach from venoarterial extracorporeal membrane oxygenation (v-a ecmo) to protekduo circuit for right ventricular support in cardiogenic shock: a case report" which speaks to a bipella cp-rp support. two days after cp explanted the patient decompensated and when taken to the cardiac catheterization laboratory the rp was found to be dislodged. the team placed va ecmo due to this mispositioning. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7124 | 1220648-2025-30677 | 22588304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588304&pc=OZD | 08/03/2020 | 07/24/2025 |  | 07/24/2025 |  |  | 08/03/2020 | 07/24/2025 |  |  | 07/31/2021 |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1429966 |  | 0046-0023 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient had a left ventricle thrombus. the impella was removed, graft secured and sentinel devices successfully removed. upon interrogation of the carotid filters, a thrombus was found trapped. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7125 | 1220648-2025-30680 | 22588314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588314&pc=OZD | 07/10/2025 | 09/18/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608936 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | It was additionally reported that the patient expired on the subsequent device, an impella 5.5 (serial number (b)(6)) after experiencing a large hemorrhagic stroke during impella 5.5. support. the patient¿s family decided to withdraw care. the patient expired. this complaint involves 2 pumps: pump 1: impella cp serial number (b)(6). pump 2: impella 5.5, serial number (b)(6). this mdr specifically addresses impella cp with serial number (b)(6) which is the subject of this report. a separate medwatch report will be submitted for the other device associated with this complaint. | The complainant reported that thirteen days into impella cp support, the pump stopped unexpectedly. attempts to restart the pump were unsuccessful and the device was subsequently removed. the patient was escalated to impella 5.5 support following impella cp pump removal. patient is stable post pump replacement. |  |  |  |  | Rereview of the event occurred and noted the demise outcome is associated with impella cp as there is not enough information to exclude the impella cp as an associated factor given the pump stop/loss of support. refer to manufacturer report number 1220648-2025-45845 for the report on the impella 5.5. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of pump stop has been completed. impella cp was returned for evaluation. upon visual inspection, a large purge gap was present on the pump. data logs were reviewed and on day 12 of support, a motor current spike with speed deviation and full saturation of pump flow occurred 30 seconds before a pump stop #13 alarm. pump was attempted to be restarted multiple times with immediate pump stops. pump was able to be started with fully saturated flows for approximately 30 minutes before another pump stop occurred. clinical details noted that pump stop occurred on day 12 of support. the cause of the pump stop was due to bearing wear beyond indicated design life per instructions for use per data log review and returned product analysis. d9 device returned to manufacturer date added instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-30680. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-30680 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30680. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30680.
7126 | 1220648-2025-30678 | 22588315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588315&pc=OZD | 07/10/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681648 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, the pump need to be repositioned. staff attempted to reposition but the pigtail was firmly wedged underneath papillary muscle. the pump was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7127 | 1220648-2025-30679 | 22588355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588355&pc=OZD | 07/10/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  | 06/30/2025 |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024360157 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that eighteen days into impella cp support, a purge pressure low alarm appeared. no visible leaks were observed. three hours later, a clot ingestion resulted in a spike in motor current and an eventual pump stop. attempts to restart the pump were unsuccessful and swapping the automated impella controller did not resolve the issue. the patient remained supported via ecmo, but later passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿
7128 | 1220648-2025-30681 | 22588398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588398&pc=OZD | 07/10/2025 | 09/05/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680755 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Other | Required Intervention | [23 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a 5.5 pump for the patient admitted in with cgs and cardiomyopathy. the pump had supported for over 22 days the pump was explanted and stroke occurred simultaneously upon explant. the documentation from field rep notes the cerebral vascular accident to be hemorrhagic. patient survived the procedure. | [5 empty] | The investigation of stroke/cva has been completed. stroke/cva: the product was returned for investigation and a small kink was observed on the cannula. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. b2 selection was was erroneously entered on the initial manufacturer device report number 1220648-2025-30681. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-30681. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7129 | 1220648-2025-30682 | 22588580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588580&pc=OZD | 07/10/2025 | 09/23/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667762 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the pump dislocated and the pump was retracted into the aorta. the pump was disconnect and explanted. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: clinical details state that the pump was pulled into the aorta while the patient was being transferred. data logs are not applicable and product was not returned. the cause of the positioning issues was most likely patient movement since the clinical details state that the pump was pulled into the aorta while the patient was being transferred. g1 reporting contact email revised on manufacturer device report 1220648-2025-30682 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7130 | 1220648-2025-30684 | 22588713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588713&pc=OZD | 07/10/2025 | 09/23/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678152 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not advance through the patient's vasculature during attempted delivery. the implant was aborted and the patient was placed on ecmo for ongoing support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because insertion was aborted due to patient's vascular anatomy. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number.
7131 | 1220648-2025-30683 | 22588859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588859&pc=OZD | 06/06/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669902 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a protect iv that found that the patient endured anemia with a possible relationship to the impella and the impella procedure. the study stated, "hgb dropped from 11.9 -> 9.9 with groin access, expected with procedural blood loss; now stabilized at 10.0 on (b)(6) 2025. there was no visible bleeding at access site." the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7132 | 1220648-2025-30686 | 22588865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588865&pc=OZD | 07/10/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690182 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 74 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support developed gastrointestinal bleeding requiring transfusion of multiple blood products. eventually, patient care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7133 | 1220648-2025-30685 | 22588965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588965&pc=OZD | 07/10/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2025595638 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complainant reported the patient was on impella cp support and had a large hematoma at the access site. one unit of packed red blood cells was provided to the patient. support continued. patient is stable post issue. | [5 empty] | Section a5: patient race and ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7134 | 1220648-2025-30687 | 22588982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588982&pc=OZD | 07/11/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/11/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694476 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon placement of the 14fr peel-away sheath, the distal flow became sluggish, and a distal reperfusion sheath to proactively prevent limb ischemia. the physician did not want to attempt to use a repositioning sheath and opted to leave the introducer in place. the impella support level was turned to p-7 and the patient was sent to the unit on impella mcs. the patient was reported to be in stable condition following the event. support continued for an additional four days before being successfully weaned, and the device explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7135 | 1220648-2025-30688 | 22588987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22588987&pc=PYX | 07/10/2025 | 09/05/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026651124 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella rp flex could not be delivered via the femoral vein or the internal jugular vein. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unbale to pass through patient vascular anatomy despite multiple attempts and insertion was aborted. d6a and d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30688. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7136 | 1220648-2025-30690 | 22589059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589059&pc=OZD | 07/10/2025 | 09/05/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613405 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Defective Device | Patient-Device Incompatibility |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp could not be completed due to the patient's tortuous anatomy. the sheath kinked in multiple places and the pump failed delivery via the introducer sheath. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of product damage (sheath kink) and difficulty inserting/removing pump through introducer has been completed. the impella device was not returned for evaluation. product damage (sheath kink): the cause of the product damage (sheath kink) was most likely patient condition related due to the patients noted tortuous vessels. difficulty inserting/removing pump through introducer: the cause of the difficulty inserting the pump through the introducer was an introducer sheath kink preventing the pump from going through the sheath. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a and d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30690. h6 component code 463 was erroneously entered on the initial manufacturer device report number 1220648-2025-30690. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-30690.
7137 | 1220648-2025-30691 | 22589061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589061&pc=OZD | 07/11/2025 | 08/19/2025 |  | 07/24/2025 |  |  | 07/11/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122512 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and ethnicity) undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support (mcs) and during support a controller error (yellow alarm) triggered and prompted replacement of the automated impella controller. impella mcs continued, and there was no harm to the patient as a result of the event. the patient was noted to be in stable condition following the event. | [5 empty] | The automated impella controller device (aic) was not received from the customer at the time of this mdr writing; therefore, evaluation/analysis of the aic was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the crc error. device analysis: the console was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the optical bench was evaluated for 5s parameters and the analog signals were checked for distortion; no issues were found. additionally, the aic was tested with know good pump and purge cassette and no problems were observed. root cause: the root cause of the ¿controller error (opm comm crc invalid) - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue (optical bench fw anomaly) d9 device returned to manufacturer date added g2 foreign; was missing from manufacturer device report 1220648-2025-30691.
7138 | 1220648-2025-30692 | 22589097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589097&pc=OZD | 07/10/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658587 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following one day of impella 5.5 support, the patient endured sustained ventricular arrythmias. an amiodarone bolus was given and the dex infusion was decreased to stabilize the patient. two days later, a hematoma was seen at the access site. heparin was discontinued in response to the condition. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7139 | 1220648-2025-30694 | 22589099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589099&pc=OZD | 07/11/2025 | 09/15/2025 |  | 07/24/2025 |  |  | 07/11/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689040 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately two days after start of the support, a purge leak with low purge pressure occurred. a purge filter bypass was completed and resolved the issue. purge pressure and purge flow were within normal range thereafter. the patient continued on impella mcs with no harm and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. purge leak - pump: no product was returned for analysis. clinical mentioned leakage from the purge pressure reservoir after which they performed purge filter bypass. the root cause of the purge leak - pump was most likely due to a damaged sidearm as evidenced by clinical details but leak location was not confirmed as no product was returned. instructions for use: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿ h.10. instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30694.
7140 | 1220648-2025-30695 | 22589120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589120&pc=OZD | 07/10/2025 | 09/05/2025 |  | 07/24/2025 |  |  | 07/11/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677200 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and following the implant, bleeding/hematoma from impella groin site was observed. pressure was applied, the site was redressed, and medication (lidocaine with epinephrine) was injected around site to resolve the adverse event. later, overnight, the patient was taken to the catheterization lab for a pericardial drain and while in the catheterization lab the patient went into atrial fibrillation requiring cardioversion. the impella mcs continued for an additional three days before the patient was successfully weaned, and the device explanted with a survived outcome. | [5 empty] | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding is determined to be anti-coagulation management because activated clotting times were higher than the therapeutic range at the time of bleeding while moving the patient from cath-lab to icu. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7141 | 1220648-2025-30689 | 22589175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589175&pc=OZD | 07/10/2025 | 09/04/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688005 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 could not be completed despite attempting with an alternative wire and using propofol on cannula to assist with insertion due to the patient's anatomy. the implant was aborted. no patient harm was reported. | [5 empty] | A4 and a5, race, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of severe stenosis of subclavian artery causing resistance for advancement of impella despite using alternat wire. contrast revealed that subclavian artery cannot accommodate impella 5.5 and insertion was aborted d6a and d6b was inadvertently omitted on the initial manufacturer device report number.
7142 | 1220648-2025-30693 | 22589215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589215&pc=OZD | 07/10/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/10/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584342 |  | 0048-0045 |  | No |  |  | Not Available |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and approximately 6 hours after start of the support labs came back hemolyzed. impella position was noted mid-ventricle measuring 4.1cm, tricuspid annular plane systolic excursion of 1.9cm. and ejection fraction was 30-35%. groin site was free of hematoma or bleeding; the dressing was clean, dry, and intact. no further update regarding the labs was made. the patient continued on impella mcs for approximately 11 days until being weaned and device explanted with a survived outcome. the patient decided to move forward with coronary artery bypass graft surgery. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7143 | 1220648-2025-30696 | 22589301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589301&pc=OZD | 07/11/2025 | 09/08/2025 |  | 07/24/2025 |  |  | 07/11/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590013 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. calcified lesion was noted at exit of sheath in femoral artery. the physician was, therefore, unable to advance impella past the calcified lesion. the implant was aborted with the decision to implant via axillary approach next week. there was no harm to the patient as a result of the event. | [5 empty] | The investigation of delivery issues has been completed. the pump was returned and no kinks or abnormalities were observed. the short sheath was not returned for review. the cause of the delivery issues was most likely patient condition, as clinical reports the impella could not passed through calcification in the vessels. | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7144 | 1220648-2025-30697 | 22589378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589378&pc=OZD | 07/14/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694493 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, the patient went into atrial fibrillation and was cardioverted. support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7145 | 1220648-2025-30698 | 22589495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589495&pc=OZD | 07/14/2025 | 09/09/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689226 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during patient support, an impella cp pump was positioned shallow in the left ventricle. attempts to reposition the device were unsuccessful prompting the patient to be brought to the cath lab for more detailed troubleshooting and va ecmo cannulation. once in the cath lab, it was discovered that the pump was under the papillary muscle. the doctor was able to reposition the device under fluoroscopy. upon planned escalation to impella 5.5 support the following day, a large hematoma developed at the impella cp access site. the vascular team was called and the site required extensive repair as part of the removal process. blood products were given as part of the treatment. patient support continued with the newly implanted impella 5.5 pump following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues and access site bleeding has been completed. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. pump was returned for analysis. visual inspection found no issues on the pump. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no problem was found on the pump. the root cause of access site bleeding was most likely anticoagulation management related since act was greater than 261 seconds. d9 device returned to manufacturer date added g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-30698.
7146 | 1220648-2025-30699 | 22589874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22589874&pc=OZD | 07/14/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp never achieved flows about 1.5. the placement signal did not correlate with ao tracing. both the cardiology and heart failure team suspected error with the device despite efforts to reposition. the impella catheter appeared to have a kink under fluoroscopy resulting in minimized flow. a decision was made to replace the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7147 | 1220648-2025-30700 | 22591230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22591230&pc=OZD | 07/14/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694618 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that persistent suction issues coinciding with a low placement signal alarm appeared during impella cp support. the patient's left ventricle looked underfilled, but despite giving volume and blood products, the suction remained. the decision was ultimately made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7148 | 1220648-2025-30701 | 22591270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22591270&pc=OZD | 07/14/2025 | 09/24/2025 |  | 07/24/2025 |  |  | 07/16/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689326 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Great Vessel Perforation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported that following impella 5.5 implant, the patient began bleeding heavily from the newly placed mediastinal chest tube. the tissue was reopened down to the sternum and the damaged vein was sewn and clipped. protamine, two units of cryo, three units of platelets, four units of plasma, and four unit of blood were given to counteract the blood loss. as support continued, the patient developed hemolysis. it was noted that the patient's lv was small and decompressed post transplant. pfhgb measurements were around 38 and ldh levels were up to 3440. the team was aware of the complication and planned to monitor the situation. support continued uninterrupted. |  |  |  |  | The investigation of vascular damage - great vessel and hemolysis has been completed. the impella device was not returned for evaluation. vascular damage great vessel: the root cause of vascular damage great vessel cannot be determined since the product was not returned and insufficient clinical details were provided. hemolysis: there were suction alarms when the p-level was turned above p4. there were no motor current spikes but motor current and placement were very noisy and pump flow was also affected by patient condition. clinical did not confirm bad pump position. the cause of the hemolysis was most likely patient condition related as evidenced by the data log abnormalities aligning with clinical details concerning patient's condition. b2 death and date of death added. b5 revised to include patient outcome. d6b explant date added. h1 type of report revised to death. h6 health effect - impact code 1802 death added g1 reporting contact email revised on manufacturer device report 1220648-2025-30701 in accordance with updated procedures. | D6: device has no explant date at this time. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ vacular damage: great vessel. ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7149 | 1220648-2025-30703 | 22591525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22591525&pc=OZD | 07/13/2025 | 09/24/2025 |  | 07/24/2025 |  |  | 07/13/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644567 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, insufficient flow was noted and suction alarms appeared. in addition, impella in aorta alarms began to appear coinciding with inconsistent positioning curves and left ventricle pressure readings despite proper positioning of the pump in the ventricle. when reducing the p-level failed to resolve the issue, the pump was removed and red foreign material was found wrapped around the motor. it was unknown what the material was. the device was flushed and run in saline before being reinserted into the patient. however, upon reinsertion, the pump once again experienced suction issues. the pump was subsequently removed and replaced as a result of the ongoing issue. the patient was noted to be in stable condition. | [5 empty] | The investigation of low pump flow, optical signal issue, and thrombus has been completed. the impella cp was returned for evaluation. low pump flow: upon visual inspection, small biomaterial was present next to impeller blade. no other abnormalities with pump noted. pump was connected to console and when started, the small biomaterial was kicked off the pump. the pump was able to run for approximately 10 minutes and no low pump flows reproduced. data logs were reviewed and real time (rt) log of case showed variable motor current with multiple instances of motor current spikes with speed deviations during initial pump run. during short pump run outside of the body, one motor current spike during outside of body run had a large step up in pump flows, resolving low flows, likely when material was kicked off the pump at that time. second motor current spike when pump was run outside of body shortly after with slight decrease in pump flows before pump was turned to p-0. when pump was re-inserted into body, suction alarms again as soon as pump was started but no motor current spike was present. clinical details noted that insufficient flows and suction alarms were present after pump was started. pump was explanted and clinical details noted something red wrapped around impeller blade. pump was run in saline outside the body and red material flew out of pump and suction alarms disappeared. pump was then re-inserted into the body, however, same issues with suction and insufficient flows occurred. the cause of the low pump flow was most likely biomaterial ingestion based on data log analysis showing ingestion signatures during pump run. optical signal issue: small biomaterial present on returned pump. optical 5s parameters of returned pump showed no abnormalities or hard failures of the optical sensor. pump passed laser continuity testing. data logs were reviewed and rt log showed motor current and pump flow was low and variable at beginning of case with multiple motor current spikes. "impella in aorta" alarms triggered during first instance of pump running also noted but was resolved a few minutes into pump run. clinical details noted false detection of "impella in aorta" alarms triggered, inconsistent positioning curves and left ventricle (lv) pressures observed. the cause of the false position in aorta alarm was most likely due to biomaterial based on data log analysis and no abnormalities on returned pump. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
7150 | 1220648-2025-30702 | 22591692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22591692&pc=OZD | 07/13/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/13/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658023 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp support a patient. the pump was supporting when there were signs of hemolysis and a retroperitoneal bleed. the team gave blood and decided to explant the impella cp. the patient continued on extracorporeal membrane oxygenation and survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7151 | 1220648-2025-30704 | 22591853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22591853&pc=OZD | 07/14/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/14/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688931 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, it was found that the only flow to the left leg was from profunda. swan attempted to be placed without success in bilat fem veins due to anatomy. additionally, a perforation occurred to the left anterior descending artery. the clinical rep stated that they did not believe it was impella related however, they were not sure when it occurred. the patient expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: limb ischemia: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ ventricle perforation: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7152 | 1220648-2025-30705 | 22592613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22592613&pc=OZD | 07/11/2025 | 07/24/2025 |  | 07/24/2025 |  |  | 07/11/2025 | 07/24/2025 |  |  |  |  | 07/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668730 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day bleeding at the access site was noted, hemoglobin has dropped to 6.5 g/dl and two red blood cells bags were administered. additional information noted that patient underwent rectum surgery two days prior to the impella implant and was anemic with hemoglobin around 7.5 g/dl. while in weaning process of ventilator, the patient was hemodynamically unstable and was diagnosed with takotsubo. after the impella cp device implantation, there was a "minimal" oozing in the groin the next day. after femstop (just with bandage, not with balloon pressure) was applied for three hours; oozing stopped. the low hemoglobin was referred to the rectum surgery. it was noted the patient was stable while on the impella until successfully weaned and explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7153 | 1220648-2025-30601 | 22574123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574123&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/23/2025 |  |  | 06/24/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651703 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 112 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from acute myocardial infarction was implanted with an impella cp for mechanical circulatory support (mcs) and encountered a device sensing problem. the patient arrived with shortness of breath and was thought to have pneumonia. however, an x-ray suggested congestive heart failure. the patient arrested, was intubated and transferred to a different facility. the patient was decompensating. a ventricular septal defect from a late presenting inferior myocardial infarction was found. the patient was taken to the catheterization lab for impella cp placement transfer to a higher-level care facility. immediately after the impella cp placement, start of the support, no optical sensor was displaying on the automated impella controller. the physician, uncomfortable with the absent waveforms, decided to explant the impella cp and replace it with a new impella cp. there was no harm to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. based on the patterns seen in 5s parameters on data logs that resolved, the root cause of the optical signal issue was most likely an air gap between the console and patient cable plug.
7154 | 1220648-2025-30602 | 22574181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574181&pc=OZD | 07/05/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/05/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support and extracorporeal oxygenation membrane (ecmo). shortly after start of the impella mcs, flow was noted to be sluggish to the distal extremity. distal perfusion catheter was inserted in the unit. the patient continued on support for an additional 11 days before being escalated to an impella 5.5 using axillary and 8 millimmeter hemashield gold graft. successfully the patient was also weaned from the ecmo and decannulated. the patient was on multiple pressers/inotropes, intubated and sent to the unit to recover. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7155 | 1220648-2025-30603 | 22574249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574249&pc=OZD | 07/05/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/05/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639377 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support experienced a large stroke and care was withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7156 | 1220648-2025-30604 | 22574354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574354&pc=PYX | 07/06/2025 | 09/03/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026671845 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella rp flex was removed sooner than intended due to placement signal failure. the patient survived. | [5 empty] | The investigation has been completed. the pump was not returned for investigation. the data logs confirm pattern associated with sensor drift with placement signal increasing as pump flows decreased with stable motor current readings and no p-level change or ingestion events. the root cause of the placement signal issue was most likely due to sensor drift of the differential pressure sensor as evidenced by pattern in the logs. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7157 | 1220648-2025-30605 | 22574447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574447&pc=OZD | 07/06/2025 | 09/03/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539869 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a decision was made to place an impella cp. the impella cp was placed via radiofrequency ablation without incident. during support, it was noted that the motor current spiked on the automated impella controller. the decision was made to replace the impella cp. upon removal, a thrombus was visualized on the inlet of the cannula and in the right common femoral artery. no distal flow was visualized under fluoroscopy image. vascular surgery was consulted. cut down and thrombectomy were performed. distal perfusion was restored. a new impella cp was placed in the left groin with no issues and the patient is stable. | [5 empty] | The investigation of thrombus and pump stop has been completed. per data log investigation, low pump flow was added as a failure mode. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. data log review: data logs showed suction alarms on the day of the complaint. the motor current and speed spiked at the time of these alarms. at the time of the first suction alarm, pump flow level dropped significantly from 3 l/min to almost 0 l/min. p-levels were lowered, and motor speed deviations were observed. at time of second suction alarm, p-levels were lowered to p-0. the purge pressure increased but did not reach threshold, so no purge alarms triggered. the placement signal also drops significantly to less than 10 mmhg, potentially indicated removal of the pump. pump stop: data log review showed an impella stopped - retrograde flow alarm. this occurred just after the p-level was turned to p-0. the placement signal then dropped to zero, indicating the removal of the pump from the patient's body. there was no issue found during the case since the user turned the pump to p-0 and therefore got alarm 18 (impella stopped - retrograde flow). per impella cp instructions for use a setting of p-0 will result in retrograde flow when the impella catheter is placed across the aortic valve. the device functioned/operated to specification. low pump flow: data logs showed suction alarms on the day of the complaint. at the first suction alarm, the motor current spiked, flow levels dropped significantly, and they did not resolve as p-levels were lowered to off. the pump was returned for investigation and biomaterial was observed around the inlet cage and around the impellar. a flow test was run on the pump at p-9, and flow levels varied from 3.8 l/min to 2.6 l/min, which is low for the given p-level. the cause of low pump flow was most likely biomaterial ingestion, as flows dropped at the time of the suction alarm and motor current spike, and biomaterial was observed at the impellar during product investigation. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7158 | 1220648-2025-30606 | 22574549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574549&pc=OZD | 07/07/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/07/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672065 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intracranial Hemorrhage | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and exhibited neurological changes that prompted a computed tomography (ct) to be performed. the ct confirmed an irreversible brain bleed. the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
7159 | 1220648-2025-30607 | 22574584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574584&pc=OZD | 07/08/2025 | 11/30/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119120 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an abiomed biomedical engineer noted that an automated impella controller was not recognizing catheters. there was no report or indication of harm to a patient as a result of the event. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: the reported complaint was performed on 07/07/2025. in the case start wizard at step three, the pump connection was most likely not recognized, or the white catheter plug was not completely inserted into the aic, so after 20 seconds, epm was reset automatically. following the epm reset, the pump connection still was not recognized, confirming the reported failure mode. the case start wizard was then restarted, and at the same step, the pump connection again was not recognized. finally, the aic automatically shut down due to no user interaction. device analysis: the console (b)(4) was tested after arriving in fs. the reported issue was not reproduced; the console successfully detected the pump and purge cassette, as well as the purge disc. furthermore, the console was tested with a known good purge cassette and pump, and no other issues were observed. root cause: the root cause of the aic not recognizing the pump was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation
7160 | 1220648-2025-30608 | 22574701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574701&pc=OZD | 06/24/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 06/24/2025 | 07/23/2025 |  |  | 07/31/2025 |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440505 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 91 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a gastrointestinal bleed. heparin was withheld. later, care was withdrawn and the patient expired. | [5 empty] | The root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided.
7161 | 1220648-2025-30609 | 22574762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574762&pc=OZD | 07/08/2025 | 09/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682797 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. after ten days of support, the device experienced a pump stop after the automated impella controller indicated a jump in motor current and a 'motor current too high' alarm. three attempts to restart the pump were unsuccessful, which led to replacement of the pump. a new impella cp was successfully implanted and operated without any noted issues. the patient was stable throughout the event. | [5 empty] | The investigation of pump stop has been completed. no product or data logs were returned for evaluation. clinical details noted that a motor current spike and pump stop was observed on day 10 of support. the cause of the pump stop could not be definitively determined as no product or logs were returned for analysis and limited clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock,, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7162 | 1220648-2025-30610 | 22574856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574856&pc=OZD | 01/06/2025 | 11/03/2025 |  | 07/23/2025 |  |  | 06/25/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210876 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death |  |  |  |  | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support. while on support, placement signal unknown alarms. echo performed showed proper placement, mid left ventricle pointed towards apex, inlet measuring 5.5cm from atrioventricular (av). physician did not feel the need to reposition or make any changes. the automated impella controller (aic) was returned for investigation. this complaint was created for the aic. eventually, care was withdrawn and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: as confirmed in 25-40372-1, there was a downward trend of placement signal throughout the case until the placement signal jumped to 3000 mmhg for the duration of the case. device analysis: ic3990 was returned for analysis but the console performed as expected. 5s was good as was ccd array. root cause: there was no problem found with ic3990. the optical signal issue was caused by the pump. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7163 | 1220648-2025-30611 | 22574971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22574971&pc=OZD | 07/06/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp that was placed on a patient for support due to cardiac arrest. it was reported that they were unable to obtain distal pulses and angio performed via hand injection via wire re-access port showing no distal flow. the decision was made to place 5fr antegrade perfusion catheter on right and retrograde 5fr sheath in left femoral artery and connected sheaths as external fem-fem bypass. angio then revealed flow distal on right, however unable to doppler pulse. arterial doppler ordered and to be done on arrival to intensive care unit. dressing angle matched and impella secured. leg immobilizer in place. patient transferred to intensive care unit. during support the impella was weaned to p5 and patient tolerating well. plan to continue to wean and possibly explant soon. fem-fem bypass intact and working well with doppler pulses in bilateral lower extremities. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7164 | 1220648-2025-30612 | 22575108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575108&pc=OZD | 07/06/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676919 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that impella cp was placed to support percutaneous coronary intervention procedure. during percutaneous coronary intervention it was noted that the patient had a hematoma around 14fr short peel away. sheath was removed and repo placed. after there was an arterial bleeding around the sheath. physician put lots of forward tension on sheath and held manual pressure. the physician also went up and over with a balloon to help control bleeding. the vascular surgeon was consulted and did a cut down. after working for around 2-3 hours, the surgeon was able to get bleeding under control. during surgery, the patient received 15 units packed red blood cells , 15 units plasma, and 3 units platelets. the pocket was partially closed and the blue impella hub was sutured down. throughout the procedure the physician also held pressure on the opposite groin because there was a hematoma on that side as well. at the end of the procedure, the patient¿s pressor requirements were decreasing. the patient was transferred to intensive care unit. over the next couple of hours, the patient coded a few times, and family made the decision to not resuscitate the patient if coded. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
7165 | 1220648-2025-30613 | 22575139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575139&pc=OZD | 07/08/2025 | 09/22/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676586 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male | 150 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, bleeding was noted at the impella access site, with one wound dressing soaked in blood. the patient was given one blood bag, which elevated hemoglobin levels from 7.5 to 7.9 grams per deciliter. there were no signs of limb ischemia. it was later determined that the patient had a coagulation disorder related to the access site bleeding. fibrinogen was administered and the bleeding resolved. the impella cp continued to operate without reported issue. no patient harm was reported. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be patient condition because it was reported that patient has coagulation disorder causing the access site bleeding and the bleeding was managed with best practices. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7166 | 1220648-2025-30614 | 22575264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575264&pc=OZD | 07/06/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690293 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with an impella cp to vent left ventricle and for support with use of the venoarterial extracorporeal membrane oxygenation. it was reported that the physician placed impella in normal fashion. however, the pump was suspected to have entrained clot in the left ventricle that possibly developed from chest compressions. original impella cp explanted and exchanged for new device. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7167 | 1220648-2025-30615 | 22575374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575374&pc=OZD | 07/06/2025 | 09/03/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691769 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that patient was placed on cp support for st-segment elevation myocardial infarction and planned escalation to the 5.5 pump. the axillary insertion of the 5.5 pump failed due to tortuous anatomy. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to cross through tortuous aortic arch. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock, section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7168 | 1220648-2025-30616 | 22575413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575413&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/23/2025 |  |  | 06/24/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664288 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 112 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Unspecified Vascular Problem | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from acute myocardial infarction was implanted with a new impella cp for mechanical circulatory support (mcs) after the first one malfunctioned. after placement of the new impella cp, the decision made to not remove the introducer. the patient was sent to the unit on support with a plan to transfer to a higher-level care facility. later in the afternoon, the patient was transferred to the higher-level care facility maxed on pressor support and extremely tachycardic. the decision was made to cannulate for venous-arterial extracorporeal membrane oxygenation (va ecmo) and place a pulmonary artery catheter (pac). transesophageal echocardiogram showed a large ventricular septal defect and papillary muscle rupture with severe mitral regurgitation. the right ventricular was large but filling pressures on pac was low signifying adequate unloading. the patient, critically ill, continued on impella cp mcs and planned for rest and recover before escalating to an impella 5.5. four days later, the patient, still critically ill, was having signs of hemolysis. it was noted that the team believed the hemolysis was related to interruptions from va ecmo during the night opposed to the impella cp mcs. shortly thereafter, it was noted there were no more signs of hemolysis. urine was clear. the plan to upgrade to an impella 5.5 remained and was targeted to be completed with ventricular septal defect repair and mitral valve surgery. the patient continued on impella cp mcs and va ecmo support. following, four days later, the patient underwent ventricular septal defect repair, coronary artery bypass graft surgery times one and mitral valve surgery. the impella cp was explanted through the 14 french sheath and artery repair took place at the end of the case. the survival outcome at impella cp explant was survived with a bridge to an impella 5.5. | [5 empty] | Hemolysis: the pump was not returned. data logs were reviewed and no abnormalities were noted of the night of the compliant occurrence. suction alarms did occur on other days and an motor current (mc) spike occurred without suction alarms leading up to it. the cause of the hemolysis was not determined since product was not returned, inconclusive data log review, and insufficient clinical details. vf/vt/af/at: the root cause of the tachycardia was unable to be determined due to insufficient clinical details. | As there is no clear dissociation between the hemolysis and impella cp, the hemolysis has been captured as a health effect: clinical code. this is one of two impella cp's associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-30601 for the report on the initially implanted impella cp that malfunctioned. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7169 | 1220648-2025-30617 | 22575496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575496&pc=OZD | 07/06/2025 | 08/22/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674428 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp placement was aborted due to bilateral iliac arterial constriction. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of delivery issue was most likely patient condition since implant aborted due to bilateral iliac arterial constriction.
7170 | 1220648-2025-30618 | 22575585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575585&pc=PYX | 07/06/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/06/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026686630 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient had a rp flex placed to support the acute myocardial infarction with cardiogenic shock with multiple cardiac and systemic comorbidities. the pump was placed and there was bleeding that prompted the team to pause heparin and femstop placement. later due to dropping platelets the team gave blood while decisions being made. patient condition poor and long term decisions being made as recovery chances discussed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
7171 | 1220648-2025-30619 | 22575644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575644&pc=OZD | 07/09/2025 | 09/05/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650842 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp had purge system open alarm and noted a complete break between purge reservoir and filter after it was pushed up against bed accidentally. a purge system bypass was successfully done. patient survived. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. clinical details noted that "purge system open" alarms were triggered with a complete break between the purge reservoir and purge filter after it was pushed up against the bed accidentally. excessive force was applied. a purge bypass was performed successfully but clinical team ultimately decided to replace the pump. the cause of the purge leak - pump was due to damage to joint reservoir to filter connection based on clinical details noting purge reservoir and filter fully separated when excessive force was applied as sidearm was pushed up against accidentally. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
7172 | 1220648-2025-30620 | 22575759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575759&pc=OZD | 07/07/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662843 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during the implant, slow distal flow was observed on contrast imaging after removal of the repositioning sheath and guidewire. an external femoral-to-femoral bypass was placed to address the issue. transesophageal echocardiogram showed significant right ventricular dilation and mitral valve regurgitation. the bypass was later removed and a hematoma was observed at the left groin site. a sandbag was placed to prevent progression of the hematoma. heparin had been stopped and the patient received two units of fresh frozen plasma. the hematoma resolved. the patient was cardioverted for atrial fibrillation with rapid ventricular response. positive distal pulses were present. the impella cp was explanted and the patient outcome was reported as stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7173 | 1220648-2025-30621 | 22575801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575801&pc=OZD | 07/09/2025 | 09/04/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  | 02/28/2026 |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467799 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, bleeding started happening through the creases of the red plug that was secured above the patients belly button. pump was replaced. 2 suture holes were found in the catheter likely caused when surgeon was adding extra stitches to the anastomosis or when closing the pocket. | [5 empty] | The impella device has been received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation of product damage (hole in catheter) has been completed. the pump was returned and inspected. the cause of the product damage (hole in catheter) was use related due to the holes found in the catheter at the 31cm mark causing the blood ingress to the red pump plug.
7174 | 1220648-2025-30622 | 22575936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22575936&pc=OZD | 07/10/2025 | 10/20/2025 |  | 07/23/2025 |  |  | 07/10/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659639 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that their impella cp pump lost its placement and left ventricle signal while being used to support a patient. the pump continued to run and the signal reappeared after a short time. this issue occurred a second time and was also resolved after a short delay. the physician finished the procedure with the same device without any problems. there was no report of patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. clinical details report intermittent controller error alarms during a ventricular tachycardia ablation procedure. data logs were reviewed, and the controller error alarm was confirmed. roughly 20 minutes into the case, the alarm triggered, and all optical signal metrics railed to some multiple of -16384 two times over a five minute period. there were no further issues with the optical signal for the remainder of the case. returned product was investigated and there were no abnormalities noted. there was no visual damage to the pump, it passed laser continuity, and all 5s parameters were within specification when connected to a console. it was also uson tested up to 150 mmhg and there were no issues with optical signal metrics. upon review of data logs and returned product, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated.
7175 | 1220648-2025-30623 | 22576036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576036&pc=OZD | 07/09/2025 | 08/26/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691675 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp insertion was aborted due to the patient¿s anatomy. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7176 | 1220648-2025-30624 | 22576094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576094&pc=OZD | 07/09/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689420 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed in a patient with an effusion from the left anterior descending which failed percutaneous coronary intervention. while on support, there was bleeding complications in the groin. the pump was explanted and escalated to an impella 5.5. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7177 | 1220648-2025-30625 | 22576283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576283&pc=OZD | 05/27/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660258 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured left subclavian deep vein thrombosis with a "possible" relationship to the impella device used: ¿lue duplex showed l subclavian dvt. team believed clot developed while on impella. impella was explanted on 5/27/25, patient was switched to intravenous heparin then eliquis. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
7178 | 1220648-2025-30626 | 22576366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576366&pc=OZD | 07/14/2025 | 09/09/2025 |  | 07/23/2025 |  |  | 07/14/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659394 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient who came into the emergency room with st-elevation myocardial infarction, biventricular failure and ventricular tachycardia arresting. the impella cp was placed and then the impella rp flex was placed. the impella cp pump had low flows despite all measures and troubleshooting and was removed and replaced. a new impella cp and impella rp flex support the patient while awaiting the family decisions as not a candidate for surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the pump was returned for investigation. no physical damage was noted on the returned product. significant biomaterial was found on the impeller and could not be removed. the pump was then tested in a bucket of water, where low pump flow was reproduced. data log shows suction alarms, low pump flow, and motor current spikes from the start of the case run, which supports the biomaterial ingestion event. thrombus failure mode was added as an investigated failure mode based on data log review and returned product investigation. the cause of the low pump flow issue was biomaterial, based on data log review and returned product information. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added.
7179 | 1220648-2025-30627 | 22576515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576515&pc=OZD | 07/09/2025 | 09/19/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681152 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male | 96 KG |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump experienced a sudden increase in motor current to above 1000 ma followed by an unexpected stop. attempts to restart the pump were unsuccessful and the pump was removed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. data logs were not available for investigation. returned product was electrical tested, and all 3 motor phases were within specification. there were no signs of biomaterial in the outflow cage/cannula. there was dried biomaterial/blood wrapped around the motor housing, however, the way that it was wrapped suggests that it was picked up on explant, and unlikely to be related to the reported pump stop. inspection of the outflow cage revealed an enlarged purge gap. the pump was then attempted to be restarted, and it was unable, triggering alarm 13 (high motor current pump stop). based on the fact that the pump was on the 14th day of use, beyond the design indication of 8 days for impella cp and returned product had an enlarged purge gap, the cause of the pump stop was determined to most likely be bearing wear.
7180 | 1220648-2025-30629 | 22576646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576646&pc=OZD | 12/16/2024 | 10/03/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | NA |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Publication from copenhagen entitled "unexpected percutaneous ventricular assist device failure in a complex case of infarct-related cardiogenic shock" was received and noted the following: the patient experienced acute onset of chest pain at midnight accompanied by near-syncope but did not seek medical attention until the following morning. the pre-hospital ambulance electrocardiogram demonstrated anterior st-segment elevations with tachycardia, with a frequency of up to 130 beats/min). the patient was directly transferred for acute coronary angiography (cag) at a tertiary heart center. on arrival, the patient was clammy and pale with a blood pressure of 104/72 mm hg. an acute bedside echocardiogram revealed a left ventricular ejection fraction of approximately 10%, with no substantial right-side heart failure, evidence of mechanical complications or severe valvular disease. owing to progressive hypotension and an arterial lactate level of 10.2 mmol/l, norepinephrine infusion was initiated. the shock team was gathered for a discussion of the treatment strategy. a microaxial flow pump was decided on before the cag. while the device and femoral access were being prepared, the patient experienced a short episode of cardiac arrest with pulseless electrical activity and need for chest compression. return of spontaneous circulation was achieved. however, the patient subsequently experienced another episode of cardiac arrest, prompting the decision to intubate. during mechanical cardiopulmonary resuscitation, a microaxial flow pump (impella cp) was implanted through the right femoral artery (14-f sheath) started at a low ¿performance (p) level¿ and gradually increased to a flow of 3.5 l/min at p-level p8. hemodynamic stability was immediately achieved. during the pci, the microaxial flow pump triggered a ¿high purge pressure¿ alarm. the purge tubes were checked for kinks, which were not found, and it was decided to replace both the purge fluid and the purge cassette; the problem was thus resolved, and the patient was transferred to the cardiac intensive care unit (cicu) after the ppci. in the cicu, motor current increased on the device, leading to device malfunction with pump stop. bedside echocardiography was performed to ensure correct position of the microaxial flow pump and to rule out presence of thrombus formation, which was not present. several attempts to run the device were attempted, but it was only possible to run at p-level 2 with an inadequate flow and massive elevation of lactate up to 21 mmol/l. of note, there was elevation of lactate between pump 1 and pump 2. the impella pump 1 device was replaced by a new device via the same access site. in the following weeks, the patient¿s cardiac function increased to an ejection fraction of 30%. the patient was hemodynamically stabilized to the extent that a microaxial flow pump was no longer required. however, the patient experienced a prolonged complicated icu course. the patient's condition was complicated by the development of melena, atrial flutter, acute kidney failure, pneumonia, and ileus. the patient's condition continued to decline owing to multiorgan failure, and he died after 34 days. | On 2025-07-09 abiomed japan forwarded 7 publications to complaints team for review. one publication from copenhagen entitled "unexpected percutaneous ventricular assist device failure in a complex case of infarct-related cardiogenic shock" which details a cp support that was explanted after high purge pressure alarms, purge cassette replacement, and a pump stop occurred. pump was replaced. |  |  |  |  | The investigation of pump stop and high purge pressure has been completed. the impella device was not returned for evaluation. high purge pressure: clinical details report that a high purge pressure alarm triggered during a pci procedure. the purge line was checked for kinks, and none were noted. however, when the purge cassette and fluid were replaced, the issue resolved. neither product nor data logs were available for investigation. since neither product nor data logs were returned for investigation and clinical details provided are confounding, the cause of the high purge pressure could not be determined. temporary pump stop: clinical details report that shortly after a high purge pressure event, motor current increased, resulting in a pump stop. the pump was able to be restarted at p-2, but flows were inadequate, so the pump was replaced with a new one. neither product nor data logs were available for investigation. since neither product nor data logs were returned for investigation, the cause of the temporary pump stop could not be determined. medical safety review of the event was completed. the event describes a pump stop failure. a different impella pump was placed, same access used. the first pump had device malfunctions (pump stop and restart at p-level 2 with an inadequate flow with massive elevation of lactate; patient not well supported). additionally, the patient experienced the atrial flutter which could be caused from a number if things given the patient's condition including possibly impella pump 1; although unlikely. and the experienced acute kidney failure can be associated with pump 1 with the lactate going to 21 mmol/l. the patient stabilized but eventually expired (while on the replacement pump). in this case, patient underlying condition likely contributed most to an outcome of death (cardiogenic shock totally occluded vessels etc.). however, the impella pump 1 cannot be dissociated from the demise outcome given the pump stop and untoward impact to the patient's hemodynamic stability in this complex case." the article was erroneously omitted from the initial report. (attached now) b1 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30629. b2 death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30629. date of death is unknown b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30629. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2025-30629. g1 manufacturing site fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30629. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-30629. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30629. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30629. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. pump stop: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7181 | 1220648-2025-30628 | 22576651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22576651&pc=OZD | 07/09/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690971 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 67 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient developed hemolysis. the decision was made to wean the pump. during the weaning, the plasma free hemoglobin was trending down. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7182 | 1220648-2025-30630 | 22577055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577055&pc=OZD | 07/09/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682567 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had echocardiogram performed showing impella deep at 5.5cm, however, sitting mid ventricular space. the team elected to hold off repositioning until the evening hours. impella was pulled back to 3.6cm. shortly after repositioning, the patient was rolled and placed on a bedpan, and after removing, the patient went into sustained ventricular tachycardia (vt) and remained conscious the whole time. the team elected to reposition once again to 3.4cm middle of ventricular space followed by optimization of flows at 3.5 liter/min. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7183 | 1220648-2025-30631 | 22577186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577186&pc=OZD | 07/10/2025 | 11/05/2025 |  | 07/23/2025 |  |  | 07/10/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170094 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they received a controller error on their automated impella controller (aic) and lost its placement and signal while being used along with an impella cp pump for patient support. the pump continued to run and the signal reappeared after a short time. this issue occurred a second time and was also resolved after a short delay. there was no report of patient harm. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. cloud logs show that (b)(6) had a controller error alarm triggered during the clinical procedure. real time (rt) logs confirmed that all signals the issue. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d4 model number updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7184 | 1220648-2025-30632 | 22577274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577274&pc=OZD | 07/09/2025 | 09/11/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679813 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the pump was being prepped and the automated impella controller (aic) showed optical sensor error when pump was connected. perfusion disconnected and reconnected pump and error went away. the pump and aic were replaced. there was no patient harm. | [5 empty] | The investigation was completed. clinical details note that there was an optical sensor error upon initial connection of pump sn (b)(6) to (b)(6). once the pump was disconnected and reconnected, the error went away. data log review confirmed the system sensor error after plugging in the pump. the user was directed to re-plug in the pump to the console. the error resolved upon reconnection but the pump and console were both replaced as noted in clinical details. the raw optical was -3276.8 mmhg upon first connection with snr near zero, which likely indicates an air gap. furthermore, the pump was returned and evaluated. the 5s parameters were in spec and the pump passed the laser continuity test. the pump was successfully connected to the lab console without optical sensor errors. the cause of the optical signal issue was most likely an air gap from between the aic and the patient cable plug which was resolved after disconnecting and reconnecting the pump as noted in clinical details and seen in the data logs. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7185 | 1220648-2025-30634 | 22577518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577518&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  | 09/30/2025 |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411515 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient presented as a st-elevation myocardial infarction due to stent re-stenosis. it was noted that the patient had a successful balloon angioplasty of the coronary artery branch to saphenous vein graft. during the case, a placement signal not reliable alarm was noted. the procedure was completed and to provide the best intensive care, the impella cp was removed for a new device with a functioning placement signal. | [5 empty] | The investigation of optical signal issue has been completed. impella cp returned for evaluation. upon visual inspection, no abnormalities were observed. pump passed laser continuity testing and all optical 5s parameters were within spec. no data logs were returned for evaluation. the cause of the optical signal issue could not be definitively determined as no abnormalities were observed with returned product, and no logs and limited clinical details were provided. | The impella pump and purge cassette were device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7186 | 1220648-2025-30635 | 22577574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577574&pc=OZD | 07/09/2025 | 09/11/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1671526 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the pump was being prepped and the automated impella controller (aic) showed optical sensor error when pump was connected. perfusion disconnected and reconnected pump and error went away. the pump and aic were replaced. there was no patient harm. | [5 empty] | The investigation of optical signal issue and pump not detected has been completed. the impella device was not returned for evaluation. analysis: cloud logs confirmed the complaint. on 7/9, 5.5 pump 597025 was connected to (b)(6) but then the user received a system sensor error and asked to re-plug in the console. the error did not return but the pump was removed from (b)(6) anyway. the real time (rt) logs show that on the first attempt raw optical sensor was -3276.8mmhg as snr was very low signifying a likely air gap. root cause: the cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug that was resolved by unplugging and re-plugging in the pump as instructed by the console. there was no problem found during the case with pump detection. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7187 | 1220648-2025-30636 | 22577679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577679&pc=OZD | 07/08/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657718 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 68 KG | Non Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was running for 42 days when there was a leak observed. they troubleshot the purge leakage with dressing and tegaderm. the pump remains on for support while decisions/transplant was being discussed with the medical team. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
7188 | 1220648-2025-30637 | 22577821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577821&pc=OZD | 07/09/2025 | 09/04/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679710 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed hemolysis, noted with high lactate dehydrogenase and red urine. impella was repositioned and preload augmented. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and positioning issues has been completed. the pump was not returned for investigation. data log shows a motor current spike during the reported hemolysis on (b)(6) 2025, with no positioning-related issues or significant trends in placement or motor current. according to the clinical report, hemolysis resolved later during the case run. the pump was also repositioned at that time. the cause of the hemolysis issue was most likely biomaterial ingestion based on data log review. thrombus failure mode was added as an investigated failure mode based on data log review and hemolysis investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provide.
7189 | 1220648-2025-30638 | 22577857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22577857&pc=OZD | 07/08/2025 | 10/27/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1571192 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reportedwhen turning on the automated impella controller (aic), a "system check failed. replace console immediately." appeared. the console was replaced immediately. no patient was on the aic. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Data analysis: on the complaint date, (b)(6) 2025, the console was automatically rebooted due to the powermod_1 communication issue. during the reboot, the persistent powermod_1 communication issue caused the console to display the ¿system self check failed¿ screen, confirming the reported failure mode. device analysis: this console was tested after arriving at fs. visually confirmed corrosion on both the top and bottom sides of the pbm. corrective actions: events of this type will be monitored for trends. escalation criteria as per (b)(4) does not require reescalation. escalation (b)(4) had assessed this issue. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. d9 and h3 updated. h8 updated.
7190 | 1220648-2025-30639 | 22578199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22578199&pc=OZD | 05/13/2025 | 10/30/2025 |  | 07/23/2025 |  |  | 05/13/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654938 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 67 KG | Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial tachycardia. the patient was defibrillated and treated with lidocaine, magnesium, and amiodarone. the patient also exhibited dark urine and hemolysis was suspected, so the purge flow was doubled. the patient survived. | [5 empty] | The investigation of vf/vt/af/at and hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: the available data logs are for the period from (b)(6) 2025 to (b)(6) 2025. a clear relation between the reported hemolysis and it's occurrence due to the usage of the impella 5.5 s2 pump could not be established as there is not enough data. to determine the root cause of the hemolysis, the clinical information mentioned above would need to be assessed in conjunction with evidence from the data logs. the root cause of the hemolysis was not determined due to insufficient clinical details and logs. vt/ vf/ af/ at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-30639. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis". ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
7191 | 1220648-2025-30640 | 22578236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22578236&pc=OZD | 07/08/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654579 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 84 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had a three vessel coronary artery bypass graft earlier in the year. however, they did not follow up with cardiology or continue medications. they were presented to the hospital with shortness of breath, pulmonary congestion, and had a cardiac arrest on the floor and was taken to the operating room in cardiogenic shock. an impella 5.5 was placed. throughout impella 5.5 support, the patient experienced multiple episodes of ventricular tachycardia and was defibrillated several times. amiodarone was started and the pump was repositioned. however, the patient continued to have ventricular tachycardia runs requiring defibrillation. it was noted that the patient was not waking up and computed tomography was negative. the patient continued to be jaundice, ventilated and requiring pressors. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7192 | 1220648-2025-30642 | 22578669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22578669&pc=OZD | 01/11/2025 | 10/30/2025 |  | 07/23/2025 |  |  | 01/11/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594092 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple runs of supraventricular tachycardia during impella 5.5 support. an amio bolus and drip were given to treat the condition. it was also reported that the pump experienced several suction issues during support. three units of blood and 150ml of albumin were given to the patient to treat the issue. the pump was also repositioned to troubleshoot the suction. it was later reported from an impact study documented that the patient experienced hemolysis and a post op bleed diagnosed as acute blood loss anemia. the patient received multiple blood products to treat the condition. per the impact study, the conditions were believed to have a definite relationship to the impella device and procedure. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of vf/vt/af/at, vascular damage - peripheral vessel, and hemolysis has been completed. no product was returned for investigation. data logs were returned for analysis. data log analysis confirmed suction alarms outside of implant around (b)(6) 2025, which got resolved after dropping the p-level. there was no trending placement signal, and the motor current had no spikes or trends outside of p-level changes. no abnormalities outside of the suction alarms were observed. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: the root cause of the vascular damage-peripheral vessel was not determined due to lack of sufficient clinical details. hemolysis: the root cause of hemolysis was not determined due to insufficient clinical details, unreturned product, and inconclusive evidence from the data logs.
7193 | 1220648-2025-30641 | 22578725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22578725&pc=OZD | 07/08/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657739 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 60 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 pump was placed to support a patient who had previously been on impella cp support but had new post explant hypoxia and ventricular tachycardia. the impella 5.5 was placed and was supporting when there was a hematoma at the access site. the heparin was paused due to the hematoma. the pump remains on despite the harm. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7194 | 1220648-2025-30645 | 22579011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579011&pc=OZD | 07/09/2025 | 10/27/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686800 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a bipella cp-rp support of 5 days. the impella rp was explanted and the cp began to have suction and low flows. the team tried to troubleshoot but explanted the cp and found that the pump had ingested a clot. the patient remained stable. | [5 empty] | Thrombus: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. low pump flow: the root cause of low pump flow was most likely biomaterial ingestion since there was a visual confirmation of clot ingested into inlet of cp upon explant. there is a very small increase in purge pressure at the time of the event. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7195 | 1220648-2025-30643 | 22579043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579043&pc=OZD | 07/08/2025 | 09/10/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683910 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male | 100 KG |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella cp support and during support, the pump got caught in the trabecular meshwork and was then under tension. the pump was explanted. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use issue since the pump got caught in trabecular meshwork and was then under tension during the repositioning of the pump.
7196 | 1220648-2025-30644 | 22579052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579052&pc=OZD | 07/09/2025 | 09/05/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688244 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed hemolysis. the impella was too deep and was repositioned. the patient was on cp support for 10 hours and was escalated to impella 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues, hemolysis, and renal failure has been completed. the product and data were not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was most likely due to pump was not in the proper position since clinical team confirmed pump was deep in ventricle and hemolysis was managed after repositioning of the pump. the cause of the renal failure was not determined due to insufficient clinical details. h6 health effect - clinical code 4499 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30644.
7197 | 1220648-2025-30646 | 22579121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579121&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655889 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 68 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and had bleeding at the access site. manual pressure was held, a femostop was applied, and a hemostatic suture was placed. impella support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because hemostasis was achieved by manual pressure, femstop, and suture.
7198 | 1220648-2025-30647 | 22579237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579237&pc=OZD | 07/08/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677186 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 101 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 and experienced ventricular tachycardia requiring defibrillation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7199 | 1220648-2025-30648 | 22579262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579262&pc=OZD | 07/10/2025 | 09/05/2025 |  | 07/23/2025 |  |  | 07/10/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629971 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant was placing a 5.5 pump for the support of the patient admitted in with post cardiotomy cardiogenic shock/low cardiac output syndrome but, due to anatomic limitations the 5.5 failed to deliver and as such the team instead placed a cp. no harm or injury noted from attempt. the cp successfully supports the patient to date. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's small vascular anatomy.
7200 | 1220648-2025-30649 | 22579406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579406&pc=OZD | 07/09/2025 | 10/09/2025 |  | 07/23/2025 |  |  | 07/09/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510616 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the pump had high purge pressure. tissue plasma activator was added to the purge. staff monitored. additionally, the patient became heparin-induced thrombocytopenia positive. anticoagulation was switched to argatroban. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure and thrombocytopenia has been completed. the impella device was returned for evaluation. data logs were not available. product was returned and evaluated. no kinks were observed on the returned product. there was some biomaterial on the impeller and some around the purge gap. however, high purge pressure did not reproduce. it was unable to be determined whether the biomaterial found around the impeller blades and purge gap affected the purge pressure in any way and if it was ingestion, buildup, or pulled during pump explant as no data logs were returned to make the distinction. high purge pressure: the cause of the high purge pressure was not determined due to lack of sufficient clinical details, since issue did not reproduce with return product, and unreturned data logs for analysis. thrombocytopenia: the root cause of the thrombocytopenia was not determined due to insufficient clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added
7201 | 1220648-2025-30650 | 22579425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579425&pc=OZD | 07/08/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/08/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676347 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced ventricular tachycardia. the patient was defibrillated and cardiopulmonary resuscitation was performed to achieve return of spontaneous circulation. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7202 | 1220648-2025-30651 | 22579598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579598&pc=OZD | 07/10/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/10/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686376 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was no flow distal to common femoral artery. vascular surgery with right femoral cutdown, found to have femoral dissection flap with repair. flow was regained after the repair. additionally, the patient had bloody secretions. heparin was used in the purge prior. support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7203 | 1220648-2025-30652 | 22579656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579656&pc=OZD | 07/10/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/10/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686725 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the removal of the impella the closure of the femoral access and perclose placement was complicated. due to a malfunction with the perclose the vessel dissected. the team troubleshot by application of the femstop and vascular surgery was consulted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7204 | 1220648-2025-30653 | 22579883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579883&pc=OZD | 07/18/2025 | 09/24/2025 |  | 07/23/2025 |  |  | 07/18/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679054 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that roughly 9 days into patient support, an impella cp pump stopped unexpectedly. it was noted that the pump was being weaned at the time of the stoppage and that a high motor current alarm appeared shortly before the event. the pump was removed as a result of the failure. no patient harm was reported. | [5 empty] | The investigation of pump stop has been completed. insufficient data logs were returned. as per clinical, purge pressure rise from 400 to 600 mmhg observed in relevance to high motor current pump stop event on (b)(6) 2025. impella cp pump was returned and visual inspection revealed large purge gap at outflow. no open electrical lines were found during electrical resistance test. bearing wear issue was confirmed with return product finding. the cause of the pump stop was bearing wear which occurred beyond design life. the impella cp optical device is not indicated for support more than 8 days of design life period according to design trace matrix per instructions for use. | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7205 | 1220648-2025-30654 | 22579892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22579892&pc=OZD | 07/11/2025 | 07/23/2025 |  | 07/23/2025 |  |  | 07/11/2025 | 07/23/2025 |  |  |  |  | 07/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697821 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and post implant, same day, the patient experienced limb ischemia, access site bleeding and subsequently expired. the patient presented with dizziness and general malaise. electrocardiogram revealed st-segment elevation myocardial infarction which prompted emergent intervention. the impella cp device was implanted given the profound hypotension and started in auto. further work up revealed a "large" ventricular septal defect. the team hypothesized that the condition was a takotsubo cardiomyopathy ventricular septal defect. the patient then began reporting foot pain/numbness. there was no distal signal on doppler. a perfusion catheter was successfully placed. now bleeding at the access site was noted, and consequently, the sheath was repositioned, angle augmented, perclose sutures were retightened, and manual pressure was applied at the site. the patient would continue on support. however, two days later overnight, the patient went into cardiac arrest and expired. care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7206 | 1220648-2025-30555 | 22563616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563616&pc=OZD | 07/02/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/02/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613567 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male | 115 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the decision was made to place an impella 5.5 as a bridge to a heart transplant. during support, it was reported that the patient was taken back to operating room for a washout and exploration of right axillary site. ten days later, the patient experienced a brief run of ventricular tachycardia over the night which reportedly seems associated with when the patient sleeps on one side. the patient remained on support until a bridge to transplant. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7207 | 1220648-2025-30556 | 22563712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563712&pc=OZD | 07/06/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/06/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660842 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the pump was supporting with the perfusion of the limbs by a distal pefusion catheter (dpc) line. the team used the contralateral leg for the dpc. additionally, there was ventricular fibrillation which required shocking six times. the following day, a clot was observed in the left ventricle by echocardiogram and the team troubleshot with bivalirudin. the pump remained on despite these harms and was successfully weaned off after a few days. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7208 | 1220648-2025-30557 | 22563782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563782&pc=OZD | 07/06/2025 | 08/28/2025 |  | 07/22/2025 |  |  | 07/06/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED.INC |  |  | US |  | 2026693921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 51 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient developed a bleed and hematoma from the access site. the site was re-sutured, two units of packed red blood cells were given, and heparin was held. the bleed and hematoma resolved. heparin was to be restarted. impella support was continued for a few more days until the patient was successfully weaned and explanted. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since hemostasis was achieved by re-suturing. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30557. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7209 | 1220648-2025-30558 | 22563853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563853&pc=OZD | 07/06/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/06/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665239 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. upon explant of the impella cp, the perclose at the right femoral artery failed. the 14 french sheath was placed for temporary management until a femoral cutdown and surgical repair of the artery could be performed. the patient was deemed stable following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7210 | 1220648-2025-30559 | 22563854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563854&pc=OZD | 06/22/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/08/2025 | 07/22/2025 |  |  | 04/30/2021 |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1416741 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on cp support for high risk cardiac procedure. the cp was placed and allowed for the coronary artery bypass graft surgery (cabg) and was then explanted. there were no alleged issues during support. patient survived to explant of the pump per case documentation. however, a month post-implant, abiomed/jnj legal forwarded an allegation of the patient's death. the report/death certificate states the death the day after explant and cause of death as cardiogenic shock and severe coronary artery disease. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7211 | 1220648-2025-30560 | 22563928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563928&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/22/2025 |  |  | 07/08/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619274 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was on impella cp support and during support, there was a hematoma at the access site. the hematoma caused the device to explant across the aortic valve. the device was explanted. | [5 empty] | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding - major: the cause of the bleeding issue most likely was patient condition related due to coagulation disorder since the patient had disseminated intravascular coagulation (dic). positioning issues: the cause of the positioning issue was not determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7212 | 1220648-2025-30561 | 22563934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563934&pc=OZD | 07/14/2025 | 09/03/2025 |  | 07/22/2025 |  |  | 07/14/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1072987 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, there were three controller errors with impella defective alarms. the decision was made to continue support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) and pump not detected has been completed. data analysis: the logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. device analysis: console was sent back for further investigation but the failure mode was not reproduced. root cause: the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received. the aic sw operated as designed. pump not detected: there was no issue found during the case. the device functioned/operated to specification. d9 device returned to manufacturer date added g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30561. h6 component code 925 was erroneously entered on the initial manufacturer device report number 1220648-2025-30561.
7213 | 1220648-2025-30562 | 22563997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22563997&pc=OZD | 07/07/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665215 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 87 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient had a hematoma at the access site. heparin was withheld. the procedure was successful and the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7214 | 1220648-2025-30563 | 22564159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564159&pc=OZD | 07/07/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/08/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691765 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had a cp support for the patient admitted in with a st-segment elevation myocardial infarction. the pump was placed and was supporting for the percutaneous coronary intervention of the coronary artery disease that was multi vessel. the pump was explanted after 75 hours. on the day of explant the team stopped the heparin due to pulmonary hemorrhage, which was later found not to be a pulmonary hemorrhage but rather nasal bleed post nasogastric tube. at the time of explant, a thrombus was observed and medically managed in the iliofemoral access. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7215 | 1220648-2025-30564 | 22564218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564218&pc=OZD | 05/29/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/09/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that placement signal of an implanted impella cp was lost during patient support. support was continued with the implanted pump until planned escalation later that day. upon later review of the data logs by an abiomed engineer, it is believed that the automated impella controller may have caused or contributed to the noted failure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7216 | 1220648-2025-30565 | 22564222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564222&pc=OZD | 06/23/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 06/23/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653819 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced a stroke which prompted discontinuation of angiomax. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7217 | 1220648-2025-30566 | 22564259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564259&pc=OZD | 07/08/2025 | 09/24/2025 |  | 07/22/2025 |  |  | 07/09/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786118 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller (aic) had a controller failure. the aic was replaced and support was continued. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console and the data logs were received for investigation. aic - controller failure (red alarm): there was no occurrence of a controller failure alarm in the case associated with the reported complaint. the unexpected controller shutdown alarm was misreported as a controller failure alarm. aic - shutdown or restart issue: the root cause of the unexpected controller shutdown and the ossiopsens.core process crash was not determined due to the inability to reproduce them. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7218 | 1220648-2025-30567 | 22564276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564276&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/22/2025 |  |  | 06/25/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662875 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the implant was subsequently aborted as a result of the noted difficulty. the patient passed later that night. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the pump could not be inserted due to the patient anatomy and high amounts of inotropic and vasoactive drops per staff. b1 product problem inadvertently omitted on the initial manufacturer device report number 1220648-2025-30567. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7219 | 1220648-2025-30569 | 22564359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564359&pc=OZD | 06/27/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 06/27/2025 | 07/22/2025 |  |  | 01/31/2026 |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505510 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced acute renal failure. it is unknown if this was preexisting condition or if it was impella related. the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7220 | 1220648-2025-30568 | 22564360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564360&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669129 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed to support a high-risk percutaneous intervention (hrpci) procedure. during the hrpci, the placement signal was lost when the shockwave was performed. impella flows remained stable and the motor current was stable. the impella cp was noted to be in good position and the patient¿s hemodynamics remained stable. no patient harm was reported. | [5 empty] | The investigation of optical signal issue has been completed. the impella device was returned for evaluation. optical signal issue: pump optical parameters were measured and 5s parameters were consistent with a sensor face break. data logs were reviewed and the logs shows that a "placement signal not reliable" (psnr) alarm occurred with a sudden increase in placement signal to 3000. optical parameter changes occurring at time of psnr alarm are consistent with a sensor face break. clinical details noted that the placement signal was lost after shockwave was performed. per instructions for use indicating that the use of shockwave at a distance of less that 20mm from the optical sensor may interfere or damage the optical sensor. the distance between the shockwave device and the pump was not noted. the cause of the optical signal issue was due to a damaged sensor as a result of shockwave use based on clinical details provided and returned product and data log analysis consistent with damage to the optical sensor face during pump use. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7221 | 1220648-2025-30570 | 22564438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564438&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648609 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, a semi-firm softball size hematoma developed with some oozing around the repositioning sheath in the impella groin (right femoral artery). manual pressure was held and the sheath angle was adjusted and re-sutured to manage the oozing/hematoma. support was continued following the intervention and the patient is stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be most likely use issue because the bleeding was resolved by sheath angle adjustment, re-suturing and applying manual pressure at the site. | Section a, b.7, d.6a and d.6b, e.1 name prefix/title, first name and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7222 | 1220648-2025-30571 | 22564452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564452&pc=PYX | 06/28/2025 | 08/26/2025 |  | 07/22/2025 |  |  | 06/28/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026675696 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 95 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced cardiac ectopy. the patient developed a fever on support and antibiotics were administered. furthermore, the patient had blood pooling in their mouth from the endotracheal tube and bleeding at the access site. one unit of blood was transfused and a new mattress suture was placed. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of access site bleeding, vf/vt/af/at, infection/sepsis, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was most likely use issue since hemostasis was achieved by placing a new mattress suture. vascular damage peripheral vessel: the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. vt/vf: the root cause of the vt/vf was unable to be determined due to insufficient clinical details. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer: ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7223 | 1220648-2025-30572 | 22564583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564583&pc=OZD | 06/28/2025 | 08/26/2025 |  | 07/22/2025 |  |  | 06/28/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. hemostasis was achieved by suturing the site. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because bleeding was controlled and hemostasis was achieved by placing additional sutures at the bleeding site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7224 | 1220648-2025-30573 | 22564675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564675&pc=OZD | 07/01/2025 | 08/28/2025 |  | 07/22/2025 |  |  | 07/01/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661747 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female | 100 KG |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 could not be completed due to the patient's tortuous anatomy. the patient was implanted with an impella cp instead. | [5 empty] | The investigation of delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related due to severe tortuous anatomy and anomalous take off of innominate artery the 5.5 was aborted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7225 | 1220648-2025-30574 | 22564769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564769&pc=PYX | 07/05/2025 | 12/04/2025 |  | 07/22/2025 |  |  | 07/05/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1683279 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA RP FLEX SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the buttons on an automated impella controller (aic) became unresponsive and the user was unable to change screens or manage impella flows. the aic was replaced to resolve the issue. no patient harm was reported. | [5 empty] | Data analysis: on 7/3, 5.5 pump 600436 was connected to ic10058. the console performs as expected with user keypresses logged in the imclogs through 7/5. starting at 10:43, accompanying keypresses were imc-kbr errors which indicate that the keyboard senses an out of bounds press. this happens a two more times before the pump is disconnected from the console with the last set of kbd errors not having corresponding logged keypress. device analysis: console was run over many days and routinely tested for keyboard functionality which never faltered. the console performed as expected throughout testing. root cause: the cause of the keyboard interface issues could not be determined because the failure mode did not reproduce during testing. | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7226 | 1220648-2025-30575 | 22564887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564887&pc=OZD | 07/08/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/08/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638702 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient was on impella cp support and during support, the patient had decreased perfusion in the impella leg. the impella was also deemed to be underneath papillary muscle with inlet close to mitral valve. the decision was made to bring the patient to operating room to do a cutdown and repair on right groin and replace new impella in left groin. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7227 | 1220648-2025-30576 | 22564979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22564979&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536678 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was attempted to be implanted with an impella 5.5 for mechanical circulatory support. during the implant, the staff was unable to place the pump. the patient had a difficult anatomy and the pump could not be placed due to the angle. the procedure was aborted. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation and no issues were found on the returned product. the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's difficult vascular anatomy. | Section a.2, a.3, a.4, a.5, b.7, d.6a, d.6b, e.1 name prefix/title, first name and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7228 | 1220648-2025-30577 | 22565029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565029&pc=OZD | 07/08/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/08/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689537 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a cp pump placed for possible right and left sided support. the team had to place perfusion to the cp accessed limb, the delivery of the perfusion was difficult and the team had to enlist multiple myelodysplastic syndromes and vascular. the limb was perfused and the pump remains on for support despite the harm while neurological status and possible escalation to the 5.5 pump is discussed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7229 | 1220648-2025-30578 | 22565090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565090&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691848 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with high risk surgery was implanted with an impella 5.5 for mechanical circulatory support. during support, high purge pressure was noted. tissue plasminogen activator was discussed but, it is unknown if it was used in the purge. support was continued. | [5 empty] | The investigation of high purge pressure has been completed. the pump was not returned for investigation. data log shows a motor current spike followed by a small increase in purge pressure. approximately 36 hours into the case, purge pressure began to gradually increase when the p-level was reduced to p2. a high purge pressure alarm was observed after 20 hours of gradual rise but resolved after 6 hours with a 2-minute gradual decrease trend. the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review. thrombus failure mode was added as an investigated failure mode based on data log review and high purge pressure investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7230 | 1220648-2025-30579 | 22565192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565192&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/22/2025 |  |  | 07/08/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681761 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the pump was 2.8cm in the left ventricle. they decided not to adjust the position unless there was signs of hemolysis. later in the day, dark colored urine was reported. the doctor repositioned the pump to 3.8cm and the pump was turned down to p-8. r lactate dehydrogenase and haptoglobin was drawn to confirm hemolysis, they resulted 1184 and 21 respectfully. the urine cleared since repositioning. additionally, the patient had a hematoma and a femstop was placed. support continued. furthermore, the patient had ventricular tachycardia. the pump was turned down to p 2 and dobutamine was started. patient survived. | [5 empty] | The investigation of hemolysis, positioning issues, and vf/vt/af/at has been completed. the complaint device not returned for analyzing. data log reviewed: few suction alarms. quick resolved impella in aorta position alarms. low placement signal and unknown impella position alarms occurred. motor current (mc) spike observed at p5 which not impacting any issue. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. hemolysis: significant hemolysis which resolved with pump repositioning and adjusting p level, urine cleared. the cause of hemolysis issue most likely was improper positioning issue since it resolved after device repositioning and p-level adjusting. positioning issues: the cause of positioning issue cannot be determined due to insufficient clinical details. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7231 | 1220648-2025-30580 | 22565316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565316&pc=OZD | 07/07/2025 | 09/04/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671957 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.5 for mechanical circulatory support. during support, a bubble-like appearance on the transesophageal ehocardiogram was noted after the impella 5.5 was started. the left ventricle was under filled and the physician was requesting volume to be given. while running at p2, there appeared to be bubbles in the aortic root while the impella was running. a needle was inserted into the ascending aorta with only normal, not frothy, blood coming out. the physician did not feel that there was any significant amount of air in the aorta. the procedure was successfully completed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of embolism and hemolysis has been completed. embolism: the product was returned and evaluated. there were no visual abnormalities on the pump and while purging the pump, there was no sign of purge leaking or damage to the filter. data log review showed that the pump was successfully primed prior to insertion into the patient and purge pressure and purge flow were stable throughout the case. the cause of the embolism was not determined since insufficient clinical details were provided and the returned product showed no abnormalities. d9 device returned to manufacturer date added.
7232 | 1220648-2025-30581 | 22565360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565360&pc=OZD | 07/15/2025 | 09/11/2025 |  | 07/22/2025 |  |  | 07/15/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660280 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella cp implant and after the implant, bleeding near the insertion site was noticed. after a contrast cat scan, it was found that the bleeding point was not at the insertion site. the cause was that during wiring of the guidewire when inserting the 8fr sheath for the first insertion, the guidewire got into a small blood vessel on the inside of the thigh. when the wire position was adjusted, the blood vessel was damaged, and caused the bleeding from there. it was not bleeding from the insertion site of the impella. it was confirmed that this was not the guidewire from the abiomed kit. 12 units of red blood cells were given to the patient. | [5 empty] | The investigation of vascular damage peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of vascular damage peripheral vessel was most likely use issue since during wiring of the guidewire when inserting the 8fr sheath for the first insertion, the guidewire got into a small blood vessel on the inside of the thigh. when the wire position was adjusted, the blood vessel probably got damaged, and caused the bleeding from there. d6b explant date added. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7233 | 1220648-2025-30583 | 22565371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565371&pc=OZD | 07/16/2025 | 12/09/2025 |  | 07/22/2025 |  |  | 07/16/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026679198 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that blood was seen slowly leaking from the connection between the implanted impella 5.5 catheter plug and the connection cable during patient support. it was believed that the bleed was a result of damage on the catheter shaft that was inside the patient's body which allowed blood to flow backwards up to the connection. the blood leak stopped without the need for intervention and support continued uninterrupted. patient survived. | The following day, the automated impella controller was swap was performed. shortly following the change, an impella failure alarm occurred. the pump was reconnected but the issue persisted and the pump eventually stopped. the device was replaced as a result of the failure. |  |  |  |  | Clinical details note that on (b)(6) 2025, it was discovered that blood was slowly leaking from the connection between the pump plug and patient cable and catheter damage was suspected. on (b)(6) 2025, the aic was replaced after a "controller error" alarm, but an impella defective alarm occurred. it could not be resolved so the pump was replaced. imc log review confirmed the impella defective alarm. product was returned and evaluated. the impella defective alarm was reproduced when connecting the pump to a lab console. a hole was found in the catheter near the motor housing and blood was observed inside of the red pump plug. the cause of the pump detection issues was blood ingress into the red pump plug due to a hole in the catheter. the cause of the hole in catheter was most likely use-related as a hole was found in the returned pump just above the motor housing and clinical details noted a direct insertion approach with some resistance when passing through the graft anastomosis. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | D9 device returned to manufacturer date added. the impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30583. h6 health effect - impact code 4627 and medical device problem code 1503 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30583. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30583.
7234 | 1220648-2025-30582 | 22565376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565376&pc=OZD | 07/06/2025 | 09/02/2025 |  | 07/22/2025 |  |  | 07/06/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662753 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ¿impella position in aorta¿ alarms. the motor current was flat though an echocardiogram confirmed appropriate positioning of the pump. the pump was explanted on suspicion of a clot in the pump however, no clot was found. | [5 empty] | The investigation of positioning issues has been completed. positioning issues: there was no true positioning issue found during the case per clinical and log analysis. the device functioned/operated to specification. optical signal issue: the return pump was inspected for 5s check and all parameters were in normal range. light continuity test passed. biomaterial wrapped under impeller and small piece biomaterial found on optical sensor/viser line face and no other abnormalities were found. impella in aorta alarms seen post motor current (mc) spike at p-4. position was confirmed to in right place per echo and false alarms were observed. placement signal (ps) monitoring disabled. during false alarms period, all 5s were within normal range and no significant variations and motor current was almost flat after mc spike. pump flows saturated after mc spike. this failure trend indicates biomaterial ingestion falsely triggering position alarms causing several impella in aorta alarms to be triggered. clinical stated, ¿impella position in aorta¿ alarmed, the bedside rn reportedly said the motor current was flat with echo confirming appropriate positioning per the physician. upon rounding early the next morning, team decided there was a probable clot in impella and wanted to remove impella. the cause of the (optical) false impella position alarms are likely due to biomaterial ingestion based on log, return product and clinical analysis. failure mode thrombus added as a result of biomaterial ingestion observation. thrombus: returned 5.5 pump visually inspected. biomaterial observed wrapped around impeller. no broken blade or cannula kink or other abnormalities observed. the root cause of the thrombus was unable to be determined due to insufficient clinical details. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-30582. | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7235 | 1220648-2025-30584 | 22565584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565584&pc=OZD | 07/05/2025 | 08/27/2025 |  | 07/22/2025 |  |  | 07/05/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695480 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. heparin was withheld. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely anticoagulation management since hemostasis was achieved by stopping the heparin along the angle matching and redressing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7236 | 1220648-2025-30585 | 22565702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565702&pc=OZD | 07/05/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/05/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650888 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed a clot on the impella. the pump was explanted and the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7237 | 1220648-2025-30586 | 22565706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565706&pc=PYX | 07/10/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/10/2025 | 07/22/2025 |  |  | 02/28/2026 |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478781 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that resistance was met when attempting to implant an impella rp flex. the guidewire was pulled back to make a better rail, with no advancement of the pump. with multiple manipulations, the catheter became bent and could not be advanced. a new impella rp flex was used for patient support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
7238 | 1220648-2025-30587 | 22565927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22565927&pc=OZD | 07/10/2025 | 10/01/2025 |  | 07/22/2025 |  |  | 07/10/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Mechanical Jam | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a console failure in the cardiac catheterization lab. the door release button sticks due to dextrose. the console was removed from service and abiomed field service was notified. no patient was involved in this event. | [5 empty] | The investigation for aic - purge door won't close has been completed. console logs did not contain any data relevant to the complaint. console was tested after arriving in field service. the reported release door issue was reproduced during testing on an engineering lab bench. after removing and reseating the door release button, the sticking issue was resolved. dextrose residue was also observed around the purge assembly during the inspection of the console. the root cause of the purge door release button issue was likely contamination. before sending this console back to the customer, field service was instructed to clean the area around the purge door release button and the purge assembly, and perform a full functional check on the console and optical system. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7239 | 1220648-2025-30588 | 22566115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566115&pc=OZD | 07/09/2025 | 08/26/2025 |  | 07/22/2025 |  |  | 07/09/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645661 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female | 91 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the pump failed delivery due to the patient's tortuous anatomy. the case was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7240 | 1220648-2025-30589 | 22566333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566333&pc=OZD | 07/09/2025 | 09/04/2025 |  | 07/22/2025 |  |  | 07/09/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690985 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, there was an "impella in aorta" alarm. the doctor attempted to reposition to no avail. the pump was subsequently explanted. | [5 empty] | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was out of position while the patient ambulating. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7241 | 1220648-2025-30590 | 22566393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566393&pc=OZD | 07/05/2025 | 09/02/2025 |  | 07/22/2025 |  |  | 07/05/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681817 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 79 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Vascular Dissection | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted as it was identified the patient had a massive aortic dissection. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to resistance in aorta and the insertion was aborted. vascular damage (great vessel): vascular damage failure mode was added due to noted massive dissection of aorta. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. b1: adverse event was inadvertently omitted on the initial manufacturer device report number: 1220648-2025-30590. b2: other serious injury was inadvertently omitted on the initial manufacturer device report number: 1220648-2025-30590. h1: type of event was erroneously selected on the initial manufacturer device report number: 1220648-2025-30590. h6: health effect - clinical code 4582 and health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number: 1220648-2025-30590. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7242 | 1220648-2025-30591 | 22566484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566484&pc=OZD | 07/09/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/09/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689378 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 116 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient developed limb ischemia. the patient was taken to the operating room for escalation to an impella 5.5 for support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7243 | 1220648-2025-30592 | 22566567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566567&pc=OZD | 07/10/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/10/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026690255 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced episodes of ventricular tachycardia and ventricular fibrillation requiring defibrillation. the patient's family decided to withdraw care and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
7244 | 1220648-2025-30593 | 22566756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566756&pc=OZD | 07/02/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/02/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623541 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump placed to support a patient with escalation from an impella cp and underlying cardiac and systemic comorbidities inclusive of renal issues on dialysis, metabolic acidosis, sepsis, and liver disease. the pump was supporting when there was a diagnosis made of a large posterior cerebral artery territory infarct. the patient had survived til the family made the decision to withdraw care. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7245 | 1220648-2025-30594 | 22566826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566826&pc=OZD | 07/10/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/10/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681725 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 78 KG | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced tachycardia. amiodarone was administered. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
7246 | 1220648-2025-30595 | 22566988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22566988&pc=OZD | 07/05/2025 | 09/12/2025 |  | 07/22/2025 |  |  | 07/05/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1879998 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella cp for mechanical circulatory support. the complainant reported that an automated impella controller displayed controller error alarms coinciding with placement signal not reliable alarms. the audio alert for the placement signal not reliable alarm was disabled. despite the alarms, mechanical circulatory support continued without interruption. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the data logs were reviewed. automated impella controller (aic) - controller error (yellow alarm): the root of the controller error (event 1045) was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. optical signal issue: failure mode added based on investigation findings. the root cause of the placement signal not reliable alarms was not determined since aic was not returned. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7247 | 1220648-2025-30596 | 22567028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22567028&pc=OZD | 07/10/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/10/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686640 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 100 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed a hematoma at the access site. a surgical washout was performed. the infusion of heparin was replaced with bivalv. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7248 | 1220648-2025-30597 | 22567133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22567133&pc=OZD | 07/07/2025 | 07/22/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676252 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 75 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced ventricular tachycardia. treatment was not reported. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7249 | 1220648-2025-30598 | 22568412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22568412&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the waveform of an implanted impella cp became uncoupled due to the pigtail getting caught in the mitral valve. initial attempts to free the pigtail were unsuccessful, but the pump was eventually advanced and the waveforms stabilized. two days later, the left femoral impella side developed sanguineous drainage after the patient moved. the dressing was changed, heparin was withheld, and one unit of packed red blood cells was transfused to treat the bleed. the patient was transferred to comfort care and two days later passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issue most likely was use related since the bedside echocardiogram showed impella appearing shallow with pigtail in mitral valve (mv), unable to free pigtail from mv and advancing the position 1.5 cm later resolved issue. access site bleeding - major: the cause of bleeding issue most likely was due to patient movement since the patient had become agitated/confused and was moving affected extremity before leg immobilizer was placed. b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30598. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7250 | 1220648-2025-30599 | 22568663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22568663&pc=PYX | 07/07/2025 | 09/04/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025507437 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female | 133 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella rp flex and had a high purge pressure alarm. tissue plasminogen activator was initiated. additionally, angiomax was given to treat the patient for heparin induced thrombocytopenia. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of high purge pressure and thrombocytopenia has been completed. the impella device was not returned for evaluation. high purge pressure (hpp): data log review showed a gradual rise in pp up until a cassette change and then a rise from around 600 to greater than 1000 mmhg pp within another hour. motor current was gradually trending downward during this time but there were no motor current spikes prior to the rise in pp. the hpp was sustained for 7 hours before a small decrease from 1100 to 1000 mmhg over 30 minutes. it was noted that lysis protocol was initiated during the sustained hpp. the p-level was then turned down from p-6 to p-5 and a motor current spike occurred seconds later followed by a quick 4 minute drop from 1000 to 600 mmhg in 4 minutes. the cause of the high purge pressure was most likely biomaterial buildup since the purge pressure gradually trended upward over 2 hours and there were no motor current spikes. thrombocytopenia: the patient was noted to be hit positive and started on angiomax on (b)(6) 2025. no other details known. no product was returned and data logs are not applicable. the cause of the thrombocytopenia was not determined due to insufficient clinical details.
7251 | 1220648-2025-30600 | 22568848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22568848&pc=OZD | 07/07/2025 | 09/04/2025 |  | 07/22/2025 |  |  | 07/07/2025 | 07/22/2025 |  |  |  |  | 07/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689362 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Cardiac Perforation | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the a patient developed hemolysis during impella 5.5 support. the condition was monitored as labs continued to trend downward during ecpella support. eight days into support, the pump was pulled out of the ventricle to facilitate a planned ventricular septal defect patch. upon reinsertion, the pump perforated the left ventricle. the pump was removed and the left ventricle was repaired. | [5 empty] | The investigation of ventricle perforation and hemolysis has been completed. the impella 5.5 was not returned for investigation. data log shows no abnormalities related to motor current spikes or positioning issues during reported hemolysis. the patient was on prior support with the cp pump and was escalated to the 5.5 pump on 6-july-2025. clinical reports also suggest that hemolysis was improving on the 5.5 pump support. the cause of the hemolysis issue was not determined since the product was not returned and sufficient clinical information. the cause of the ventricular perforation issue was most likely use-related, as it occurred accidentally during placement after vascular repair in a patient with a small and diseased heart. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7252 | 1220648-2025-30517 | 22553343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553343&pc=OZD | 06/28/2025 | 09/15/2025 |  | 07/21/2025 |  |  | 06/30/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566959 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced reduced flow of the purge solution due to a thrombus in the pump. the p-3 flows were 3.3 liters per minute with a motor current of 500. saturated pump flow was noted. the low flow in the impella was due to the use of extracorporeal membrane oxygenation. the pump was explanted and the patient survived. additional information was received. a thrombus or clot in the cannula after the pump had been explanted was noticed. to manage the problem with the purge flow, p-level was initially tried to be ramped up a little bit to p-4 or p-5 and then sodium bicarbonate was replaced with heparin with the aim of having better anticoagulation. | [5 empty] | The investigation of thrombus, high purge pressure, and saturated flow has been completed. one the same day as implant, reduced purge flow and pump flows at p-3 of 3.3l/min noted with thrombus in catheter. data logs were not recovered. the pump was returned and inspected. biomaterial was on the impellor blades and a thin film of biomaterial was found in the purge gap. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. high purge pressure: the cause of the high purge pressure was biomaterial due to the biomaterial film found in the purge gap. saturated flow: the cause of the saturated flow was most likely biomaterial found on the impellor blades. d.9. device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7253 | 1220648-2025-30518 | 22553569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553569&pc=OZD | 07/07/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/07/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during priming of an impella cp pump, it was noted that the placement signal was giving faulty readings. the staff proceeded with the implant despite the noted reading, however the impella was unable to cross the aortic valve. the placement signal continued to be irregular and the decision was made to replace the pump. new impella cp then inserted without issue and good placement signals. no patient harm was reported due to issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7254 | 1220648-2025-30519 | 22553578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553578&pc=OZD | 07/03/2025 | 09/24/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672307 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump stopped unexpectedly during patient support. the staff was hesitant on explanting the device and the pump was restarted successfully. despite the motor current normalizing over the next three hours, the decision was ultimately made to withdraw care before another stop could potentially occur. a clot was not confirmed at the time of the incident. the patients outcome is stable. | [5 empty] | The investigation of pump stop has been completed. on day 8 of support, cp experienced a temporary pump stop but was able to be restarted. the occurred after the 192 hour cp design life. 6 hours after that, pump stop occurs that cannot be restarted. the pump was explanted. data logs were not recovered. the pump was returned and inspected. inspection revealed a large impellor purge gap on the returned product, indicative of bearing wear. the cause of the pump stop was bearing wear due to the large impellor purge gap found on the returned device. clinical details state this occurred after the 192 hour design life of the cp pump, making this misuse. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7255 | 1220648-2025-30520 | 22553604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553604&pc=OZD | 07/07/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/07/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680697 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon explantation of an impella cp closure of the access site was unsuccessful. vascular surgery was required to repair the vessel. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
7256 | 1220648-2025-30522 | 22553655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553655&pc=OZD | 07/08/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026649897 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 had a placement signal not reliable alarm. correct pump positioning was verified. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7257 | 1220648-2025-30523 | 22553686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553686&pc=OZD | 07/03/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661830 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp pump was placed to support a younger patient with seizure history found down for an unknown time. there was cardiogenic shock complicating the acute myocardial infarction diagnosis made and the pump was placed after the team had placed a reperfusion sheath to the leg. the limb was perfused but after four days of the impella cp, the pump was explanted and replaced by escalation to an impella 5.5 with concerns of limb ischemia. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7258 | 1220648-2025-30521 | 22553715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553715&pc=OZD | 07/03/2025 | 08/26/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651581 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during in-service training there was an unexpected controller shutdown. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - a/c power issues has been completed. console logs from the reported day of event are consistent with complaint and show that after demo pump was disconnected from the console and console shutdown was initiated, the shutdown procedure was incomplete, and based on some missing log entries, certain steps were not executed. last log entry is at 11:25:39. upon the next boot up, which started at 11:55:48, console presented with unexpected controller shutdown alarm (#603). console was idle for about 30 minutes, and when shutdown was initiated, the shutdown script was fully executed and console was able to correctly shut down. further analysis of console logs indicates that shutdown script imcshdn_bg appeared to have stopped executing somewhere between steps 4 and 20. because of that, the impellaconnect_bridge process kept running, which is evidenced by corresponding rlm journals: rlm keeps waiting for usb data from the console, and eventually restarting due to perceived loss of communication. as the result of incompletely executed imcshdn_bg script, its step #23 (¿touch /userflash/shutdown¿) had not been executed, which had consequence during the next boot-up, where imcstart script did not detect the shutdown file, which then triggered ¿unexpected controller shutdown¿ alarm (#603). console was tested, but issue was not reproduced. power-cycling the device 100 times and did not reveal any anomalies, and each time the imcshdn_bg script was executed correctly. there were no issues with power delivery to the console. the device functioned/operated to specification. the unexpected controller shutdown alarm (#603) was caused by prior inability to shut down due to incompletely executed shutdown process/script within aic software, after console shutdown was initiated by operator. issue could not be reproduced during testing, and root cause was unable to be determined. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7259 | 1220648-2025-30524 | 22553769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553769&pc=OZD | 07/08/2025 | 10/27/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678023 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The impella 5.5 was noted to be in proper position providing adequate support for the patient. there were plans to leave the patient on support until a transplant was available; however. the patient expired while on support. | It was reported that an impella 5.5 had been implanted in a patient when the placement sensor stopped functioning. the pump was explanted and a new pump was implanted. the patient survived. |  |  |  |  | The investigation for optical signal issue has been completed. the pump was not returned for investigation. data log shows all 5s optical parameters were within the normal range before the pump was disconnected from the console. the pump was reconnected to the same console with active psnr and air gap trends on optical signal parameters. the cause of the optical signal failure was most likely an air gap, based on data log review; however, the cause of the air gap was unknown without returned product investigation or sufficient clinical details. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. additional information was received: b1 adverse event updated, b2 updated to death, b5 expired while on support, h1 updated to death, h6 updated to include death b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | D6b "if explanted, give date" corrected. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7260 | 1220648-2025-30525 | 22553776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553776&pc=OZD | 07/03/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690212 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 84 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intracranial Hemorrhage | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient with anterior st-elevation myocardial infarction was found to be in acute myocardial infarction/cardiogenic shock. the decision was made to place an impella 5.5. during placement, before the graft could be anastomosed, the patient decompensated. the patient was cannulated for peripheral extracorporeal membrane oxygenation (ecmo) when the right ventricle free wall and interventricular septum ruptured simultaneously. the patient was crashed onto a peripheral venoarterial ecmo and an extensive repair surgery was completed. ecmo was reconfigured to central ecmo with left ventricle (lv) vent extending from the right upper pulmonary vein into the lv through the mitral valve. two days post procedure, the decision was made to bring the patient back to hybrid operating room for impella 5.5 placement. the impella 5.5 was successfully placed. during support, it was reported that the patient had computed tomography images taken and there were three areas of infarct and a brain bleed noted. ultimately care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7261 | 1220648-2025-30526 | 22553815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553815&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691038 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female | 48 KG | Non Hispanic | White | Injury | Yes | No | Difficult or Delayed Positioning | Malposition of Device | [5 empty] | Arrhythmia | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implantation of an impella 5.5 the physician had difficulty crossing the aortic valve. a straight 0.035 wire and jr4 catheter was used to successfully cross into the left ventricle. a transesophageal echocardiogram showed the pump position was still above the aortic valve, and a small pericardial effusion was noted. impella insertion continued and the patient became hemodynamically unstable. the patient was transfused five units of packed red blood cells, two units of platelets. kcentra and protamine were given. a hole was found on the lateral wall of the left ventricle. hemostatic patches and packing were applied to site, and two chest tubes were placed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of ventricle perforation and vf/vt/af/at has been completed. the impella device was not returned for evaluation. ventricle perforation: during implant procedure while pump placement was occurring, pericardial effusion was noted. investigation pointed to a "wire- sized hole" found on lateral left ventricle (lv) wall. the pump was not returned. the cause of the ventricle perforation was most likely a guidewire puncture due to the damage occurring prior to pump placement in the lv and physician's details provided. vf/vt/af/at: the root cause of the hypotension was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7262 | 1220648-2025-30527 | 22553903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553903&pc=OZD | 07/03/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660057 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient presented in cardiogenic shock and on extracorporeal membrane oxygenation (ecmo). the decision was made to escalate care to an impella 5.5 as a bridge until further continuation of care could be discussed and to decannulate ecmo. the implantation of the impella 5.5 and decannulation of ecmo was successful and the patient was sent to the unit to rest and recover. during support it was reported that the impella 5.5 was exchanged due to mispositioning. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7263 | 1220648-2025-30528 | 22553965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22553965&pc=OZD | 07/04/2025 | 09/22/2025 |  | 07/21/2025 |  |  | 07/04/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666479 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced heparin-induced hemorrhaging during impella cp support. the general hemorrhage included bleeding from the access site. it was discovered that the bleed was due to an excessive heparin dosage. upon returning the patient to the target partial thromboplastin time range, the bleeds stopped and support continued. after nine days of support, the motor current suddenly rose above 1050ma and the pump stopped. attempts to restart the pump were unsuccessful and the decision was made to explant the device. the lack of pump support was compensated for by increased catecholamine requirements. | It was reported that the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of access site bleeding, vascular damage - peripheral vessel, and pump stop has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of the access site bleeding was anticoagulation management related due to use of excessive heparin that resolved with likely heparin reduction. vascular damage - peripheral vessel: the cause of the vascular damage - peripheral vessel was anticoagulation management related due to use of excessive heparin that resolved with likely heparin reduction. pump stop: after 9 days on support, pump stop occurred after high motor current noted. the pump was not returned. data log review showed pump running for more than the 192 hours of cp design life. logs show motor current spike occur at pump stop with purge pressure increasing and purge flow decreasing; however, the purge metrics remain within operating parameters. the cause of the pump stop was not determined since product was not returned and insufficient clinical details. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30528. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30528. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-30528. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30528.
7264 | 1220648-2025-30529 | 22554007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554007&pc=OZD | 07/05/2025 | 09/02/2025 |  | 07/21/2025 |  |  | 07/05/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690188 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient developed a hematoma at the access site. manual compression was applied, and heparin was withheld. the patient continued on impella support. | [5 empty] | The investigation of access site bleeding has been completed. access site bleeding: clinical reported that when the patient got out of bed, a hematoma was noted at the axillary site. product was returned for investigation, and no abnormalities were found. the cause of the access site bleeding was most likely use related since the bleeding occurred when the patient got out of bed. | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7265 | 1220648-2025-30530 | 22554012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554012&pc=OZD | 07/03/2025 | 08/30/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651500 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 93 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed with best practices. the physician felt like the impella cp was underneath the papillary muscle and wanted to reposition. while attempting to reposition, the impella cp was pulled back into the aorta and was unable to recross. the physician was adviced to use a new impella but decided to use the same impella cp for placement again. the impella cp was rinsed in heparinized saline. the pump was examined for a clot or defect and was reinserted with no issues. impella cp turned on and flows of 3.5 were achieved. on day four of impella cp support, an echocardiogram showed a thrombus in the left ventricle. the patient was not a candidate for advanced therapy and care was withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the thrombus and positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump pulled out of the lv while repositioning of the pump. without sufficient clinical details, the root cause of the thrombus was unable to be determined.
7266 | 1220648-2025-30531 | 22554072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554072&pc=OZD | 07/03/2025 | 08/30/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026695517 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 111 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp pump placed to support for acute myocardial infarction/cardiogenic shock. the pump was persistently out of appropriate position within the left ventricle. the pump was adjusted but was along the mitral valve and also went to the papillary muscle. the pump remains on despite the malposition while awaiting possible impella 5.5 and extracorporeal membrane oxygenation placement. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue has been completed. pump was not returned for investigation. data log review was not required for this case run. the cause of the positioning issue was most likely use related since it was found deep during initial placement, based on given clinical details.
7267 | 1220648-2025-30532 | 22554152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554152&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677009 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced an ischemic stroke. the clot was noted to be too deep for a thrombectomy. anticoagulation was used to dissipate the clot, and then the patient became unresponsive. an echocardiogram showed a potential foreign body on the impella cannula above the inlet cage. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | Stroke/cva: clinical reports patient had acute ischemic stroke while on support confirmed by ct. the product was returned for investigation and no abnormalities were observed. a small kink on the cannula was discovered, but it is not related to this failure mode. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. thrombus: clinical reports echo showed a potential foreign body above inlet cage on the cannula; anticoagulation was increased and the clot dissipated. the product was returned for investigation and no foreign bodies were observed on the pump. a small kink on the cannula was discovered, but it is not related to this failure mode. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7268 | 1220648-2025-30533 | 22554235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554235&pc=OZD | 07/08/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665842 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited signs of hemolysis and hematuria. the pump was repositioned and the pump flow was reduced the labs continue to be hemolyzed. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
7269 | 1220648-2025-30534 | 22554243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554243&pc=OZD | 07/03/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586181 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that following impella cp implant, a planned revascularization attempt was performed during which the patient suffered from ventricular fibrillation. the patient was shocked 23 times in response to the condition. it was additionally reported that the patient endured a persistent gastrointestinal bleed throughout support. two units of packed red blood cells were given to treat the bleed and support was continued uninterrupted. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7270 | 1220648-2025-30535 | 22554447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554447&pc=OZD | 07/03/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685600 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 105 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during the implant, the impella cp was advanced and the patient went into ventricular fibrillation. cardiopulmonary resuscitation was imitated. the patient received 10 defibrillations, 450 milligrams of amiodarone, and 150 milligrams of lidocaine. support was continued. the patient had a small hematoma at the access site. manual pressure was held for 10 minutes with resolution of the hematoma. support was continued and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7271 | MW5172995 | 22554479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554479&pc=OZD | 10/15/2023 | 07/16/2025 |  | 07/21/2025 |  |  |  | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | P9 IMPELLA SWAN CL |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Patient | [6 empty] | 76 YR |  | Male | 84 KG | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Cardiogenic Shock | Obstruction/Occlusion | Pallor | Respiratory Failure | [8 empty] | Death | [5 empty] | ALOGLIPTIN, | ALTORVASTATIN | EZETIMIBE | GLIMEPIRIDE | ISOSORBIDE MONONITRATE | MEPOTROLO SUCCINATE | MEPOTROLO SUCCINATE, ANDISOSORBIDE MONONITRATE | METFORMIN | [15 empty] | Initial submission | [9 empty] | On (b)(6) 2023, 76-year-old (b)(6) walked into (b)(6) hospital, located at (b)(6) complaining of shortness of breath. within minutes he was diagnosed with a congestive heart failure exacerbation and was scheduled for a catheterization the next day due to fluid surrounding heart and lungs. he was given the drug lasix to help. it was delayed another day due to waiting for fluid to drain. dr. (b)(6) performed the catheterization around 11am and came to the conclusion there were no blockages, and it was not a "heart problem." his hospitalization then became complicated by cardiogenic shock and transferred to intensive care unit. on (b)(6), nephrology was consulted, diuretics were stopped, and they are started crrt and pressor therapy. this dialysis was supposed to be temporary. he had sustained v tach on (b)(6) requiring brief cpr, subsequent defibrillation, then intubation; also temporary. he went in for another catheterization with dr. (b)(6), who found five blockages and placed five stents. he then placed the impella supported pci. on (b)(6) his coloring was good and according to the nurse on duty, he was trending positive. when i arrived on (b)(6) his coloring was gray, his blood was purple/brown, and i had noticed that his bank of monitors had been changed and were plugged in on the other side of the room. soon after, the doctors were leading me by the arm to a quiet place to tell me i needed to make a decision of whether or not to take him off of life support, as he was getting worse, and there was no hope for him. after consulting with family, the decision was made to take him off of life support. he died within minutes on (b)(6) 2023.
7272 | 1220648-2025-30536 | 22554693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554693&pc=OZD | 07/03/2025 | 09/02/2025 |  | 07/21/2025 |  |  | 07/03/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622736 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 84 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was successfully placed. during the removal of the peel away and inserted repo sheath, the impella cp was pulled back into the aorta and could not recross the valve. the impella side wire port was used to introduce a .035 buddy wire to attempt to cross the valve, but it was unsuccessful. the impella cp was removed but wire access remained. a 25 centimeter peel away was placed. the impella was cleaned, flushed, and inserted with successful placement. the decision was made to leave the peel away in and send to the intensive care unit (icu). it was further reported that a femoro-femoral bypass was placed for acute loss of distal pulse and pale foot. once initiated, color returned and extremity was warm, pink, and perfused. the next day, ischemia was noted, there was an attempt to troubleshoot with distal reperfusion catheter, somewhat successful but ultimately the left lower extremity (lle) at risk. the impella was removed due to leg ischemia. impella closure was successful with contralateral angiogram used to verify blood flow down the lle. the patient was sent to the icu after explant and decompensated requiring increasing pressors. lle became cold and dusky again with increased pressor requirements and the patient was transferred to another facility for further management. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the limb ischemia and the positioning issue has been completed. the device was not returned for evaluation. clinical details state that the impella was removed due to leg ischemia. attempted to troubleshoot with distal reperfusion catheter, somewhat successful but ultimately left lower extremity was at risk. however, without additional clinical details, the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. clinical details also state that during the removal of the peel away and insertion of repo sheath the impella was pulled back into the aorta and could not recross valve. impella side wire port was used to introduce .035 buddy wire to attempt to cross valve but it was unsuccessful. impella removed but wire access remained. 25cm peel away placed and impella was cleaned and flushed and inserted with successful placement. the cause of positioning issue most likely was use related due to improper technique.
7273 | MW5173001 | 22554959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554959&pc=OZD | 07/16/2025 | 07/16/2025 |  | 07/21/2025 |  |  |  | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP AND IMPELLA AUTOMATED IMPELLA CONSOLE (AIC) |  |  | OZD | [5 empty] | JOHNSON & JOHNSON / ABIOMED, INC. |  |  | US | [5 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Nurse | [6 empty] | 53 YR |  | Male | 74 KG |  | Asian | Injury | Yes | Yes | Failure to Power Up | Therapeutic or Diagnostic Output Failure | [5 empty] | Insufficient Information | [12 empty] | Other | Life Threatening | Required Intervention | [26 empty] | Initial submission | [9 empty] | Patient (pt) was transferred from another facility (b)(6) with a johnson & johnson impella cp in place for higher level of care. the patient is transported on the transferring facilities console. it is standard policy to transfer the patient from the transferring centers console to our console upon arrival. this was attempted and the pump did not restart like it should have. the nursing staff tried 3 of our consoles and tried to put the pt back onto the other centers console and was never able to get the impella cp restarted. this could have been fatal had the patient not also been on ecmo. the impella catheter had to be removed and the patient taken to surgery urgently for a second impella implant. important to note that we had just received an urgent medical device notification (umdn) from johnson & johnson about this issue and currently do not have a fix for the issue. i will send this along with this submission.
7274 | 1220648-2025-30537 | 22554962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22554962&pc=OZD | 07/05/2025 | 09/18/2025 |  | 07/21/2025 |  |  | 07/05/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682276 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male | 100 KG |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. after repositioning, the automated impella controller (aic) displayed an ¿impella in aorta¿ alarm. attempts to reposition the device were unsuccessful, and the impella pump was subsequently explanted. care was withdrawn and the patient expired. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely patient movement per clinical that the patient was moved for an x-ray and then the pump could not be repositioned. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30537. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7275 | 1220648-2025-30538 | 22555138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555138&pc=OZD | 07/05/2025 | 09/02/2025 |  | 07/21/2025 |  |  | 07/05/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679186 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the pump developed elevated purge pressure and a purge system blockage alarm occurred. they replaced the purge cassette in an attempt to resolve the issue but ultimately tissue plasminogen activator was administered again via the purge system. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was returned for investigation. data log shows a gradual rise in purge pressure over three hours, followed by a high-pressure alarm. some resolution was observed in the purge pressure with a gradual decrease trend for around four hours, but the trend then began gradually increasing and remained high for the rest of the case run, approximately seven hours. no motor current spike was reported during high purge pressure event. the returned product did not have any physical abnormalities. the purge gap was inspected, and some dry dextrose was found. the pump was started priming, and the pressure was high at the beginning but returned to within specification range after the pump was started in a bucket of water. the cause of the high purge pressure issue was biomaterial buildup during case run.
7276 | 1220648-2025-30539 | 22555574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555574&pc=OZD | 06/28/2025 | 08/26/2025 |  | 07/21/2025 |  |  | 07/04/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632506 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 60 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced low pump flow. the patient was transfused volume and treated with bicarbonate. there were also suction alarms that resolved without treatment. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. low pump flows: data logs were reviewed and flows were intermittently low throughout the case. that being said, the ps was only variable during the early portions of the case. there were a total of 68 suction alarms, over 200 position unknown (alarm #35), and 13 preventing retrograde flow alarms. there were also 17 position in aorta alarms on 7/jul/2025, however, this is after any of the reports of low flows/suction, so the relation is unclear. early in the case, there were many suction alarms due to a low motor current (mc) min. the concentration of suction alarms on 30/jun/2025 occurred shortly after there was a clear step down in mc min values. there was a slight drop in ps values at the same time, but not as significant as the mc. the suction alarms that triggered shortly after were intermittent and diastolic. it is unclear based on data logs if the suction alarms were caused by pump position being slightly deep. clinical details did not report any blood or bicarbonate administration after 29/jun/2025. however, when the pump was being run at p-3 during weaning, the preventing retrograde flow algorithm was kicking in because mc values were low. mc was roughly 180/130 before mc/ms were being automatically ramped up. when the mc/ms ramped, ps values also increased. based on this information, data logs do not align with the clinical report that the patient's ejection fraction had improved by the end of the case. there is potential that the patient's condition early on in the case played a role in low pump flows, but the position of the pump could have also played a role. finally, data logs did not show low pump flows resolving to support one troubleshooting method resolving the complication. product was not returned for investigation. since limited clinical details were provided, data logs were inconclusive, and product wasn't returned, the root cause of the low pump flows couldn't be determined.
7277 | 1220648-2025-30540 | 22555587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555587&pc=OZD | 07/05/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/05/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 80 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced an absence of pulses and the physician planned to perform a femoral-to-femoral bypass. however, after several unsuccessful attempts, he opted to remove the impella cp due to leg ischemia. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7278 | 1220648-2025-30541 | 22555660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555660&pc=OZD | 06/28/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 06/28/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613554 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia. the patient was defibrillated and treated with amiodarone. additionally the patient had blood in the stool. bival was stopped and purge solution switched to d5 heparin. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7279 | 1220648-2025-30542 | 22555806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555806&pc=OZD | 07/05/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/05/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672658 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male | 72 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient developed bleeding at the impella access site. the puncture site was treated with adrenaline, but bleeding persisted. three units of blood were administered. the bleeding was successfully controlled. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7280 | 1220648-2025-30543 | 22555889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555889&pc=OZD | 07/08/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685376 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that upon explantation of an impella cp the patient received a thrombectomy. no further details were provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
7281 | 1220648-2025-30544 | 22555992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555992&pc=OZD | 07/05/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/05/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689538 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 56 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient had been implanted with an impella cp for mechanical circulatory support. during support, a small hematoma developed around the impella cp repositioning sheath. manual pressure was applied and one unit of blood was administered to manage the condition. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7282 | 1220648-2025-30545 | 22555998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22555998&pc=OZD | 07/08/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635388 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 101 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed a hematoma at the access site. heparin was withheld. additionally, the patient experienced atrial fibrillation and was treated with amiodarone. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7283 | 1220648-2025-30547 | 22556148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22556148&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583006 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 65 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | Malposition of Device | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced placement signal issues that prompted the team to question signal. the pump was found to be positioned deep, so it was repositioned. the patient also had limb ischemia that was treated by a fem-fem bypass. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of optical signal issue and limb ischemia has been completed. the impella device was not returned for evaluation. data log shows no abnormalities reported on all 5s optical signal parameters. several placement signal low alarms were noted during the case, with observed placement signal trend and pump flow trend. no abnormalities were noted in motor current or positioning. the cause of the optical signal issue was most likely patient condition related, based on data log investigation. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30547. h6 medical device problem code 2616 was erroneously entered on the initial manufacturer device report number 1220648-2025-30547. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7284 | 1220648-2025-30548 | 22556242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22556242&pc=OZD | 07/06/2025 | 09/03/2025 |  | 07/21/2025 |  |  | 07/06/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662002 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support (mcs). the patient had arrested prior to hospitalization and was found to be in ventricular tachycardia and st-segment elevation myocardial infarction. after implant of the impella cp, it was removed in minutes, as during initial attempts to optimize the impella cp pump's position, it became kinked, and the cannula folded onto itself. the pump was pulled shallow to correct. however, now the pump was under papillary on reinsertion and again folded onto itself limiting cannula flow. consequently, the pump was removed, and the kink was visualized. a new impella device was implanted without incident. the patient was transported to the unit in critical condition after unsuccessful revascularization of the circumflex artery. | [5 empty] | Section a and section e first name and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (cannula kink) and positioning issues has been completed. the impella device was not returned for evaluation. product damage (cannula kink): clinical reports during attempt to optimize pump position, the pump was pulled shallow and reinserted, and found to be under the papillary muscle and kinked, resulting in limited cannula flow and the removal of the pump. the product was not returned for investigation. the cause of the cannula kink was mostly likely use related based on the clinical details, which state that positioning was attempted to be optimized and the cannula became kinked during the process. positioning issues: clinical reports pump was pulled too shallow during attempt to optimize pump position, then it was found to be under the papillary muscle and kinked, resulting in limited cannula flow and the removal of the pump. the product was not returned for investigation. the cause of the positioning issues is most likely use related as clinical reports the impella was found to be under the papillary muscle after the pump was pulled too shallow during repositioning.
7285 | 1220648-2025-30546 | 22556300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22556300&pc=OZD | 07/08/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646491 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured ventricular tachycardia during impella cp support. an amino drip was started and two shocks were given to stabilize the patient. the pump was also repositioned via echo to help prevent future recurrence. after three days of support, the decision was made to withdraw care and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7286 | 1220648-2025-30549 | 22556540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22556540&pc=OZD | 07/06/2025 | 09/03/2025 |  | 07/21/2025 |  |  | 07/06/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687362 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female | 102 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support (mcs) and developed an access site bleed. manual pressure and an additional suture were initially placed as an attempt to manage the bleed. the team was unable to control the bleeding and the decision was therefore made to remove the impella cp in the catheterization lab. the patient was given one unit of packed red blood cells to treat the bleed. the patient's outcome is critical. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. access site bleeding: clinical reports bleeding around the access site after insertion of guidewire repositioning unit and despite pressure and an additional suture, the bleeding could not be controlled and the pump was removed. the pump was returned for investigation and there was mild kink in the cannula. the cause of the access site bleeding is not determined due to a lack of clinical details.
7287 | 1220648-2025-30550 | 22556582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22556582&pc=OZD | 07/08/2025 | 09/04/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648593 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 114 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during repositioning of the impella cp the contamination sleeve was torn. occlusive sterile dressing was applied. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. the root cause of product damage (sleeve damage) was most likely use issue since the sterile sleeve was torn by the doctor during the repositioning attempt at the bedside. h6 component code 4723 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30550.
7288 | 1220648-2025-30551 | 22557472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22557472&pc=OZD | 07/08/2025 | 09/11/2025 |  | 07/21/2025 |  |  | 07/08/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658354 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Unknown | 65 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemolysis | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient endured runs of ventricular tachycardia (vt) during impella 5.5 support. the condition self resolved and no further action was required. three days into support, a hematoma developed at the cutdown site. one unit of blood was transfused and heparin was withheld to manage the condition. the patient was then taken to the operating room for a washout to treat the site. due to ongoing concerns with the site, a warm compress was eventually placed and the range of motion of the right upper extremity was limited. two days later, the patient developed signs of hemolysis evidenced by pink tinged urine. the provider opted against repositioning the pump or adding fluids. as support continued, the patient once again endured runs of vt for which they were defibrillated. impella support continued following the intervention. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | A3 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of vf/vt/af/at, hemolysis, and access site bleeding has been completed. the impella device was not returned for evaluation. vf/vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. hemolysis: data logs were reviewed and a slight trend down in placement signal was observed on day of reported hemolysis. cvp was low at 7 on (b)(6) 2025 but was up to 13 on (b)(6) 2025. cm marking noted to be 51 on (b)(6) 2025 confirming that the pump position was never changed. physician did not want to reposition or give fluids. the cause of the hemolysis was most likely patient condition related as clinical details noted hemolysis resolved without repositioning and data logs showed a slight trend down on placement signal on day of reported hemolysis. access site bleeding - major: clinical details noted that the patient had a hematoma at the access site that required surgical intervention. the cause of the access site bleeding could not be definitively determined as limited clinical details were provided. d6b explant date added b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30551. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30551. h1 type of reportable event was erroneously selected on the initial manufacturer device report number 1220648-2025-30551. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30551.
7289 | 1220648-2025-30552 | 22559408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22559408&pc=OZD | 07/06/2025 | 09/16/2025 |  | 07/21/2025 |  |  | 07/06/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2025607785 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the impella cp was unable to advance into the ventricle due to the patient's existing aortic valve stenosis. the implant was aborted as a result of the noted resistance. no patient harm was reported due to the issue. | [5 empty] | The investigation for the reported delivery issues has been completed. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass previously received biological aortic valve replacement and was currently presenting with aortic valve stenosis and made it impossible for to insert impella into the ventricle and the insertion was aborted. corrections to the initial mfr report: d7a updated single device use to no. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7290 | 1220648-2025-30553 | 22559603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22559603&pc=OZD | 07/09/2025 | 07/21/2025 |  | 07/21/2025 |  |  | 07/09/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026695433 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support for a patient in acute myocardial infarction/cardiogenic shock. post implant, the patient had runs of ventricular tachycardia during repositioning and p level increases. the pump was kept at p5 setting and the runs stopped. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7291 | 1220648-2025-30554 | 22559801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22559801&pc=OZD | 07/09/2025 | 08/26/2025 |  | 07/21/2025 |  |  | 07/09/2025 | 07/21/2025 |  |  |  |  | 07/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691794 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.5 device for mechanical circulatory support. the pump was backloaded over abiomed 0.018 wire. the pump was unable to pass patient's anatomy. decision was then made to transition to the left. insertion was difficult but pump was eventually placed into left ventricle and wired removed. it was further reported that the patient¿s care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. updates to the initial mfr report: g1 mdr reporting contact email. h6 med dev prob code. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7292 | 1220648-2025-30464 | 22540252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540252&pc=OZD | 07/03/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661830 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient's right leg became ischemic following impella cp implant. a reperfusion sheath was placed, but the condition remained. the decision was ultimately made to explant the pump. following removal, a thrombectomy and stenting of the right common femoral artery were performed. the patient was noted to be in stable condition. | [5 empty] | Section a - patient information is unknown. e.1 first and last name are unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7293 | 1220648-2025-30465 | 22540299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540299&pc=OZD | 07/04/2025 | 08/22/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689038 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was attempted to be inserted but due to the patient¿s tortuosity and calcium, the impella 5.5 placement was aborted. an impella cp was placed instead and there was no reported patient harm. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of delivery issue was most likely patient condition since due to tortuosity and calcium they were unable to pass impella 5.5 despite many attempts with various wires. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7294 | 1220648-2025-30466 | 22540330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540330&pc=OZD | 07/04/2025 | 09/02/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693043 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp insertion was aborted after multiple attempts. the impella cp was unable to pass the aortic arch. there is no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since pump unable to move pump past aortic arch after multiple attempts. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
7295 | 1220648-2025-30467 | 22540359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540359&pc=OZD | 07/04/2025 | 09/02/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the a patient was on impella cp support and after the implant, waveforms were decoupled with lots of suction alarms. the impella was explanted and a thrombus was noted on the catheter. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. based on clinical description and data review, the root cause of low flow was most likely biomaterial ingestion. the root cause of the thrombus was unable to be determined due to insufficient clinical details. data review: data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for about 18 minutes, motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case.
7296 | 1220648-2025-30468 | 22540430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540430&pc=OZD | 07/02/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764459 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that upon start up, the automated impella controller powered on with a battery failure alarm. the device was plugged in, the power was reset, and a hard reset to the battery was completed but the issue remained. there was no patient involvement at the time of the noted failure. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. the logs confirmed the reported failure. there was a battery failure alarm. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. batttest results revealed that there was a cell imbalance in the battery in position 2. the root cause of the defective battery was due to single cell failure.
7297 | 1220648-2025-30469 | 22540523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540523&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657687 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 70 KG | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced a drop in flow and change in left ventricle placement waveform. the impella 5.5 was explanted and the patient was placed on venoarterial extracorporeal membrane oxygenation (ecmo) with a direct left ventricle vent. post-operative, the patient is doing well with ecmo flows approximately 4.5 to 5 liters per minute. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. the returned impella 5.5 pump was visually inspected finding no biomaterial, no broken blade or cannula kink observed. impella connected to aic to run test. low flows were not reproduced. data logs were reviewed finding a motor current (mc) spike, drop in placement signal (ps), speed deviation, and drop in flows observed with suction alarms 4 hours before pump explanted. no position issue or placement signal issue observed. a purge spike seen. clinical details states there was a spike was seen in the motor current waveform. concurrently there was a drop in flow and change in lv placement waveform and patient treated from history of lv thrombus and clot was suspected during support. the cause of the low pump flow issue was most likely biomaterial ingestion per data log review and clinical description. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7298 | 1220648-2025-30470 | 22540656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540656&pc=OZD | 07/02/2025 | 08/26/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678220 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | Company representation |  |  |  | 54 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5, a patient experienced ventricular tachycardia (vt) which required cardioversion. the patient had been going into vt numerous time prior to impella support. the impella was intentionally kept shallow in an abundance of caution. the patient continued to experience vt multiple times while on impella support and continues to remain on support. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. no product was returned. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
7299 | 1220648-2025-30471 | 22540664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540664&pc=OZD | 07/01/2025 | 10/07/2025 |  | 07/18/2025 |  |  | 07/01/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678048 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Mitral Perforation | Thrombocytopenia | [9 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support. a hematoma was noted at the access site. due to concerns about heparin-induced thrombocytopenia, heparin was discontinued. the hematoma remained stable, and the patient was ambulating. systemic anticoagulation with bivalirudin was initiated, with therapeutic levels achieved. there were intermittent impella position-in-aorta alarms and erratic waveforms. despite this, echocardiography again confirmed the pump was in the proper position, suggesting a possible sensor issue. the patient also developed a nosebleed, prompting the discontinuation of heparin. one unit of platelets was administered. additionally, the patient exhibited mitral regurgitation, which was attributed to the angle of the pump in the left ventricle. the pump was subsequently repositioned. the patient remained on impella support. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." | The investigation for access site bleeding, thrombocytopenia, vascular damage, positioning issues, heart valve damage has been completed. clinical reported hematoma at impella site. no additional information on how hematoma occurred was reported. the root cause of access site bleeding - major was not determined as limited clinical information was provided. the cause of the thrombocytopenia was not determined due to insufficient clinical details. clinical reported false impella position in aorta alarms with echo positioning confirming good position. the data logs do not show any motor current spikes before onset of impella in aorta alarms. the 5s parameters were stable and within expected ranges. clinical reported patient had some mitral regurgitation due to pump angle in the left e. the root cause of the heart valve damage was most likely due to improper positioning of the pump as evidenced by clinical details of mitral regurgitation due to pump angle in the left ventricle. clinical reported pump was repositioned after mitral regurgitation occurred due to pump angle in the left ventricle. it is unknown how pump became mispositioned. the root cause of the positioning issue was not determined as limited clinical information was provide. the root cause of vascular damage - peripheral vessel was not determined as limited clinical information was provided corrections to the initial mfr report: b1 malfunction added. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product returned. the data logs do not show any motor current spikes before onset of impella in aorta alarms. the 5s parameters were stable and within expected ranges. the root cause of the optical signal issue was most likely due to patient condition related issues affecting optical signal as evidenced by no motor current spikes before alarm onset and echo confirming good positioning. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d6.b. updated with additional information corrections that were made from the initial manufacturer device report number 1220648-2025-30471. d.10. concomitant medical products and therapy dates updated.
7300 | 1220648-2025-30472 | 22540713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540713&pc=OZD | 02/25/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/01/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534525 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient underwent implantation of an impella cp device for mechanical circulatory support. during support, a pump failure occurred and was associated with the development of limb ischemia. clinical intervention was required; however, specific details regarding the intervention were not provided. the ischemia ultimately resulted in limb amputation. the patient survived the device explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
7301 | 1220648-2025-30473 | 22540854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22540854&pc=OZD | 07/03/2025 | 09/09/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147287 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SERIAL NUMBER (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella automated impella controller (aic) generated communication at the beginning of the case to check the light going out the plug. the light was confirmed and the technician connected the pump. it started working and the motor current wave was visible and the aortic and left ventricle placement signal were visible too and then suddenly disappeared. it was decided to exchange for another aic on which the case was finished without any adverse events. | [5 empty] | The investigation of optical signal issue and automated impella controller (aic) - controller failure (red alarm) has been completed. console logs show that during case start, after pump 600882 was connected to the console for the first time and optical sensor g0 calibration commenced, an error message was presented to the user ¿optical sensor system error: disconnect cable¿. real time (rt) log data shows that during this time optical signal not reliable (snr) was very low (less than10), lamp was at 100 percent and gain was 2.6 (max value). pump was disconnected, then immediately connected prior to the screen prompt to do so. at the prompt, pump was disconnected again and reconnected twice. eventually g0 could be successfully calibrated and corresponding rt log data showed correct optical signal with snr greater than 8000 and gain of 1.5, indicating that optical signal issue (psnr condition) has been resolved. shortly (20s) after pump start, communication with optical bench stopped, and aic was no longer receiving data from the bench, which manifested in ¿controller error¿ alarm (1043, opm comm stopped). the logs also reported processsamplingdatafromopticalbench: sentstopacquisitioncmd ack, indicating ossiopsens receive thread gets a data sampling packet with an unexpected length. rt log data shows that at the moment of optical data loss, last sample values for all optical parameters were recorded as ¿+16384¿, and no data points were recorded in rt log file, while all other signals were still logged. there were no controller failure alarms during the case. optical signal issues (at the beginning of case) were most likely due to air gap at the optical pump connector, and were resolved after pump plug was reconnected. the log entry processsamplingdatafromopticalbench: sentstopacquisitioncmd ack indicates the sentstopacquisitioncmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false ¿communication stopped¿ condition. the absence of snr samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. it¿s been determined that optical bench lamp power was below specification, but this did not contribute to any of the observed failure modes, and was unrelated to the complaint. d9 date was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7302 | 1220648-2025-30474 | 22541007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541007&pc=OZD | 07/03/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565279 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Moisture Damage | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) had been damaged prior to use. the aic had been stored in an area where it was damaged by receiving a spray of water. no patient was involved. | [5 empty] | E.1 name prefix/title, first name and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7303 | 1220648-2025-30475 | 22541017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541017&pc=OZD | 07/12/2025 | 08/22/2025 |  | 07/18/2025 |  |  | 07/12/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678515 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was provided and reported that the patient continued to need inotrope and pressor support. urine output continued to have brown colored urine where orders were placed for dialysis. the bedside nurse discussed lack of perfusion and need for escalation. the surgeon was notified and agreed to escalate to impella 5.5. the procedural outcome was the patient survived surgery. | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, laboratory results confirmed hemolysis. the patient also experienced a run of ventricular tachycardia, for which they were defibrillated and treated with lidocaine. mechanical support was continued following the intervention. additionally, the staff encountered resistance while attempting to reposition the pump. documentation indicated that retracting the device was difficult and required multiple personnel to reposition it successfully. the patient remained on impella support. |  |  |  |  | B5 describe event or problem updated with additional information. d6b explant date updated with additional info. the investigation has been completed. no product was returned. per the clinical investigation for hemolysis: the data log shows the trend of the placement signal and flow while running on steady at the p-level during the time of hemolysis. there were no motor current spikes or positioning issues reported during that time. according to the clinical report, ecmo was running on a high-pressure circuit during the hemolysis event, which was later resolved. the cause of the hemolysis issue was most likely patient condition related, based on data log trend and provided clinical details. positioning issue: as per the clinical report, the doctor needed nursing support to press the blue button and used significant force to retract the device. the doctor also suggested that the position of the pump could have contributed to the difficulty in repositioning. the cause of the positioning issue was not determined since product was not returned for investigation. ventricular tachycardia: in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ventricular tachycardia was unable to be determined due to insufficient clinical details. | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7304 | 1220648-2025-30476 | 22541067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541067&pc=OZD | 07/03/2025 | 09/15/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274101 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Moisture Damage | Inaudible or Unclear Audible Prompt/Feedback | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was furhter reported that upon disconnecting the battery, staff noticed that the automated impella controller (aic) did not emit any sound. this issue did not occur in a clinical setting. | It was reported that the automated impella controller (aic) was in a storage closet and the air conditioner sprayed water on the console. there was no patient involvement. |  |  |  |  | B5 additional information received from the clinical site. h6 updated to reflect the additional failure mode. | The investigation of automated impella controller (aic) - buzzer/ alarm inaudible and aic ¿ damage before therapy has been completed. data analysis: the console underwent preventative maintenance on june 20, 2025. the reported complaint date was july 3, 2025. the logs show that no commercial pump was connected to ic4846 after june 20, 2025, and there was no issue in the logs that confirmed the reported failure mode. this indicates that water spray occurred while the console was in storage. device analysis: this console was tested after arriving and visual confirmed the evidence of water spray on the console. root cause: the root cause of the water spray was most likely a user-related issue. no evidence of buzzer/alarm inaudibility issues was found in the logs, and no issues were identified during testing. d9 device returned to manufacturer date added. | E.1 name prefix/title, first name and last name are unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7305 | 1220648-2025-30477 | 22541071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541071&pc=OZD | 07/02/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644103 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced supraventricular tachycardia with a wide qrs complex followed by asystole. the patient would decline and subsequently expire. the patient was found to have left anterior descending artery disease and cardiothoracic surgery was consulted. coronary artery bypass graft surgery was scheduled two days prior to the impella cp implant but was postponed due to creatinine and white blood cells increase. over the night, the patient's respiratory status declined and was intubated due to poor o2 saturation. when the catheterization lab staff arrived at the unit today, the pt had a mean arterial pressure of 30 with neosynephrine at 50mcg. the decision was made rush the patient to the catheterization lab for impella mcs. upon arrival to the catheterization lab the patient had a mean arterial pressure of 20. the impella was placed via the right femoral artery and started. the patient converted out of atrial fibrillation and mean arterial pressure improved to 80. drips were then weaned and decision made to stent the left anterior descending artery using single access. one drug-eluting stent was placed, and the patient hemodynamically looked much improved. the patient was sent to the unit to rest and recover and plans to start continuous renal replacement therapy as tolerated. approximately two days after start of the impella mcs, the patient arrested for approximately three minutes before return of spontaneous circulation was achieved; supraventricular tachycardia converted to a wide qrs complex then asystole. the patient's oxygenation had also declined. the next morning an echocardiogram demonstrated new right ventricular failure, which made it difficult to flow above impella support level p6-p7. the plan was to restart crrt, replace the swan, and lighten sedation to better assess neuro status post arrest. the patient, however, continued to decline throughout the day, became increasingly hypotensive and expired. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7306 | 1220648-2025-30478 | 22541119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541119&pc=OZD | 07/04/2025 | 09/02/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026689371 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be advanced through the patient's vasculature during attempted delivery. provider attempt to deliver the device with excessive force caused a kink in the cannula near the inlet. the pump was removed and replaced wiht another impella 5.5 as a result of the noted delivery issue. the impella 5.5 was discarded by customer. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. product damage (cannula kink): the root cause of the product damage issue was most likely patient condition related; provider attempted to deliver device with excessive force causing a kink in the cannula near the inlet. delivery issue: the root cause of the delivery issue was most likely patient condition related; impella 5.5 pump could not be advanced through the patient's vasculature during attempted delivery.
7307 | 1220648-2025-30479 | 22541231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541231&pc=OZD | 07/03/2025 | 09/09/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664175 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, x-ray and extracorporeal membrane oxygenation were done and the pump appeared deep. staff attempted to pull back several times and patient went into ventricular tachycardia. the optical signal was damaged. the pump remained deep, but physician was fine with leaving it as is due to adequate flow. additionally, the patient had internal bleeding, systemic anticoagulation was withheld. | [5 empty] | The investigation of positioning issues, vf/vt/af/at, vascular damage - peripheral vessel, and optical signal issue has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issue was not determined since product was not returned and insufficient clinical details. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: the cause of the vascular damage - peripheral vessel was not determined due to insufficient clinical details regarding the bleeding. optical signal issue: data log review showed placement signal (ps) monitoring was disabled after both the psnr events at case start and at 1100 on (b)(6) 2025. 5s parameters were nominal and pulsatile but showed no changes during psnr alarms. cross functional review determined this issue is not related to the console. the cause of the optical signal issue was not determined since product was not returned, inconclusive data log review, and insufficient clinical details. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30479. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7308 | 1220648-2025-30480 | 22541242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541242&pc=OZD | 07/03/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636173 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced heparin-induced thrombocytopenia. the patient was started on bivalirudin. impella support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7309 | 1220648-2025-30481 | 22541249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541249&pc=OZD | 07/03/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679220 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported had a 5.5 pump placed to support the patient having coronary artery bypass grafting x2 grafts in the operating room. the pump was supporting when there was new atrial fibrillation with rapid ventricular response the day after implant when being extubated. the pump remains on despite the arrhythmia. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7310 | 1220648-2025-30482 | 22541398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541398&pc=OZD | 07/04/2025 | 09/03/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had a cp support placed in femoral for the cardiogenic shock (cgs) complicating acute myocardial infarction. after 2 days it was explanted and escalated to the 5.5. the limb needed vascular to take the patient to the operating room for revascularization for ischemia. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. b5: additional information states that after the patient was escalated to the 5.5, some time later, care was withdrawn and the patient expired. b2 adverse event of death and death date added. h6 impact code 1802 added.
7311 | 1220648-2025-30483 | 22541443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541443&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/18/2025 |  |  | 06/27/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657876 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing for the patient admitted in with known disease and awaiting transplant. during the support there were multiple ¿placement signal low¿ alarms overnight and labs returned hemolyzed. abiomed representative advised medical staff to evaluate position under echo and evaluate for hemolysis. the medical doctor acknowledged input but did not proceed with imaging or additional evaluation. the following day the device again had multiple placement signal low alarms, and saw increased creatinine from 1.3 to 1.8, ldh bump to 1300, and hemoglobin without red blood cells on urinalysis. the medical doctor elected to reposition pump under echo with final depth 4.6cm across aortic valve. the patient remained in stable condition and survived. | [5 empty] | Hemolysis: the pump was repositioned and there were no more hemolysis related issues noted. the data logs show placement signal (ps)low alarms that were accurately triggered and a ventricular appearing ps around the 26th and 27th. these alarms resolve for the reminder of the case. 5s parameters are stable. based on the clinical details and data log analysis, the root cause of the hemolysis is most likely related to improper positioning. positioning issues: due to limited clinical details provided, the root cause of how the pump was not in optimal position cannot fully be determined. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7312 | 1220648-2025-30484 | 22541527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541527&pc=OZD | 07/04/2025 | 11/13/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662318 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed hemolysis. the team was advised to turn down the p-level. the patient also developed acute kidney injury and started continuous renal replacement therapy. it was further reported that the patient had non-sustained vt's and amiodarone was started. patient survived. | [5 empty] | The investigation has been completed. renal failure & hemolysis. the patient developed aki and had a trending pfhg. positioning was checked and the pump was turned down. due to lack of product, data logs and insufficient clinical information provided, the root cause of the hemolysis cannot be determined. the cause of the renal failure was not determined due to insufficient clinical details. vt/vf. on the 5th, it was stated that the patient had non-sustaining vt. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. d10 concomitant medical products and therapy dates updated. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7313 | 1220648-2025-30485 | 22541582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541582&pc=OZD | 07/02/2025 | 08/27/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684243 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female | 84 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed bleeding at the impella access site. heparin was stopped and was restarted the next day. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely anticoagulation management since hemostasis was achieved by with holding the heparin along with pressure dressing.
7314 | 1220648-2025-30486 | 22541618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541618&pc=OZD | 07/04/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient experienced multiple complications during impella cp support. after one day of support, it was noted that the patient experienced bleeding issues from their access site, a gastrointestinal bleed, and bloody output from their nasogastric tubing. blood products were given to address the ng output and gi bleed. the dressing was changed and a sandbag was placed to manage the oozing from the access site. in addition, it was noted that the patient's impella limb was cooler with no pulse in the right dorsalis pedis or posterior tibial. a plan was made to wean pressors to stimulate flow to the leg. support continued following the interventions, however the patient expired two days later. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7315 | 1220648-2025-30487 | 22541655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541655&pc=OZD | 07/03/2025 | 08/30/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670061 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 74 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to the patient's tortuous and calcified iliac arteries. an intra-aortic balloon pump was placed instead. the patient survived. | [5 empty] | The investigation has been completed. no product was returned. the root cause of delivery issue was most likely patient condition since the pump was unable to advance impella through the femoral artery as calcification was noted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7316 | 1220648-2025-30488 | 22541719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541719&pc=OZD | 07/04/2025 | 09/02/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a 5.5 pump being placed to support the patient admitted in with cardiomyopathy and atrial fibrillation as well as other comorbidities. the pump failed to deliver and when the pump was explanted they observed the delivery attempts may have caused the vessel damage. the surgeon had to be called for axillary artery repair. patient survived. | [5 empty] | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be use issue because 0.018 steel core non abiomed guide wire is used during insertion which is not recommended as per impella 5.5 instructions for use. later, tried with 0.035 wire and unable to pass wire into axillary after multiple attempts and found that axillary dissection in the portion of aorta. vascular damage - peripheral vessel: the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. there was information related to cause of the dissection in the right axillary artery as per the clinical details. later the vascular damage was repaired and second pump was opened and inserted successfully. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7317 | 1220648-2025-30489 | 22541826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22541826&pc=PYX | 07/03/2025 | 09/02/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025533886 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 136 KG | Non Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella rp flex support. neither the signal waveform nor values were present on the console. despite flow calculations being disabled, the motor current and purge values remained stable. the failure occurred throughout the course of support. upon eventual wean and explant of the device, it was documented that oozing persisted despite application of multiple sutures. a femostop was subsequently placed to help manage the condition. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding and optical signal issue has been completed. the impella rp flex was returned for evaluation. access site bleeding - major: the cause of the access site bleeding could not be definitively determined as limited clinical details were provided. optical signal issue: optical parameters of returned pump were consistent with a damaged optical sensor face. pump passed laser continuity testing. visual inspection of the optical sensor face showed damage to the silicone with a dent in the metal surrounding the sensor face. data logs were reviewed log showed loss of optical data with a placement signal not reliable (psnr) alarm. the real time (rt) log at time of psnr alarm shows optical 5s parameters consistent with a sensor face damage, with a motor current spike and speed deviation. clinical details noted that patient had been put on dialysis on morning of second day of support around time when psnr alarm occurred per data logs. the cause of the optical signal issue was due to a damaged sensor face from external device interaction of the dialysis device based on returned product evaluation and data log analysis showing damage to the sensor with a motor current spike at time of damage and clinical details indicating that a dialysis catheter was placed around time of psnr alarm. placement signal issue: a visual inspection of the returned pump showed no damage to the dp sensor face. based on analysis of returned data logs indicating issue with the dp sensor on the rp flex pump being investigated, an additional failure mode of "placement signal issue" was added to the investigation. data logs were reviewed and log showed placement signal and pump flow drifting prior to going out of spec at time of psnr alarm #1051. additionally, cvp at time of alarm goes out of range. returned product was electrical tested and all sensor electrical lines were within range. pump was run in a flow loop test for 56 hours and sensor drift of the dp sensor was reproduced. clinical details noted that the psnr alarm occurred after patient was transferred to icu but no additional clinical details were provided at time of alarm. the cause of the placement signal issue was due to dp sensor drift based on returned product analysis and data log review. | The device was received, but evaluation has not been completed. upon completion of the investigation, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7318 | 1220648-2025-30490 | 22542030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542030&pc=OZD | 06/28/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 06/28/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574808 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 70 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced arrhythmia. transvenous pacing was placed. additionally, the patient had melena in their stool. two units of packed red blood cells were transfused. the patient was successfully weaned from the pump and survived. | [5 empty] | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
7319 | 1220648-2025-30491 | 22542078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542078&pc=PYX | 06/30/2025 | 08/28/2025 |  | 07/18/2025 |  |  | 06/30/2025 | 07/18/2025 |  |  | 08/31/2025 |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025453571 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 17 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella rp was complicated as the patient was a hancock implant with four stents. the team had many attempts and troubleshooting with multiple wire exchanges. implantation was aborted and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because patient had unfavorable anatomy and after multiple attempts and after change in guide wire as per clinical details and insertion was aborted. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30491. d6a and d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30491. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30491. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-30491. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30491.
7320 | 1220648-2025-30492 | 22542191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542191&pc=OZD | 06/18/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | EMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1418929 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the automated impella controller (aic) had a controller error alarm. placement signal was lost temporarily but have since reappeared. the issue resolved without intervention. support continued. no patient harm reported due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7321 | 1220648-2025-30493 | 22542210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542210&pc=OZD | 07/03/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595430 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced migration of the pump into the left ventricle. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7322 | 1220648-2025-30494 | 22542240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542240&pc=OZD | 07/04/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683594 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient in acute myocardial infarction and cardiogenic shock. the patient had ventricular fibrillation this morning, did not require cardiopulmonary resuscitation due to impella flows. the patient was shocked x2 to break rhythm. reviewed position of impella on echo, pulled back about 1cm. later it was reported that there was some platelet aggregation vs. tissue noted on the end of the impella pigtail catheter at explant. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7323 | 1220648-2025-30495 | 22542322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542322&pc=OZD | 07/04/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586990 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 79 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support developed limb ischemia resolved with contralateral perfusion. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7324 | 1220648-2025-30496 | 22542417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542417&pc=OZD | 07/01/2025 | 08/28/2025 |  | 07/18/2025 |  |  | 07/01/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660941 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 76 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematoma | Hemolysis | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was implanted with an impella 5.5 for mechanical circulatory support. the pump had been placed via the axillary artery and was providing support when the patient developed a hematoma leading to a pause in heparin therapy. the patient had elevated plasma-free hemoglobin levels and concerns for hemolysis which prompted repositioning of the pump. iron supplementation was administered due to suspected hemolysis-related anemia which the physician believed had been caused by the pump. additionally, the patient experienced ventricular tachycardia while on impella support. despite these complications, the patient remained on impella support. | [5 empty] | The investigation of hemolysis, access site bleeding, and vf/vt/af/at has been completed. pump was not returned for investigation. data log shows that a motor current spike was reported along with a drop in pump flow on 6/28 (two days prior to the reported hemolysis). purge parameters were not affected during the time of the reported motor current spike. there is no indication of trendy placement signal, positioning issue, or motor current spike during the time of the reported hemolysis event, from 06/30 to 07/02. hemolysis: data log and clinical information was not sufficient to derive hemolysis root cause. the cause of the hemolysis issue was not determined since without returned product investigation. access site bleeding: the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. vt/vf/at/af: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7325 | 1220648-2025-30497 | 22542566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542566&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661323 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | High Blood Pressure/ Hypertension | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and became hypertensive and hemodynamically unstable requiring intervention. the patient presented to the emergency department with severe chest pain and shortness of breath and an electrocardiogram showed an anterior st-segment elevation myocardial infarction. the patient went into ventricular fibrillation arrest requiring one round of advance cardiovascular life support and 10 defibrillator shocks. return of spontaneous circulation was achieved and the patient was transported to the catheterization lab. catheterizations were completed with percutaneous coronary intervention (transluminal coronary angioplasty and a drug-eluting stent) to the left anterior descending artery. an impella cp was implanted via the right femoral artery without incident. while prepping the patient was transfer via lifeline. the patient became extremely hypertensive for an extended amount time. sedation was increased and the patient was transferred uneventfully. now into the next day, upon arrival in the unit, there was a struggle to correct gases and electrolytes. point of care ultrasound showed the impella cp pump to be slightly deep. however, the team felt no need to reposition at such time. the decision was made to decrease flows to a p-4 support level as the patient remained hypertensive. shortly thereafter, the patient had "several" events where the mean arterial pressure dropped to the 30's. the patient was started on epinephrine and levophed. systemic heparin was started. the decision was made to not complete the automated impella controller (aic)-to-aic exchange as the patient was unstable. approximately eight hours later, the patient was noted to be hemodynamically stable and off all drips (continuing on heparin and amiodarone). pulses palpable for bilaterial lower extremities. the aic-to-aic exchange was completed. the next day it was noted the patient was successfully weaned from the impella cp support and the device was explanted with a survived outcome. | [5 empty] | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the hemodynamic instability and hypertension has been completed. clinical details state that hypertension and pump flow level was reduced to p4. the pump was returned for investigation and no kinks or abnormalities were observed. the root cause of the hypertension and hemodynamic instability was unable to be determined due to insufficient clinical details.
7326 | 1220648-2025-30498 | 22542673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542673&pc=OZD | 07/02/2025 | 09/15/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1484084 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella pump for mechanical circulatory support and while transferring the patient via helicopter from the hospital to another, the automated impella controller (aic) had a controller failure alarm. as there was no available replacement, the impella pump continued to support the patient on this aic. the pump ran without incident and the alarm was said to have disappeared after approximately five minutes. it was reported there was no harm to the patient as a result of this event. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the impella device was not returned for evaluation. data analysis: review of the log confirms the complaint. an error of ¿bad_pressure_sample_mask¿ occurs, followed two minutes later by a controller error alarm due to ¿optical bench ambient pressure out of range.¿ the console software has hard coded limits for minimum and maximum ambient pressure. these are 550mmhg and 825mmhg, respectively. ambient pressure measurements outside of this range and lasting for 2 minutes or more will trigger a controller error alarm due to what is perceived as an invalid measurement. the complaint states that the alarm occurred during helicopter transport between facilities. the decrease in ambient pressure that occurs with increased altitude during air transport has been known to trigger ambient pressure controller error alarms in the past. there was no issue found during the case. the device functioned/operated to specification. d.3. email should have been left blank on 1220648-2025-30498. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7327 | 1220648-2025-30499 | 22542847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22542847&pc=OZD | 07/02/2025 | 08/30/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487327 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 108 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support during protected high-risk percutaneous coronary intervention. during the support, the impella cp was found to be shallow and directed towards the mitral valve. while attempting to reposition the pump, the device was inadvertently pulled into the aorta and was subsequently explanted with a survived outcome. subsequent treatment was unnoted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue has been completed. product was not returned for investigation. clinical details state the pump was found to be shallow and directed towards the mitral valve. while attempting to reposition the pump, the device was inadvertently pulled into the aorta. the cause of the positioning issues was most likely due to a use related issue, as the clinical reports the pump was inadvertently pulled out by the physician.
7328 | 1220648-2025-30500 | 22543008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22543008&pc=OZD | 07/02/2025 | 08/30/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experience low pump flow. the impella cp device was placed with through the left femoral artery with noted minimal difficulty, although the device was inserted through a graft due to the patient¿s previous abdominal aortic aneurysm. the companion sheath was utilized for single access. once placed, the impella was turned on to auto and immediately suction alarms triggered. the support level was reduced to p-5 with continued suction and limited flows. it was noted the patient had "extremely" unique arch and cardiac anatomy which appeared to have kinked the device at the angle of insertion across the new prosthetic valve. the patient hemodynamics remained stable at support level p-5. the percutaneous coronary intervention was completed and the impella cp device was weaned and explanted with a survived outcome. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation for the low pump flow and product damage has been completed. product was not returned for investigation. data logs were reviewed which confirmed low pump flow when pump started and remained to the rest of the case. brief position alarms which resolve. suction alarms confirmed. clinical states almost immediate suction when pump turned on and low flow occurred when pump started. patient had extremely unique arch and cardiac anatomy which appeared to kink the device at the angle of insertion across the new prosthetic valve. the root cause of the low flow was most likely kinked cannula per clinical description and data log review. the root cause of the cannula kink was most likely patient condition related per clinical stating patient had extremely unique arch and cardiac anatomy which appeared to kink the device at the angle of insertion across the new prosthetic valve. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7329 | 1220648-2025-30501 | 22543313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22543313&pc=OZD | 07/02/2025 | 08/30/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677193 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 158 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and two days after start of the support, upon waking up, the patient inadvertently pulled the impella cp pump out of its position. consequently, the decision was made to replace the impella cp which was successfully completed. the patient remained on impella cp mcs for one additional day as a right ventricular assist device was placed along with valve replacement and the impella cp was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue has been completed. the device was not returned for evaluation. the clinical details state that the patient woke up and pulled the impella out. based on the clinical details provided, the root cause of the positioning issues is most likely due to patient movement as the patient pulled out the impella.
7330 | 1220648-2025-30502 | 22543885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22543885&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support there was an impella stopped/ mc high alarm that self resolved, although the pump never stopped. there was a spike in the motor current. the flows were now saturated. additionally it was reported of possible thrombus ingestion. the impella was successfully removed due to possible thrombus ingestion. patient is stable post explant. | [5 empty] | The investigation of saturated flow, pump stop, thrombus, and optical signal issue has been completed. the pump was not returned for investigation. the data log shows that the pump stop occurred on the 9th day with alarm 13 ¿ pump stop. the pump was able to restart with saturated pump flow. also, the lvp reading spiked along with the motor current spike and remained high afterward. there was no known pump optical failure trend found during reported left ventricle (lv) signal issue. the pump was explanted within 2 hours after saturated pump flow was reported. pump stop: the data log alone is not sufficient to determine the cause of the issue, as the pump was used beyond its design life. the cause of the pump stop issue was not determined since product was not returned. saturated pump flow: the cause of the saturated pump flow issue was not determined since product was not returned. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. optical signal issue: data log review was not sufficient to address the optical signal issue. the cause of the optical signal issue was not determined since product was not returned. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30502. h10 instructions for use were incomplete on on the initial manufacturer device report number 1220648-2025-30502. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ' section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7331 | 1220648-2025-30503 | 22544027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22544027&pc=PYX | 07/04/2025 | 09/02/2025 |  | 07/18/2025 |  |  | 07/04/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625171 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during support, there was bleeding at the access site. an additional suture was placed. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause access site bleeding is determined to be use issue because the bleeding around the sheath was stopped and hemostasis was achieved by placing additional suture. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7332 | 1220648-2025-30504 | 22544037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22544037&pc=OZD | 07/02/2025 | 08/30/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655791 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful. the device was unable to pass due to the patient's anatomy. there was no report or indication of adverse effects to the patient as a result of the event. subsequent actions taken to treat the patient were unnoted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
7333 | 1220648-2025-30505 | 22544184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22544184&pc=OZD | 07/07/2025 | 08/21/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680111 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Increase in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after impella cp was delivered, high purge pressure was noted. the physician was able to resolve the high purge pressure alarm. then the cp had suction alarms and was unable to flow greater than 1.5 liters per minute even at a lower p-level. additionally, there was difficulty positioning the impella. impella cp pump was replaced with a new cp pump. | [5 empty] | The impella device has been received, but evaluation has not been completed. upon completion of the investigation, a supplemental mdr will be filed. | The investigation for high purge pressure, low pump flow, and positioning issues has been completed. no physical damage was noted other than a cannula kink on the returned product. the purge gap and impellor was free of biomaterial. the pump was primed without issue and was able to run according to specifications in a bucket of water. data logs show low pump flow from the start of the case; the pump was not able to exceed approximately 3.0 l/min flow. additionally, a gradual purge pressure rise was noted after case start but returned to normal limits after 15 minutes of case run. as per the clinical report, the troubleshooting method for resolving the high purge pressure issue was unknown. the doctor reported suction, difficulty in positioning, and low pump flow. the pump was then replaced with a new one. the cause of the high purge pressure issue was most likely biomaterial, based on returned product investigation and data log review. the cause of the low pump flow issue was most likely kinked cannula, based on returned product investigation and data log review. the cause of the product damage (cannula kink) issue was not determined without provided sufficient clinical information. the cause of the positioning issue was not determined without provided sufficient clinical information.
7334 | 1220648-2025-30506 | 22544257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22544257&pc=OZD | 07/07/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Unspecified Vascular Problem | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was noted that in the lab the vessel size was borderline with concern for calcium lesion however, decision was made to proceed with impella placement for hrpci. the impella was advanced in normal fashion, flows were initiated and single access was obtained for percutaneous coronary intervention (pci). once pci was concluded, impella was weaned and ready to explant. explanted and the peel away was removed, perclose was then advanced, however perclose would not deploy despite multiple efforts. final 2 perclose took however, an 8 fr introducer had to be advanced to maintain hemostasis. at this time, distal flow was assessed. there was no flow past insertion site and the patient¿s blood pressure also began to drop. vascular surgery was consulted. levophed and fluid boluses were administered for hypotension. iliac was reassessed and saw possible perforation. decision was then made to advance a balloon into the iliac to tamponade while vascular surgery prepped the operating room. two units of packed red blood cells was administered. anesthesia then took patient to the operating room for vascular surgery intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿
7335 | 1220648-2025-30507 | 22544756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22544756&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552538 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon inserting the impella cp, the catheter kinked due to the patient having severe vascular disease. the impella placement was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of access site bleeding is determined to be patient condition because the catheter was kinked during insertion. patient was having a vessel condition and the catheter was kinked during the insertion through the vessels. product damage(catheter kink): the root cause of catheter kink was determined to be patient condition because of the patient's severe vascular disease and catheter was kinked during insertion. second pump was used and placed successfully. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2025-30507. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30507.
7336 | 1220648-2025-30508 | 22545290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22545290&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/18/2025 |  |  | 07/02/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026697263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported impella cp and extracorporeal membrane oxygenation (ecmo) support was placed to support a patient with angina and ventricular septal defect. the patient experienced ventricular tachycardia with turns. vascular addressed a large hematoma in the arterial side of the ecmo. the impella cp was assessed on echocardiogram and was confirmed to be shallow (2 centimeters) but the pigtail was touching the posterior wall with a very small left ventricle cavity. they opted to leave as is since there was adequate flow. many attempts at repositioning were done but the small left ventricle still was not optimal. the patient survived. | It was reported that there were hemolysis signs confirmed with elevated plasma free hemoglobin levels. plasma free continued to increase, witch resolved after escalation to impella 5.5. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation has been completed. complaint device not returned for analyzing. data log reviewed: on (b)(6) 18 hours into support there is a mc spike at p6 followed by slightly dampened mc. this event is days prior to the reported hemolysis and the pump metrics, such as mc, shortly resolve. there are suction alarms throughout support and position unknown alarms. the root cause of the tachycardia was unable to be determined due to insufficient clinical details. the cause of hemolysis issue was most likely improper positioning per clinical information that unable to get optimal positioning due to small lv. the cause of positioning issues was most likely patient condition related due to small lv. the cause of vascular damage was not determined due to lack of clinical details. b5 updated with information that was inadvertently omitted on the initial manufacturer device report number. h6 health effect clinical code 1884 hematoma was inadvertently omitted on the initial manufacturer device report number. h6 health effect impact code 4629 device revision inadvertently omitted on the initial manufacturer device report number.
7337 | 1220648-2025-30509 | 22545720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22545720&pc=OZD | 07/07/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615119 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was placing the impella 5.5 pump for the support of the patient admitted in with cardiogenic shock and cardiomyopathy. the pump was going to be inserted via the graft/axillary site but, there was a vessel tear and blood was lost. the team performed surgical repair and a new graft was placed. pump was successfully placed and remains on for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7338 | 1220648-2025-30510 | 22545776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22545776&pc=OZD | 07/03/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678216 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 60 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient presented for planned impella 5.5 insertion with likely transplant listing. the patient was prepped and draped in standard fashion. the graft was sewn on in standard fashion. left ventricle access was obtained. the impella 5.5 was primed and inserted without complication. during support, it was noted that the patient had ventricular tachycardia issues unrelated to impella as its baseline. the impella 5.5 was successfully weaned and explanted as a bridge to transplant. the patient outcome was notes as successful transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7339 | 1220648-2025-30511 | 22545865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22545865&pc=OZD | 07/05/2025 | 09/08/2025 |  | 07/18/2025 |  |  | 07/05/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675303 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support ongoing for acute myocardial infarction and cardiogenic shock when there was pump positioning issues during cardiopulmonary resuscitation. the pump dislocated in aorta and the pump was able to be repositioned by the medical doctor successfully. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of positioning issues have been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely patient movement per clinical that the pump dislocated in aorta after patient's convulsive fit and then cardiopulmonary resuscitation (cpr). | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7340 | 1220648-2025-30512 | 22545893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22545893&pc=OZD | 07/02/2025 | 09/16/2025 |  | 07/18/2025 |  |  | 07/03/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683156 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp insertion was aborted due to the patient¿s vascular anatomy. no harm was reported. | [5 empty] | The investigation of delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since pump could not be advanced via the aorta. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7341 | 1220648-2025-30513 | 22545899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22545899&pc=OZD | 07/07/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026691056 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient experienced some slight oozing at their impella access site during impella 5.5 support. no noted intervention was required to manage the condition. in addition, the md stated that despite functioning correctly, some clots/thrombus were visible following the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular. perforation)¿
7342 | 1220648-2025-30514 | 22546068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22546068&pc=OZD | 07/07/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656110 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was admitted for inferior st-elevation myocardial infarction. the impella was inserted into left ventricle and position was confirmed. during support, a hematoma was observed. manual pressure was held in addition to re-suturing and heparin was stopped. the groin was bruised. ultimately the patient expired on support | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
7343 | 1220648-2025-30515 | 22546129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22546129&pc=OZD | 07/04/2025 | 07/18/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683073 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. post explant of impella the patient developed a leg ischemia. physician confirmed that the leg ischemia is not related to the impella but there is not enough information to exclude the impella device as an associated factor with the outcome. | [5 empty] | A5: the patient race and ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
7344 | 1220648-2025-30516 | 22546421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22546421&pc=OZD | 07/07/2025 | 09/03/2025 |  | 07/18/2025 |  |  | 07/07/2025 | 07/18/2025 |  |  |  |  | 07/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663932 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a groin bleed. perclose were in place to help manage the condition and heparin remained on hold. in addition, the placement signal went out unexpectedly. the signal reappeared after roughly fifteen minutes. support continued uninterrupted. | [5 empty] | The investigation of access site bleeding and optical signal issue has been completed. no product was returned for evaluation. access site bleeding: the cause of the access site bleeding could not be definitively determined as limited clinical details were provided at time of bleeding event. optical signal issue: data logs were reviewed and all optical parameters are showing intermittent spiking to 3000 with "controller error" alarms triggered each time consistent with a failure of the pump console. optical parameters the remainder of the case are stable and within spec. upon review of the returned data logs, it is apparent that the console is the potential cause of the issue. please refer to manufacturer device report 1220648-2025-30660 for an investigation into the console. no problem was found with the pump per data log review and clinical details. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2025-30516. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30516. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
7345 | 1220648-2025-30409 | 22529441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529441&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/17/2025 |  |  | 06/27/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679899 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 63 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Difficult to Remove | Malposition of Device |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hematuria | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced low pump flow, thrombocytopenia, hematuria, and removal issues resulting in bleeding and limb ischemia requiring intervention. the patient was taken to the catheterization lab for suspected st-segment elevation myocardial infarction and was hemodynamically unstable and on multiple vasoactive drips. an angiogram showing left anterior descending artery occlusion. the decision was made to place an impella cp prior to the percutaneous coronary intervention (pci). the impella cp was successfully placed and pci via the right radial artery was completed with a stent to the left anterior descending artery. pressors were weaned down status post pci and the sheath was removed. the patient was sent to the unit on impella mcs at performance level p-9. a couple of hours later, suction alarms triggered, which improved with 500 cubic centimeters of fluid bolus. urine output was pink. the patient to be monitored and impella cp was to be adjusted if urine color remained. the following day, the patient had position alarms that were resolved under echocardiogram guidance and advancement of the impella. additionally, labs revealed a drop in hemoglobin and platelets. the patient was transfused with one unit of packed red blood cells and one unit of platelets. heparin was discontinued and a switched was made to bivalirudin. heparin-induced thrombocytopenia panel was drawn. the patient had suction alarms above performance level p-5. the patient responded well to the volume. it was noted that there was concerns for flow down the right lower extremity due to a purple/discolored toe the next day. however, a right lower extremity ultrasound confirmed blood flow down the leg. the physician noted likely removal of the impella cp at bedside later in the day which was eventually completed. the patient had a constant ooze from right groin sheath site. during the attempt to remove the pump and the sheath together, the unit was unable to come out of the body. a "large" amount of blood loss occurred during this attempt to remove. the pump and introducer were advanced back into the artery and the patient was taken to the operating room for an angiogram to ensure no vascular injury at the site of the arteriotomy had occurred. the pump was removed in the operating room and the physician used the previously placed percloses to achieve hemostasis. no further surgical intervention was needed to close the site. the patient received an additional two units of packed red blood cells and one unit of platelets after this removal event. however, prior to the discharge from the operating room back to the unit, the patient's right lower extremity was "heavy" and there were concerns for compartment syndrome, which was confirmed. consequently, a fasciotomy was performed. the patient survived the explant and was noted to be in critical condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, thrombocytopenia, removal issues, positioning issues, low pump flow, hemolysis, and limb ischemia has been completed. the product was not returned for investigation. the data log shows that several suction alarms were reported during the case, with a trending flow and placement signal while running on a steady p-level. no motor current spikes or major positioning issues were noted during the case run. low pump flow: according to the clinical details, suction events were managed with fluid administration, and it was also noted that the patient could not tolerate the high p-level during the case run. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. hemolysis: according to the clinical notes, hemolysis was reported during the case, with slight resolution noted on the second day of pump use. the clinical notes also confirm low pump flow and fluid blood volume issues. the cause of the hemolysis issue was most likely patient condition related, based on data log review and given clinical details. access site bleeding: the cause of the access side bleedings issue was patient condition related to coagulation disorder, based provided clinical details. positioning issues: the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provide. removal issues: the cause of the removal issue was not determined since product was not returned and sufficient clinical information was not provided. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ h10 instructions for use were incomplete on initial manufacturer device report number: 1220648-2025-30409.
7346 | 1220648-2025-30410 | 22529505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529505&pc=OZD | 10/14/2024 | 07/17/2025 |  | 07/17/2025 |  |  | 03/06/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122733 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal not reliable alarm with erratic placement signal waveforms. there was no patient harm. additionally, during investigation of the device, it was found that there was an automated impella controller alarm inaudible issue. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show psnr events recurring frequently over some periods of time, which coincided with low snr, often with same pulsatility as placement signal. console was tested in danvers, and some anomalies were found within its optical system. the ccd detector on the optical bench seemed to have a defect (a ¿notch¿ in the baseline) that distorted fringe pattern, preventing proper placement signal detection. although the unit initially passed ¿5s¿ specification with the console¿s housing open, it¿s been noted that optical signal degraded when aic housing was closed. this was due to improper routing of the optical connector fiber (¿pigtail¿) which, when caught between the housing and ethernet cable, resulted in fiber kink, which significantly lowered snr. also the ferule on the connector, that holds optical fiber, was mechanically compromised. console inspection and testing also revealed malfunction of the pbm which caused alarm buzzer to be out of spec (unrelated to complaint).
7347 | 1220648-2025-30411 | 22529511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529511&pc=OZD | 06/26/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  | 05/31/2025 |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024347570 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was escalated to an impella 5.5 for a patient admitted in with acute myocardial infarction/cardiogenic shock. the impella 5.5 was placed and did support for 35 days when the pump was explanted for a left ventricle assist device placement. during the support, the pump reached the expiration date but was not replaced. on the day of explant there was a pump thrombosis observed. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7348 | 1220648-2025-30413 | 22529586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529586&pc=OZD | 03/20/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593561 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had bleeding from the extracorporeal membrane oxygenation site. anticoagulation was withheld. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used: ¿aki secondary to atn. started on crrt then transitioned to hd. hd catheter removed prior to discharge due to kidney function improvement.¿. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7349 | 1220648-2025-30414 | 22529642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529642&pc=OZD | 12/29/2024 | 10/03/2025 |  | 07/17/2025 |  |  | 06/09/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136384 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon initiation of impella cp support display appeared ¿connect grey connectors, arrow to arrow¿ while automated impella controller is updated 12.2 and device recognized as impella cp with smart assist. after impella started, ¿lv disabled¿ immediately. impella gave adequate flows and motor current. was able to enable lv waveform manually. there was no patient harm. additionally, during investigation it was found that the pump was unable to be recognized by the console. this includes impella defective alarms. | The complainant reported that upon initiation of an impella cp support, the lv waveform was immediately disabled on the automated impella controller (aic). the staff was able to manually re-enable the lv waveform for the planned support. this complaint was related to a clinical procedure. there were no adverse events caused by the reported console issue. care was withdrawn and the patient expired. |  |  |  |  | The investigation has been completed. the automated impella controller was returned for investigation. the logs showed the case associated with the reported complaint was initiated on december 29, 2024. at case start - step 3, the epm was reset because the pump connection was not recognized for 20 seconds. the console then displayed ¿connect cable: connect grey connectors screen,¿ which aligns with the reported claim. this type of intermittent pump detection issue aligns the previously investigated epm crystal issue. after 3 minutes, although the pump (sn: (b)(6) was detected, the user unplugged, re-plugged, and started the pump. then the support continued for 3 days. the epm crystal issue is characterized by an absent i2c clock signal which prevents the console from recognizing a pump connect. software release 8.3 should reset the epm sau, which will act as a mitigation for ¿pump not recognized due to epm crystal issue¿. device analysis: this console was tested after arriving in fs. unable to reproduce the pump detecting issue. the console was tested by power cycling 20 times while observing the i2c signals, and each time, after initiating the case start and following the epm reset, the test pump connection to the aic was recognized without any issues. during the testing, the console displayed ¿controller error (defective optical sensor lamp) ¿ alarm,¿ event # 1039, and the log recorded abnormal optical signals, particularly the light parameter was 0.48v. the power to the lamp was 4.8v, and there was no issue with the analog output, confirming that there was no problem with the optical bench. this indicates that event was most likely due to a defective lamp (which is unreported and unrelated to the reported failure mode). found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). root cause: the pump detection issue was due to a rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a software mitigation in place. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. device analysis: this console was tested after arriving in fs. unable to reproduce the pump detecting issue. the console was tested by power cycling 20 times while observing the i2c signals, and each time, after initiating the case start and following the epm reset, the test pump connection to the aic was recognized without any issues. during the testing, the console displayed ¿controller error (defective optical sensor lamp) ¿ alarm,¿ event #1039, and the log recorded abnormal optical signals, particularly the light parameter was 0.48v. the power to the lamp was 4.8v, and there was no issue with the analog output, confirming that there was no problem with the optical bench. this indicates that event #1039 was most likely due to a defective lamp (which is unreported and unrelated to the reported failure mode). found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). root cause: the pump detection issue was due to a rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a software mitigation in place. before returning this console (ic2502) back to the customer, fs was instructed to perform the following, replace the front panel optical connector (0042-2736). [due to the debris] replace the optical bench lamp (2450-0038), or the entire optical bench (0042-1012) if parts are not available. [due to the controller error-unrelated to the reported failure mode] perform section 23 of the pm procedure (b)(6). perform a full functional check (b)(6). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added death "1802" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
7350 | 1220648-2025-30412 | 22529671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529671&pc=OZD | 06/30/2025 | 09/15/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672573 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp had a high purge pressure alarm. after inspection, the purge system had a broken purge filter. the impella cp was replaced and the patient survived. | [5 empty] | The investigation of high purge pressure and purge leak ¿ pump has been completed. impella cp was returned for evaluation. high purge pressure: upon visual inspection, biomaterial was present around impeller blade and in purge gap. pump was unable to be purged as purge reservoir to filter joint was returned broken. no data logs were returned for evaluation. clinical details noted that a "high purge pressure" alarm was triggered during support. the cause of the high purge pressure could not be definitively determined as limited clinical details were provided and no logs were returned for analysis. purge leak - pump: clinical details noted that physician found that the purge filter was broken during support. physician requested instructions for purge filter bypass. upon visual inspection of returned pump, purge reservoir to filter joint was fully separated. the cause of the purge leak pump was most likely damage to the purge reservoir to filter joint per returned product analysis. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
7351 | 1220648-2025-30415 | 22529699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529699&pc=OZD | 06/30/2025 | 08/27/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686703 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 57 KG | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced low flows and a perforated left ventricle. implantation was complicated and required rewiring the catheter to avoid the papillary muscle. the perforation was repaired surgically, but the patient expired overnight. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data logs confirm low flows during case. no persistent positioning alarms noted on the logs. repeated suction alarms occurring throughout log. first real time (rt) log of case shows no ingestion events occurring when low pump flow first occurred. clinical details noted that pump was not able to flow above p-4, and pump was found to be behind the papillary muscle. pump was repositioned and pump was still not able to flow above p-4. pump was re-wired to continue support. the cause of the low pump flow could not be definitively determined as no definitively ingestion event was observed in the logs, and clinical details noted that pump repositioning did not resolve the low flows. ventricle perforation: clinical details noted that pump was found to be behind the papillary muscle. pump was able to be repositioned but patient decompensated and was taken to operating room. when patient's chest was opened a left ventricle perforation was noted. perforation was repaired and patient remained on impella support. the cause of the ventricle perforation was most likely due to improper positioning based on clinical details of pump in papillary muscles.
7352 | 1220648-2025-30417 | 22529785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529785&pc=OZD | 07/01/2025 | 08/29/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1952532 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) was having issues with configuration. after they turned the controller on, the cloud and airplane button did not illuminate. troubleshooting included: power cycle, usb plugged in, airplane mode, and 60-second factory reset, which were all unsuccessful. this issue occurred during configuration, not in a clinical setting. | [5 empty] | The investigation of optical signal issue has been completed. teh console was received for investigtaion. based on the data and device analysis there was no optical signal issue. the console functioned within specifications. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
7353 | 1220648-2025-30416 | 22529786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529786&pc=OZD | 06/30/2025 | 08/27/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677857 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced a gastrointestinal bleed. the patient was transfused a bolus and two units of packed red blood cells. anticoagulation was withheld. the patient also experienced ventricular tachycardia so the p level was reduced. there were persistent suction alarms and the pump was found to be positioned under the papillary muscle. it is unknown if the positioning issue or suction issue was resolved. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel, positioning issues, vf/vt/af/at, and low pump flow has been completed. the impella device was not returned for evaluation. vascular damage - peripheral: the cause of the gi bleed was not confirmed, however. product was not returned for investigation. since insufficient clinical details were provided, the root cause of the vascular damage could not be determined. low pump flow: data logs were reviewed and the sudden suction alarms were confirmed. at 01:30 on (b)(6) 2025, there was a sudden drop in placement signal (ps) values but increased pulsatility that aligned with a very sudden drop in mean motor current (mc) but increased modulation. though the ps values dropped, the pressure trace still had an aortic characteristic, and mc modulation increased, so it doesn't seem likely that the pump dropped into the ventricle. the pulsatile suction algorithm, or continuous suction triggered shortly after. there was also a concentration of position in aorta alarms later that day, however, they appear to be unrelated to the initial onset of the suction alarms. pump flows remained low for the remainder of the case and the ps continued to trend down. based on the clinical details provided, pump position seems to be the most likely cause of the suction alarms. however, it is unclear whether or not the very sudden drop in ps and mc at the time suction began to trigger correlates with any pump positioning issues. additionally, the ps on data logs trended down from that point forward, aligning with the patient having a gi bleed and then expiring on support that day. since there are insufficient clinical details, product wasn't returned for investigation, and data logs were inconclusive, the root cause of the low pump flows could not be determined. positioning issues: data logs were reviewed and there were 3 position in aorta alarms on (b)(6) 2025. based on the sudden change in mc pulsatility, it appears this shift of the pump was intentional (repositioning) but there were no repositioning events reported. there were also 2 position in ventricle alarms at the end of the case, so they were likely during pump explant. product wasn't returned for investigation. since insufficient clinical details were provided regarding how the pump came out of position and product wasn't returned, the root cause of the positioning issue could not be determined. vf/vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
7354 | 1220648-2025-30418 | 22529855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529855&pc=OZD | 07/01/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691891 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a pleural effusion, and right-sided chest tubes were placed. a total of 1.5 liters of blood were administered, along with two units of blood products. heparin had been administered prior to these events. additionally, the patient experienced mottling in the limb used for impella access, and initially, distal pulses could not be detected via doppler. however, pulses were eventually identified. the patient survived the explant procedure and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7355 | 1220648-2025-30419 | 22529917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22529917&pc=OZD | 06/30/2025 | 09/24/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647190 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp could not be complete as the pump was not able to move above the artery bifurcation. after many attempts, the impella procedure was aborted and percutaneous coronary intervention was completed without impella support. no harm was noted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition. related vessel size was less than 5mm and while inserting the impella device the motor was not able to move above the artery bifurcation after many attempts and hence the impella procedure was aborted. d6.b. explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30419.
7356 | 1220648-2025-30420 | 22530114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530114&pc=OZD | 07/01/2025 | 08/28/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620187 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella device was unable to generate flows above p4 without triggering suction alarms. the patient expired while on mechanical circulatory support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B.2 updated outcomes attributed to adverse event to not reflect other serious (important medical events) as death is the adverse event since manufacturer device report number 1220648-2025-30420 was submitted. h.6 updated health effect - clinical code to 4580 since manufacturer device report number 1220648-2025-30420 was submitted. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-30420 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: complaint device not returned for analyzing. data log reviewed, suction alarms observed during impella support above p2, low flow at p-9 observed. no positioning or placement signal issues observed. a purge spike seen. clinical confirmed suction alarms above p2, no echo details, no positioning performed. no enough clinical details provided to confirm root cause of the reported low flow issue. the cause of the low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided.
7357 | 1220648-2025-30421 | 22530168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530168&pc=OZD | 02/24/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589883 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient decompensated prior to the procedure and underwent successful cardiopulmonary resuscitation. an impella device was implanted for mechanical circulatory support, resulting in hemodynamic stabilization. during support, the patient experienced pulseless electrical activity (pea) and worsening cardiogenic shock. both events were considered to have a possible association with impella support. the patient remained on support for three days until a planned wean and explant. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7358 | 1220648-2025-30422 | 22530293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530293&pc=OZD | 06/30/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679391 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 43 KG | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the blue button lock on the repositioning unit of an implanted impella 5.5 was stuck in the open position. the surgeon placed a curved needle through the hole designed for the yellow pin and used it to spring the button into the locked position. impella support continues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7359 | 1220648-2025-30423 | 22530325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530325&pc=OZD | 06/23/2025 | 09/23/2025 |  | 07/17/2025 |  |  | 07/09/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661743 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 support for hemodynamic support. during support it was reported that the motor current rose to 32767 without warning and the pump stopped. restart was attempted according to the flowchart; however, the motor current remained as 0. extracorporeal membrane oxygenation was added urgently due to hemodynamic collapse, and the patient was taken to the operating room and impella 5.5 was reinserted from the same side. it was noted that the patient's condition, including hemodynamics, was being supported without any problems. patient survived. | [5 empty] | The investigation of pump stop has been completed. the data logs show that motor current, purge pressure and pump flow were stable before the pump stopped. the motor current (mc) spikes to 30,000 and triggered alarm 115. the returned product did not have any indications of biomaterial or any kinks seen in the catheter. electrical testing was performed and open lines were found. the red pump plug was opened and all 3 motor cable lines were torn on the patient cable side. based on the clinical details provided and the returned product analysis, the root cause of the pump stop is due to electrical open lines likely due to excessive force. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7360 | 1220648-2025-30424 | 22530476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530476&pc=OZD | 07/10/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/10/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026652791 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and shortly after the implant, the pump stopped. there were no abnormalities in the purge flow rate, purge pressure, or motor waveform before the pump was stopped. the pump was restarted and there was no increase in the current consumption value, and the operation was normal. support continued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7361 | 1220648-2025-30425 | 22530549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530549&pc=OZD | 07/02/2025 | 10/07/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657847 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | During preparation for implantation of an impella 5.5 the pump was found to have a malfunctioning optical signal sensor. a new impella was used to support the patient. no patient harm was reported. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. pump was returned for investigation. no physical damage was noted on the returned product. the pump passed the laser continuity test for the optical signal line, and all the 5s parameters were within specification during optical bench testing. data log shows the placement signal value below 0 with reported spike 3000 before pump placement. the cause of the optical signal issue was not determined since it could not be reproduced on the returned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7362 | 1220648-2025-30426 | 22530622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530622&pc=OZD | 07/02/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661142 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 70 KG | Non Hispanic | White | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was being implanted. after the single access sheath was placed, the physician attempted to pass a wire through the sheath but was unable to pass through the descending aorta. multiple attempts were made but unsuccessful. implantation was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7363 | 1220648-2025-30427 | 22530748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530748&pc=OZD | 06/30/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 06/30/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623500 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 88 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient had low platelets with suspected heparin-induced thrombocytopenia. platelets were monitored and transfused. additionally, the patient had ectopy and the pump was repositioned. the patient was started on amiodarone. support continued. | [5 empty] | Impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7364 | 1220648-2025-30428 | 22530782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530782&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687446 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a pump stop. the pump had ingested suture material. the pump was explanted and the patient survived. | [5 empty] | The investigation for the pump stop has been completed. no product was returned for evaluation. data logs were reviewed and real time (rt) log at time of pump stop showed loss of motor current pulsatility and pump flow prior to pump stop alarms occurring. pump was attempted to be restarted multiple times with immediate pump stops each time. clinical details noted that the when patient was being move to bed, a "impella stopped - high motor current" alarm was triggered. additionally, when pump was explanted, suture material was present around the impeller. the cause of the pump stop could not be definitively determined as no product was returned for analysis. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7365 | 1220648-2025-30429 | 22530839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530839&pc=OZD | 07/02/2025 | 08/30/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694426 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) undergoing a high-risk percutaneous coronary intervention to the left main/left anterior descending artery was implanted with an impella cp device for mechanical circulatory support. upon repositioning sheath insertion, the impella cp device advanced and at that time, the patient abruptly lifted a leg and the impella cp device came back past the valve. the physician was unable to properly reposition the impella cp device and, therefore. the decision was made to place a secondary impella. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was patient movement related since patient lifted leg abruptly and impella came back past the valve and it was unable to reposition properly. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7366 | 1220648-2025-30430 | 22530870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22530870&pc=OZD | 06/23/2025 | 08/26/2025 |  | 07/17/2025 |  |  | 06/23/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136137 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was being escalated from an impella cp to an impella 5.5. it was noted that during 5.5 support, the pump stopped. attempts were made to restart the pump with no success. alarm stated the pump was defective. the impella 5.5 was explanted and a new impella cp was placed. | [5 empty] | The investigation for the aic-pump stop has been completed. review of the data log confirms the complaint. at approximately 62 hours of run time there was a pump stop due to detection of excessive differences in three phase motor current. the pump was re-started, but it shut down again almost immediately, again due to differences in phase current. the pump was disconnected and then reconnected, but failed to re-start. two core dumps were noted in the days preceding the complaint, but attempts to recover the file were unsuccessful. the rt logs show two rapid increases in motor current a few minutes before the failure occurred, but otherwise seem to be unremarkable. the console was returned for evaluation. a test pump was connected to the console and flow was started at p-level 1. the p-level was then increased incrementally until reaching p-level 4. it was then allowed to run for approximately 4 hours. no pump stops or other issues occurred in this time. the optical bench 5s parameters were checked and all were within acceptable limits. the analog output of the ccd array showed minor distortion on the right side of the waveform where no fringe pattern occurs. the console was opened for inspection. there was a slight bend in the ribbon cable connecting the impellatronic to the 4-in-1, but this is unlikely to be problematic. no other issues were noted. the root cause for the aic-pump stop could not bet determined. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7367 | 1220648-2025-30431 | 22531278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531278&pc=OZD | 07/01/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669458 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female | 80 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella cp pump placed for the acute myocardial infarction/ cardiogenic shock. the site of the pump had bleeding and a hematoma developed. the bleeding prompted blood transfusion. the team also questions limb ischemia as there is less flow. additionally there was suction noted above p4. the patient survived and is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7368 | 1220648-2025-30432 | 22531356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531356&pc=OZD | 07/02/2025 | 08/30/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682070 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp device to a patient (unknown ethnicity) undergoing a high-risk percutaneous coronary intervention was unsuccessful. access was obtained; however, the device was unable to cross through the aorta due to calcium. consequently, the implant was aborted. there was no report or indication of adverse effects to the patient as a result of this event. subsequent treatment was unnoted. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of the pump was unable to cross aorta due to calcification.
7369 | 1220648-2025-30433 | 22531376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531376&pc=PYX | 07/01/2025 | 08/28/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659135 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 133 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella rp flex support and there were placement signal issues. there were placement signal not reliable alarms. support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The failure mode was changed from optical signal issue to placement signal issue because there was no issue found with the optical metrics via data log review and the issue was confirmed to be dp sensor related. placement signal issue: the pump was not returned. data log review showed the placement signal not reliable (psnr) (dp out of range) alarm # 1051 shown. motor current (mc) spikes without p-level change occur. alarm #1052 also occurs. ps drifts down before railing to 3000mmhg at 0018 on (b)(6) 2025, recovering 0630, drifting down more, and railing again at 0820. flows trend up during periods of ps trend down. no impact to 5s parameters. after 10 hours on support, mc begins periodically spiking from m 64 to 1452 ma. the cause of the placement signal issue was most likely sensor drift due to the ps trending down and the flows trending up at the same p-level.
7370 | 1220648-2025-30434 | 22531401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531401&pc=OZD | 07/03/2025 | 08/27/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687925 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pain | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown ethnicity) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and developed an access site bleed approximately three days after start of the support. the bleed also resulted in a hematoma at the site. a pressure dressing was applied, and heparin was placed on hold to manage the condition. it was additionally noted that the patient had some numbness / pain in the extremity. pain medication was given to the patient to assist with the sensation. pulses were verified to be present in the limb. the impella 5.5 mcs continued. approximately 10 days later the patient had bleeding at the insertion site and heparin was stopped and, however, scheduled to resume later that evening. subsequently, the patient was noted to be feeling better and doing well. the patient was sitting up in the chair and ambulating throughout the day, waiting for transplant and still on impella 5.5 mcs. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely anticoagulation management since hemostasis was achieved by with holding the heparin along with pressure dressing. the root cause of pain at insertion site could not be determined since the product was not returned and insufficient clinical details were provided. d6b explant date added.
7371 | 1220648-2025-30435 | 22531454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531454&pc=OZD | 07/01/2025 | 11/18/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1410103 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) had a controller error display on the screen upon rebooting. this issue occurred during preventative maintenance. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. data analysis: ¿ on the complaint date, (b)(6) 2025, the console displayed ¿controller error (opm comm stopped) [pump not connected] ¿ alarm,¿ event #1035. the optical parameters were abnormal, with the light at 0.08v and the snr at 1.4 (imc_logs_ic7110.pdf). this confirms the reported failure mode. note: before the complaint date, on (b)(6) 2025, when the console was used with the demo pump sn: (b)(6), it displayed the same event #1035 due to abnormal optical parameters. device analysis: this console was tested at fs, and the reported failure mode, ¿controller error,¿ was reproducible. root cause: the root cause of the controller error was a lamp malfunction.
7372 | 1220648-2025-30436 | 22531590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531590&pc=OZD | 06/22/2025 | 08/21/2025 |  | 07/17/2025 |  |  | 06/22/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685042 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a implantation of an impella cp could not be completed due to the patient's small arteries. an intra-aortic balloon pump was used instead. the patient survived. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the delivery issues has been completed. clinical details state that the sheath would not advance into artery due to size and anatomy. 10 dilator was largest access. impella pump was not opened. insertion aborted. based upon clinical details, the root cause of delivery issue is determined to be patient condition because the sheath would not advance into artery due to size and anatomy.
7373 | 1220648-2025-30437 | 22531598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531598&pc=OZD | 07/08/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/08/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616166 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon explant of an impella 5.5, a donut-shaped thrombus was found in the ascending aorta. the thrombus was surgically removed following discovery. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7374 | 1220648-2025-30438 | 22531650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531650&pc=OZD | 07/03/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a cp pump placed to support patient for acute myocardial infarction related cardiogenic shock (ami/cgs). the pump was delivered and placed within the left ventricle but pump position was suboptimal and within the papillary per the transesophageal echocardiography (tee) imaging. the pump was being repositioned and doing so the pump kinked and was removed/replaced. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7375 | 1220648-2025-30439 | 22531683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531683&pc=OZD | 07/03/2025 | 08/29/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687477 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The facility reported that a patient developed hemolysis during impella cp support. the patient's urine was red tinged and their free hemoglobin (pfhg) levels were 570. the pump was found to be too deep via echo and was repositioned at bedside to manage the condition. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation has been completed. no product was returned. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle. no other issues were found on the logs. support continued for 2 days after the repositioning performed to troubleshoot the hemolysis. product was not returned for review. based on clinical description & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position.
7376 | 1220648-2025-30440 | 22531809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22531809&pc=OZD | 07/10/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/10/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679208 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male | 47 KG | Non Hispanic | Asian | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella 5.5 via direct aortic access for a surgical patient. it was reported that when the impella 5.5 was removed, the metal part of the outflow cage and the cannula became detached. the cannula and the inflow cage remained inside the patient's body (inside the artificial blood vessel). after that, those were successfully removed. an x-ray showed no metal residue inside the body. there were no health hazard to the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7377 | 1220648-2025-30441 | 22532117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532117&pc=OZD | 07/02/2025 | 09/11/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1754606 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a battery failure occurred during regular check. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. e1 is unknown. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. data analysis: on boot up on 7/1, the user got a battery level low for voltage below 50 percent confirming the complaint. the power data from the complaint date showed cell imbalance with battery 2 which caused the battery pack¿s internal electronics to shut down. device analysis: batttest confirmed there was cell imbalance with battery 2 cell 4 and battery would not charge as a result. root cause: the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. d9 device returned to manufacturer date added.
7378 | 1220648-2025-30442 | 22532146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532146&pc=OZD | 07/12/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/14/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623357 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp to support the patient admitted in with ventricular fibrillation, implantable cardioverter-defibrillator activation, st-elevation myocardial infarction and cardiogenic shock. the pump was placed prior to imaging of the left ventricle. the team later (84 hours after implant) had a pump stopped due to a thrombus. it is thought that the thrombus at the apex of the heart was not identified before the impella was inserted, and that the thrombus was aspirated by impella. the pump was explanted and replaced by intra-aortic balloon pump. | [5 empty] | The impella device was discarded by the customer and therefore. an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7379 | 1220648-2025-30443 | 22532175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532175&pc=OZD | 07/02/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660943 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Intraoperative Cardiac Valve Injury | [11 empty] | Required Intervention | Disability | Hospitalization | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) was implanted with an impella 5.5 device for mechanical circulatory support and the next day after start of the support, the patient experienced multiple runs of ventricular tachycardia. an amiodarone drip was started. the patient remained sensitive to repositioning and would go into ventricular tachycardia any time after coughing or gagging. impella flows were adjusted each time the ventricular tachycardia occurred. approximately six days into support, it was noted that the patient's aortic valve leaflet was "frozen," and it is believed that this was due to intrathoracic pressure changes as a result of the patient's respiratory function. support was eventually weaned, and the impella 5.5 device was explanted with a survived outcome. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿.
7380 | 1220648-2025-30444 | 22532415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532415&pc=OZD | 07/01/2025 | 08/28/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670414 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 69 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with impella cp with acute myocardial infarction and coronary artery disease for support during a high risk cardiac procedure. during the procedure, the patient was noted to have multiple blockages and with the previous stents placed. the patient was on multiple pressors, intubated, and continued to decline. it was then decided to place the impella cp with the goal to leave in place to support the patient prior to proceeding with any further interventions. shortly after the repositioning sheath was advanced, the patient was noted to be in ventricular tachycardia and required two shocks per the defibrillator. the patient was given bolus and started on a drip at this time. the physician noted the color and warmth to the right foot and was unable to obtain doppler pulse. however, it was stated with current pressor support and decreased pulsatility due to amio bolus, there were satisfied with the right foot being warm. the patient had suction events and gross hemolysis of lab. echocardiogram measured the impella at 2.0 centimeters from atrioventricular annulus. the pump was repositioned due to the labs hemolyzing. after repositioning, the impella measured 3.6 centimeters. unfortunately, the patient continued to decompensate. continuous renal replacement therapy was attempted. however, it only ran for 30 minutes due to volume issues. fluids were given and bicarbonate drip was started. suction continued and p-level was dropped as needed. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis, low pump flow, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. hemolysis: data log review showed placement signal (ps) trending down over the course of the case. low flows noted and patient in suction conditions the majority of the case, independent of p-level. no motor current (mc) spikes outside of p-level changes. the cause of the hemolysis was most likely patient condition since patient declined for 3 days prior to implant and volume noted to be low with a downward trending ps. low pump flow: data log review showed suction events with lowering flows independent of p-level over the entire case runtime. no mc spikes noted with purge pressure decreasing slightly over the runtime and purge flow increasing slightly. the cause of the low pump flow was most likely patient condition related due to lack of blood volume to match support required of the pump. positioning issues: the cause of the positioning issues was not determined due to insufficient clinical details. vf/vt/af/at: the root cause of the vt was unable to be determined due to insufficient clinical details.
7381 | 1220648-2025-30445 | 22532438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532438&pc=OZD | 07/03/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657795 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient (pt) presented as an out of hospital cardiac arrest. downtime 15-20 minutes. return of spontaneous circulation achieved by emergency medical services. pt intubated due to being unresponsive. electrocardiogram showed st-segment elevation myocardial infarction. pt was transferred for left heart catheterization evaluation. in the cath lab, pt had heparin drip and dopamine drip infusing. the pt continued to be hypotensive after interventions. decision made to place impella cp. the impella was prepped and inserted per ifu. peel-away sheath removed, repo advanced, and site secured. dressing applied. during support it was noted that the pt had absent doppler pulse to distal extremity and was rapidly deteriorating having multiple runs of ventricular tachycardia (vt). the physician stated likely related to inferior myocardial infarction. pt went pulseless vt and advanced cardiac life support measures were initiated and given for several minutes. code ended and medical doctor pronounced time of death. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7382 | 1220648-2025-30446 | 22532561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532561&pc=OZD | 07/03/2025 | 10/07/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680828 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported patient had a 5.5 pump placed for support for multiple cardiac and systemic comorbidities inclusive of had a 5.5 pump support ongoing for the patient with multiple cardiac and systemic comorbidities inclusive of end stage renal disease on huntington disease, valve disease, and abdominal aortic aneurysm. the pump was supporting when there was elevated purge pressure that prompted the team to add tissue plasminogen activator to the purge. the pump remains on despite this. additionally there were ps issues observed that self resolved on the aic without any product replacement. patient survived | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. product was not returned for investigation. data logs were reviewed, and the controller alarm was confirmed. there were a couple other instances of this pattern in optical 5s parameters, however, none of which lasted long enough to trigger an alarm. the 5s parameters were stable throughout the case. upon reviewing data logs, it is apparent that no problem with the pump. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h.6. updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-30446. d.10. concomitant medical products and therapy dates updated. | The investigation for high purge pressure and limb ischemia has been completed. data logs were reviewed and the high purge pressure event was confirmed. purge pressure began to rise gradually on (b)(6) 2025 from the 500's to the 900's over the course of 5 hours. there were no motor current (mc) spikes that correlated with the onset. once the pressure was in the 900's it plateaued for 41 hours prior to another 12 hour gradual rise that triggered the first high purge pressure alarm on (b)(6) 2025. product was not returned for investigation. the gradual rise in purge pressure without a mc spike is a signature of biomaterial buildup. this is supported by the fact that the purge pressure gradually returned to specification after tpa was added to the purge. for these reasons, the cause of the high purge pressure was most likely biomaterial buildup. upon further data log review, an abnormality was noted in purge pressure that resolved after the purge cassette was swapped out to troubleshoot high purge pressure. data logs showed that there were cyclic dips of ~400 mmhg in purge pressure that occurred once every cycle of the purge cassette piston which is indicative of priming issue. in this instance, they occurred every time the piston would pull back/reset. however, without additional clinical details, the root cause of the limb ischemia and priming issue could not be determined. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7383 | 1220648-2025-30447 | 22532577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532577&pc=OZD | 07/01/2025 | 09/17/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683208 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported that a patient was on impella 5.5 support. during support, high purge pressure was observed. a lysis protocol was initiated and tissue plasminogen activator was added to the purge solution. the patient remained on impella support. |  |  |  |  | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. high purge pressure: data logs were reviewed and on day of reported issue, a step up in motor current with a small spike in pump speed and an increase in purge pressure of 2 minutes and drop in purge flow was observed. pump flow became fully saturated after step up in motor current. purge pressure remained high for approximately 3 hours before slowly resolving. clinical details noted that no kinks were present in purge line and lysis was added to the purge and resolved the issue. the cause of the high purge pressure was most likely biomaterial ingestion based on returned data log analysis and clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. b1 adverse event selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30447. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30447. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30447. h1 type of event was erroneously selected on the initial manufacturer device report number 1220648-2025-30447. h6 health effect - clinical code 4582 and health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2025-30447. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7384 | 1220648-2025-30448 | 22532690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532690&pc=PYX | 07/03/2025 | 08/29/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659371 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported decision for impella rp flex was made and placed using best practices. during support, it was noted of possible entrapment of clot in impella upon arrival to icu. during transfer to ct flows restored to pump form 0.6 to 3.0 and patients hemodynamics stabilized. the rp was explanted and the patient continued on extracorporeal membrane oxygenation support. patient survived. | [5 empty] | The investigation has been completed. data logs confirmed the failure mode & clinical narrative. logs showed after pump started & run for about 40 minutes, mc spiked followed low flow that triggered auto suction response & suction alarms. this event remained for 4 hours then flow back to the normal range. near the end of the case, mc spiked again followed low flow & remained to the rest of the case. visual inspection found biomaterial inside if cage. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion since the data meet the low flow pattern & biomaterial was found on the pump. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7385 | 1220648-2025-30449 | 22532791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532791&pc=OZD | 07/01/2025 | 08/21/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026691750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced bleeding from multiple sites, accompanied by a drop in hemoglobin and hematocrit levels, which necessitated transfusion of packed red blood cells and heparin was discontinued. additionally, the patient developed hemolysis but no further interventions were noted to address this complication. the impella cp was ultimately weaned and explanted successfully without complication. the patients outcome was stable. | [5 empty] | The investigation for access site bleeding and hemolysis has been completed. the device was not returned for evaluation. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. no motor current (mc) spikes outside p-level changes. the root cause of access site bleeding was most likely anticoagulation management related since the bleeding ceased with stopping anticoagulation per clinical information. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7386 | 1220648-2025-30450 | 22532811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532811&pc=OZD | 07/03/2025 | 09/17/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656889 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 119 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the "placement signal unreliable" alarm occurred on an impella cp implanted in a patient. the pump was explanted and replaced with another impella. the patient survived. | [5 empty] | The investigation of optical signal issue has been completed. the data logs show that immediately after plugging in the pump there is raw optical and placement signal (ps) that are within respected range, however the signal not reliable (snr)drops from about 10000 to about 500. shortly after the snr, contrast and gain begin to trend down with fluctuations in the signal until the ps goes out of range and psnr alarms are triggered, this occurs right as the pump is fully ramped up to p-9. it remains this way for about 9 minutes and then the parameters return to be within range but then the pump is then explanted. the returned product passed laser continuity, had nothing on the sensor face, all 5s parameters were within specification, reproduced no alarms and was able to hold pressure in the uson test. the other case on the console for the pump that was used right after did not present with similar behavior. due to the lack of evidence on what was going on clinically and that the issue was not able to be reproduced, the root cause of the optical signal issue cannot be fully determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7387 | 1220648-2025-30451 | 22532918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22532918&pc=OZD | 07/03/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577132 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after impella cp was successfully implanted there was no flow to the impella leg. a distal perfusion catheter was placed to re-establish flow. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7388 | 1220648-2025-30453 | 22533040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22533040&pc=OZD | 07/03/2025 | 09/02/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  | 07/31/2025 |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437195 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 60 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to inability to pass through the subclavian artery. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since the impella couldn't pass through the subclavian artery and the tortuous anatomy.
7389 | 1220648-2025-30454 | 22533080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22533080&pc=OZD | 07/01/2025 | 11/25/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1202891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 86 KG | Non Hispanic | White | Malfunction | No | Yes | Overheating of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller had a high battery temperature. this issue occurred during support. no further information was provided. | [5 empty] | The investigation has been completed. data analysis: console logs from the complaint date confirm occurrence of two ¿battery temperature high¿ alarms (#49, due to battery temperature over 60degc) that self-resolved in exactly 10 seconds. there were no preceding battery temperature alarms #48 (battery temperature is between 50degc and 60degc.) power and battery data embedded in ltlog files showed stable battery temperature throughout the case, not exceeding 30degc. seemingly false "battery temperature high" alarms occurred while aic was operating on batteries, and ac cord had been unplugged 17 minutes and 43 minutes earlier. device analysis: the console (b)(6) was returned to field service for further investigation. the reported issue was not reproduced during testing, but was most likely caused by momentary corrupt data communicated over smbus between battery, smart charger (on pbm), and pbm microcontroller, which falsely triggered the alarm. it had not been determined which specific component was responsible for the malfunction but is most likely limited to pbm and batteries. root cause: the root cause of the battery temperature high alarm (#49, due to battery temperature over 60degc) was most likely due to a momentary communication glitch between the batteries and pbm, where a single sample from battery data (in this case value of battery temperature) was corrupt. however the exact component has not been determined. d9 updated. h3 updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7390 | 1220648-2025-30455 | 22533951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22533951&pc=OZD | 07/01/2025 | 08/28/2025 |  | 07/17/2025 |  |  | 07/01/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026690277 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was undergoing implantation of an impella 5.5 for mechanical circulatory support. during the procedure, complications arose while delivering the pump through the aortic arch. the clinical team removed the pump and guidewire and discovered that the wire had sheared within the pump outlet. all components were successfully removed and replaced with a new pump. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The investigation of delivery issue has been completed. the impella device was not returned for evaluation. delivery issue : the root cause of the delivery issue was most likely patient condition related since the pump delivery was complicated in the aortic arch. delivery issue the root cause of the delivery issue was most likely patient condition related, the team removed the pump and wire and found the wire to have sheared in the pump outlet. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7391 | 1220648-2025-30456 | 22534520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22534520&pc=OZD | 07/03/2025 | 09/02/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664092 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [5 empty] | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp suffered ventricular tachycardia. the patient was cannulated for va ecmo. when the patient changed position the pump came out of the left ventricle. the pump was explanted and a new pump was implanted in its place. the patient survived. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. positioning: based on clinical description, the root cause of positioning issue was most likely patient movement since the pump was out of left ventricle while turning the patient. vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30456. d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30456. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
7392 | 1220648-2025-30457 | 22534636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22534636&pc=OZD | 07/03/2025 | 08/27/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp could not be delivered via the left femoral artery due to calcification in the ostial iliac, and a hematoma developed. the pump was delivered via the right femoral artery instead. andioseal and perclose were used to close both sites. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and delivery issues has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue was most likely patient condition since the impella cp unable to be delivered in the left femoral artery due to calcification in the ostial iliac. access site bleeding: the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
7393 | 1220648-2025-30458 | 22534787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22534787&pc=OZD | 07/03/2025 | 09/03/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  | 08/31/2025 |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393748 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 73 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 suffered blood loss during graft anastomosis and was transfused two units of packed red blood cells. additionally, the patient had substantial bleeding from a fasciotomy site on the calf requiring volume replacement and pressors. the patient remains on impella support. | The medical safety officer reviewed and determined the impella 5.5 device event is a serious injury. cascading events starting with the impella cp and given the details there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome captured under manufacturer device report 1220648-2025-30482. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The medical safety officer reviewed and determined the impella 5.5 device event is a serious injury. cascading events starting with the impella cp and given the details there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome captured under manufacturer device report 1220648-2025-30482. | Access site bleeding: the root cause of access site bleeding was most likely use issue since non-abiomed product was used which is not recommended per the instructions for use. vascular damage peripheral vessel: the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. additional information received reporting the patient outcome of death. the following fields were updated: b2 death and date of death added. b5 patient outcome added. d6b explant date. h1 type of event updated. h6 health effect - impact code 1802 added.
7394 | 1220648-2025-30459 | 22535059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22535059&pc=PYX | 04/25/2025 | 11/04/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 67 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA RP FLEX SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a placement signal unreliable alarm occurred for about 10 minutes and resolved after shifting the patient's position. the automate impella controller was exchanged for another. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: on the complaint date of (b)(6) 2025, approximately 30 minutes after the case was initiated with pump sn: (b)(6) the console displayed ¿placement signal not reliable (rp dual sensor, optical signal snr out of range for 2 minutes) ¿ alarm,¿ event (b)(6). this issue was resolved in 10 minutes. this confirms the reported failure mode. note: following this case, a psnr alarm also occurred during several other cases performed with pumps sn: (b)(6) and sn: (b)(6). device analysis: this console was tested after arriving at fs. the analog output from the optical bench was abnormal, with a spike. unable to remove debris from the front panel optical connector. root cause: the root cause of the unreliable placement signal was due to an optical bench malfunction. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h6 updated.
7395 | 1220648-2025-30460 | 22535237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22535237&pc=OZD | 07/03/2025 | 09/02/2025 |  | 07/17/2025 |  |  | 07/03/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 137 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death |  |  |  |  | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed increasing runs of ventricular tachycardia through the night. the patient was cannulated for ecmo to resolve the issue. the patient developed a large clot in the aortic root and on top of impella cannula and the impella had low pump flow. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." | The investigation of vf/vt/af/at, thrombus, and low pump flow has been completed. the impella device was not returned for evaluation. data review: logs showed pump ran without issue for 14 hours & 30 minutes, then there are suction alarms and the p-level is turned to p-2 with low flow occurred with position in ventricle alarms and position unknown alarms. there is an increase in purge pressure and drop in purge flow around this time as well, which slowly resolves. the low flow remained to the rest of the case. low flow: there are a few potential contributing factors to the low flow such as the patient decompensating, potential improper position, and thrombus observed on transesophageal echocardiogram (tee) and seen interacting with the purge system. however without the returned product or more clinical information, a single root cause cannot be determined. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. b2 required intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30460.
7396 | 1220648-2025-30461 | 22535244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22535244&pc=OZD | 07/02/2025 | 07/17/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670245 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that upon explantation of an impella cp, the patient experienced access site bleeding. the use of perclose and angioseal were ineffective. surgery was consulted as a hematoma had formed. the patient survived and was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7397 | 1220648-2025-30462 | 22535302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22535302&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/17/2025 |  |  | 07/02/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161032 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preventative maintenance, an automated impella controller powered on with a system self check failure message. there was no patient involvement at the time of the noted issue. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the aic self check error has been completed. data logs were reviewed which found during initial boot-up, the system rebooted automatically due to a continuous powermod_1 communication issue. during the reboot, because of the persistent powermod_1 communication issue, the console displayed the 'system self check failed' screen. this confirms the reported failure mode. the console was returned for evaluation. during bootup, the console rebooted automatically and displayed the "system self check failure" screen. visual inspection confirmed pbm corrosion. the root cause of the controller error was determined to be pbm corrosion resulting from electrolyte leakage caused by battery failure alarm super caps.
7398 | 1220648-2025-30463 | 22535430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22535430&pc=OZD | 07/03/2025 | 09/11/2025 |  | 07/17/2025 |  |  | 07/07/2025 | 07/17/2025 |  |  |  |  | 07/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP O |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102528 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a pump support ongoing when there was a self resolving controller yellow alarms. the pump remained on despite this and the console too. this occurred on day 10 of the 5.5 support, the day of explant. no patient harm reported due to the issue. | [5 empty] | The investigation into the reported issue has been completed. console logs from the reported day of event are consistent with complaint and show an controller error alarm triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented with negative values due to the error. the console was sent back to field service for further investigation. the reported failure mode was not reproduced during testing. the controller error was due to an optical bench communication protocol anomaly, resulting in misalignment of data communication packets. the aic operated as designed. the optical bench was evaluated for 5s parameters and the analog signals were checked for distortion; no issues were found. additionally, it was observed that the purge disc flag was broken. the purge pressure sensor accuracy was checked and found to be within the specification as per the pm procedure (b)(4). the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
7399 | 1220648-2025-30373 | 22518612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22518612&pc=OZD | 07/04/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 07/04/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688195 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. at the start of the procedure, the impella was positioned 1.3 cm below the aortic valve; the physician subsequently advanced the device to 3.5 cm. the patient tolerated the procedure well. a bedside post-closure technique was attempted but was unsuccessful. to maintain hemostasis, a 14 french sheath was inserted. during support, the patient was noted to be mildly hypertensive. during flushing of the sheath, thrombus was observed, and the patient developed clinical signs of acute limb ischemia. the patient was taken emergently to the cardiovascular operating room for surgical thrombectomy and open patch repair of the affected artery. the impella cp device was subsequently weaned and explanted successfully without complication. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. a5 section is unknown. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7400 | 1220648-2025-30340 | 22518631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22518631&pc=OZD | 06/29/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/29/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655548 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 114 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 anticontamination sleeve had a tear. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7401 | 1220648-2025-30374 | 22518677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22518677&pc=OZD | 06/27/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/27/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601479 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. during ongoing support, the patient developed a hematoma at the vascular access site. heparin was withheld, and a hematoma evacuation procedure was performed. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7402 | 1220648-2025-30375 | 22518746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22518746&pc=OZD | 06/27/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/27/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601246 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a large hematoma was observed at the left groin return cannula site. multiple units of blood were transfused, and anticoagulation therapy was held due to the bleeding. the patient survived the explant procedure and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7403 | 1220648-2025-30376 | 22518903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22518903&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/27/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679721 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Arrhythmia | Stroke/CVA | Hemolysis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the patient was taken to have a cat scan and found that they had endured multiple strokes. | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella was noted to be oriented toward the mitral valve, with the inlet positioned beneath the papillary muscle. this malposition resulted in severe mitral regurgitation (mr). initial attempts to reposition the device were unsuccessful. on a second attempt, the clinical staff successfully pulled back and torqued the pump; however, the anticontamination sleeve fractured during the maneuver. at that time, the patient also exhibited signs of hemolysis, which was resolved following successful repositioning of the device. despite the correction in pump positioning, the mitral regurgitation persisted. on day nine of support, the patient experienced an episode of ventricular tachycardia (vt) during efforts to wean from extracorporeal membrane oxygenation. the patient remained on impella support following the event. |  |  |  |  | The investigation of positioning issues, product damage (sterile sleeve damage), vf/vt/af/at, hemolysis, and stroke/cva has been completed. positioning issues: the cause of the positioning issues could not be definitively determined as limited clinical details were provided on when and how the pump came out of position. hemolysis: clinical details noted that once pump was repositioned, hemolysis resolved. no additional details were provided on when hemolysis was first noted or how long it lasted before resolving. the cause of the hemolysis was most likely improper positioning based on clinical details noting that hemolysis resolved once pump was out of improper position towards the mitral valve and within papillary muscle. product damage (sterile sleeve damage): clinical details noted that sterile sleeve was torn when attempting to reposition pump out of papillary muscle. the cause of the product damage was most likely a use related issue as the damage occurred when manipulating pump for repositioning. vf/vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. d6b explant date added b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30376. h6 health effect - clinical code 1770 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30376. | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
7404 | 1220648-2025-30377 | 22518942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22518942&pc=OZD | 06/28/2025 | 08/13/2025 |  | 07/16/2025 |  |  | 06/28/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687356 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was initially supported with an impella cp device and subsequently escalated to an impella 5.5 for mechanical circulatory support. during 5.5 support, the device began triggering sudden "purge blocked" alarms. the team initiated tissue plasminogen activator in the purge line. despite this intervention, there has been no significant improvement in purge pressures. the patient remained on escalating vasopressor support. the patient's condition appears to be worsening, with multiple episodes of ventricular tachycardia and ventricular fibrillation requiring defibrillation. the patient was receiving amiodarone infusion but cannot tolerate a bolus due to bradycardia after conversion. ultimately, care was withdrawn, and the patient expired. | [5 empty] | The investigation of high purge pressure, vf/vt/af/at, and vascular damage (peripheral vessel) has been completed. the impella device was not returned for evaluation. high purge pressure (hpp): clinical details state that there were purge system blocked/purge pressure high alarms on (b)(6) 2025. no kinks/blockages were observed. purge solution was d5w with bicarb. lysis solution was added to purge solution and infusion continued for 30 hours before purge flow was restored and pressure normalized. data log review showed small temporary increase in motor current later followed by a 13 minute purge pressure (pp) increase. hpp was sustained for about 33 hours before resolving over a 6 hour period. there were no motor current spikes prior to the pp rise. there was a long cassette change but it occurred after the rise in pp. no product was returned. the cause of the high purge pressure was most likely biomaterial buildup as there was a gradual rise in pp over 13 minutes and no motor current spikes based on data logs. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. vascular damage (peripheral vessel): clinical details state that the patient experienced increased bleeding from oral and nasal cavities and a drop in hemoglobin requiring 4 units of packed red blood cells (prbc). team suspected retroperitoneal bleed from extracorporeal membrane oxygenation (ecmo) cannulas. it was noted that the heparin drip was contributing to bleeding risk. data logs are not applicable and product was not returned. the cause of the vascular damage was not determined since insufficient clinical details were provided. h6 health effect - clinical code 1888 and health effect - impact code 4643 was inadvertently omitted on the initial manufacturer device report number. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7405 | 1220648-2025-30378 | 22519149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519149&pc=OZD | 06/29/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/29/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557918 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had a 5.5 pump support placed for the patient with cgs and cardiomyopathy. the pump was supporting when after 18 days the pump was found to be having ps issues and the alarm was disabled. additionally there was ventricular tachycardia that was new to the patient and possibly due to pump positioning. the care was withdrawn and patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | Optical signal issue: clinical details note that the patient was having constant position in ventricle alarms on (b)(6) 2025. echo verified very good placement. no other details are known around this time. data log review shows a decrease in minimum snr around the time of the first impella in ventricle alarms with placement signal waveforms looking ventricular and motor current losing pulsatility. the raw optical signal did not show any clear issues with the optical sensor. product was not returned. the cause of the optical signal issue was not determined since product was not returned and insufficient clinical details were provided. vf/vt/af/at: data logs are not applicable and product was not returned. i the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30378.
7406 | 1220648-2025-30379 | 22519209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519209&pc=OZD | 06/29/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/29/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694704 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that distal pulses on both feet were difficult to find before bringing the patient to the cath lab for impella cp placement. impella leg appeared to be cooler. doppler pulses could not be found. ultrasound looked like maybe continuous or slow flow. the team will monitor. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7407 | 1220648-2025-30380 | 22519342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519342&pc=OZD | 06/30/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642529 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Ischemia | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced difficult placement as the pump could not be advanced beyond a large abdominal aneurysm. a cook 18fr sheath was used to deliver the pump. additionally, the patient's limb was perfused after low pulsatility was observed. later, the family withdrew care and the patient expired. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7408 | 1220648-2025-30381 | 22519444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519444&pc=OZD | 07/01/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 07/01/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678139 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced oozing at the insertion site. systemic heparin was withheld. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7409 | 1220648-2025-30382 | 22519587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519587&pc=OZD | 06/30/2025 | 08/22/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672003 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 90 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced atrial fibrillation. the patient was treated with vasopressors. in response to heparin-induced thrombocytopenia the patient's anticoagulation was changed to bival. additionally, the patient was oozing from all sites. one unit of blood products were transfused. impella support continues. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation has been completed. no product was returned. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the root cause of the access site bleeding issue was moat likely patient condition related since patient is oozing from all sites. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided. h6 investigation conclusions updated with additional information. h10 additional manufacture narrative updated with additional information.
7410 | 1220648-2025-30383 | 22519787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519787&pc=OZD | 06/26/2025 | 08/21/2025 |  | 07/16/2025 |  |  | 06/26/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 65 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Hemorrhagic Stroke | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient implanted with an impella cp experienced clotted distal perfusion catheters and the decision was made to explant the pump. systemic heparin had been withheld due to bleeding from the mouth. four days after explantation the patient exhibited neurological changes and a head ct confirmed an intraparenchymal hemorrhage with interventricular and subarachnoid extension resulting in herniation. care was withdrawn and the patient expired. | [5 empty] | The investigation for stroke/cva, mouth bleeding, and limb ischemia has been completed. the device was not returned for evaluation. without additional clinical details and the device evaluation, the root cause of mouth bleeding, stroke, nor the limb ischemia can be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7411 | 1220648-2025-30384 | 22519839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519839&pc=OZD | 06/26/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/26/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677535 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient became increasingly septic from pneumonia during impella cp support. the patient's condition was monitored throughout the course of support. upon planned explant, a thrombus was discovered prompting a thrombectomy and arthrotomy patch. the patient was stable following the intervention. | [5 empty] | Patient information is unknown. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7412 | 1220648-2025-30385 | 22519975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22519975&pc=OZD | 07/01/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 07/01/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660974 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a cp pump placed to support for non-st-segment elevation myocardial infarction. the pump was supporting for 1.5 hours for the percutaneous coronary intervention and explanted. when explanted there were bleeding issues when perclose could not achieve hemostasis and so contralateral leg balloon angioplasty was done for hemostasis. there was not distal perfusion observed and patient got 1 unit of blood transfused. the patient survived the harm. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7413 | 1220648-2025-30386 | 22520025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22520025&pc=OZD | 06/30/2025 | 09/15/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518916 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AIC SN (B)(6). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was being placed when placement signal unreliable alarm occurred. no kinks in the cable were seen. there was also an unsuccessful attempt to insert the impella connector more firmly into the automated impella controller. the impella cp was replaced. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the impella device and data logs were received for investigation. there was no problem was found with the pump and after cross functional review it is apparent the console was the potential cause of the issue. please refer to manufacturer device report number 1220648-2025-30387 for an investigation into the console. d9 device returned to manufacturer date added.
7414 | 1220648-2025-30387 | 22520609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22520609&pc=OZD | 06/30/2025 | 09/05/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541270 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support the supporting automated impella controller (aic) had a controller error alarm. the aic was exchanged. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue and automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: on 6/30, cp pump 521773 was connected to ic8449. throughout usage, there were placement signal not reliable alarms despite the pump being exchanged. additionally, there were placement signal not reliable alarms on the previous case on 5/23. the real time (rt) logs confirm the unreliable placement signal as well as the other 5s signals. device analysis: the failure mode was reproduced during testing. additionally, the fringe pattern on the ccd array with and without a pump was found to have irregularity on the left side that would have affected optical performance. root cause: the cause of the placement signal issues that triggered placement signal not reliable alarms was a defective optical bench. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30387.
7415 | 1220648-2025-30388 | 22520745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22520745&pc=OZD | 06/30/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1639305 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an error alarm occurred on an automated impella controller (aic) supporting an implanted impella cp. the aorta and left ventricle waveforms weren¿t visible. the nurse turned flow control to p-0 then to p-2 and all waveforms reappeared. flow control was turned up one level at a time until resumed at p-5. the controller error alarm was resolved. no patient harm was reported and the procedure was successfully completed. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench were represented as values due to the crc error. console was tested and the issue was not reproduced, however the cause of issue was identified. the controller error and loss of placement signal were due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7416 | 1220648-2025-30389 | 22520846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22520846&pc=OZD | 06/28/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142508 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preparation of the impella automated impella controller (aic) a controller error alarm occurred. troubleshooting was done and the issue was resolved. there was no reported involvement. | [5 empty] | The reported event did not involve patient use. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the console was booted up with controller error event #24 for loss of communication between the keyboard and the 4-in-1. the console was power cycled and the console booted up without error. this controller error did not occur on any of the other boot ups with this console. device analysis: the console was returned and was power cycled 100 times but the failure mode did not reproduce. the communication that failed on that boot up would have been between the qt1060 microcontroller on the kbd and an msc1210 on the 4-in-1. root cause: the root cause of the controller error could not be determined as the keyboard communication error on boot up was not reproduced. h8 usage of device was erroneously entered on the selected manufacturer device report number: 1220648-2025-30389.
7417 | 1220648-2025-30390 | 22521091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521091&pc=OZD | 06/22/2025 | 08/20/2025 |  | 07/16/2025 |  |  | 06/22/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673802 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 75 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Hematuria | Ischemia Stroke | Thrombosis/Thrombus | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced an arrhythmic event, hematuria, thrombus and a stroke. the patient would eventually pass away. the patient presented to the emergency room with chest pain and shortness of breath. the patient underwent a cardiac catheterization which revealed multivessel disease. the left anterior descending artery was the culprit and stented. the decision was made to place an impella cp and swan which was successfully place. due to the patient's ejection fraction percentage, the plan was to transfer the patient to a different hospital for advance heart failure work up. approximately nine hours later, there were suction alarms. the patient was moving and was tachycardic. the impella cp was repositioned initially measuring 5.2 centimeters with suction alarms and after repositioning measuring 3.3 centimeters within mid left ventricle and free of structures. the performance level was at p-7. it was noted that urine was previously pink tinged but cleared to yellow. the patient was accepted at the higher level of care hospital and planned for transfer, which was completed. the following day it was noted that the patient had undergone a computed tomography angiogram which was positive for a cerebral vascular accident. the patient was unable to move the left side. the care team is unclear as to where the clot originated and was noted that the patient did not receive pre-impella cp placement angiogram. eventually, the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Stroke: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. hemolysis: the pump was not returned. data log review showed suction alarms mostly resolving after a placement signal spike, around the time pump noted to be repositioned. the cause of the hemolysis was most likely improper positioning of the device as repositioning the pump from 5.2 to 3.3cm resolved the issue.
7418 | 1220648-2025-30391 | 22521322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521322&pc=OZD | 06/23/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/23/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1838437 |  | 0042-0040-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 and approximately 10 days into the impella support, the automated impella controller (aic) spontaneously shut down and rebooted itself. the aic was exchanged with no report or indication of harm to the patient as a result of the event. the patient was reported to be in stable condition following the event and continued on impella mcs. | [5 empty] | The investigation for the aic (console) shutdown/restart issue has been completed. data logs were reviewed which showed the console had uninitiated 5.5 pump. the console unexpectedly rebooted with hemodynamic support returning 25 seconds later automatically. the aic was switched out. the console was returned for testing but the unexpected reboot was not reproduced. the cause of the unexpected reboot could not be determined since the failure mode was not reproduced. | D.9. date device returned to manufacturer is unknown. the associate clinical consultant reported it was returned weeks ago. the automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7419 | 1220648-2025-30392 | 22521414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521414&pc=OZD | 06/30/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679735 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 79 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced access site bleeding. the patient was transfused multiple units of packed red blood cells. impella support continues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7420 | 1220648-2025-30393 | 22521571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521571&pc=OZD | 06/27/2025 | 08/18/2025 |  | 07/16/2025 |  |  | 06/27/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664127 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 138 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | Fluid/Blood Leak |  |  |  |  | Low Blood Pressure/ Hypotension | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with a impella cp for mechanical circulatory support. during transport, a registered nurse was instructed to apply manual pressure to the access site due to a developing hematoma and was kneeling on the patient's bed. subsequently, the patient developed hypotension, accompanied by both a suction alarm and a purge system open alarm on the impella console. troubleshooting revealed that the purge system side arm had become lodged between the bed railing, as reported by the nurse, and had fractured into two pieces. a 5 french micropuncture sheath was inserted into the damaged purge tubing to restore flow to the catheter. during this maneuver, the impella catheter migrated proximally and retracted into the aorta. despite repositioning attempts, the patient's condition deteriorated and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: cleaning, ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump, positioning issues, and low pump flow has been completed. the impella device was not returned for evaluation. the cause of the issue was a leak from the sidearm but the exact location of the leak was unknown. logs show pump was set to p9 but could not reach desired flows and was at p2 to p6 range. there were no motor current spikes. there were no positioning issues or alarms before low pump flow onset but an 'impella position in ventricle' that soon resolved occurred. no other devices or patient condition issues were noted. the root cause of the positioning issue and low pump flow was not determined as limited clinical information was provided.
7421 | 1220648-2025-30394 | 22521821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521821&pc=OZD | 06/30/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662777 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematoma. additionally, low pump flow was observed and the pump was explanted. upon explant, a clot was observed on the outflow area of the impella. | [5 empty] | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, low pump flow, and thrombus has been completed. the pump was returned for investigation. no visual abnormalities were observed on the returned product. the cannula was inspected, and a significant amount of biomaterial material was observed on the impella, which could not be removed. the pump was operated as-is, with the biomaterial material in the impella and in a bucket of water, and low pump flow was observed. the data log also confirmed a motor current spike at the end of the case, followed by low pump flow and suction alarms. the cause of the low pump flow issue was biomaterial ingestion. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the cause of the access site bleeding issue was not determined without sufficient clinical information provided. b1. product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30394. h6. medical device problem code 1248 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30394.
7422 | 1220648-2025-30395 | 22521852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521852&pc=OZD | 06/29/2025 | 08/27/2025 |  | 07/16/2025 |  |  | 06/29/2025 | 07/16/2025 |  |  | 08/31/2025 |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398351 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced intermittent atrial fibrillation during impella cp support. support was scheduled to continue, however, the pump was inadvertently pulled partially across the valve when the patient was repositioned. attempts to readvance the pump were unsuccessful and the decision was made to remove the device. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. vf/vt: the root cause of the atrial fibrillation was unable to be determined due to insufficient clinical details. positioning issues: the cause of the positioning issue was most likely due to patient movement per clinical details noting that pump was pulled out accidentally when patient was being repositioned in bed.
7423 | 1220648-2025-30396 | 22521859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22521859&pc=OZD | 07/01/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 07/01/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026687546 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 support went into ventricular tachycardia while walking, requiring amio, lido, and cardioversion. impella was reduced to p6. echo showed impella in good position. patient continued on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7424 | 1220648-2025-30397 | 22522042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522042&pc=OZD | 07/01/2025 | 08/04/2025 |  | 07/16/2025 |  |  | 07/01/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 95 KG |  |  | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Arrhythmia | Fever | Sepsis | Vascular Dissection | Ectopic Heartbeat | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella support continued to have runs of ectopy. they were trying different drips to manage heart rate and blood pressure. additionally the patient had a new fever, cultures were sent. the patient was taken to cath lab to reassess coronaries. intravascular ultrasound and pressure wire to left man artery/left anterior descending coronary artery showed to be stable so decision to leave as is. cp was weaned, however it was noted that the impella cp was removed from right femoral artery using dry closure technique with 8mm x40 balloon. impella was pulled back and balloon inflated. two perclosures were used. final pictures of femoral artery showed dissection in common femoral requiring balloon tamponade. vascular consulted, stabilized with balloon and no further intervention required. patient survived. | The impella cp was removed due to left heart recovery. there were no issues with the device. cultures all came back negative and it was decided that the fever was most likely from the acute phase of cardiogenic shock. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.4 weight was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30397. b.5 additional information was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30397. e.1 phone number was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30397. the email reported on initial manufacturer device report number 1220648-2025-30397 was incorrect and should not have been reported. h.6 codes 1721, 2067 and 4114 were reported incorrectly on the initial manufacturer device report number 1220648-2025-30397. new codes were added. h.10 updated additional device manufacturer narrative and instructions for use since the initial manufacturer device report number 1220648-2025-30397 was reported due to updated codes.
7425 | 1220648-2025-30398 | 22522136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522136&pc=OZD | 06/29/2025 | 08/26/2025 |  | 07/16/2025 |  |  | 06/29/2025 | 07/16/2025 |  |  | 07/31/2025 |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374968 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a left atrial appendage thrombus. it is unknown if this thrombus was related to the impella. additionally, the patient had a hematoma at the access site requiring return to the operating room for hematoma evacuation. support was continued and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was most likely patient movement since the bleed started occurring after the patient was moved from or and brought to icu. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7426 | 1220648-2025-30399 | 22522160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522160&pc=OZD | 07/01/2025 | 08/28/2025 |  | 07/16/2025 |  |  | 07/01/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615240 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump could not be advanced through the patient's artery due to existing calcification. a dye shot was utilized to locate the calcification and a shockwave was used, but the pump was still unable to pass into the left ventricle. the implant was subsequently aborted due to the noted resistance. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. a.5 unknown. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since impella backloaded and was unable to pass due to calcification in artery.
7427 | 1220648-2025-30400 | 22522464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522464&pc=OZD | 06/29/2025 | 08/27/2025 |  | 07/16/2025 |  |  | 06/29/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662603 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 54 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a patient admitted in for a high-risk percutaneous coronary intervention. the support was ongoing when there were oozing and suction issues treated by the infusion of red blood cells. the pump remains on for support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding - major: oozing at impella site observed. unknown how and why it was oozing. the root cause of the bleeding issue cannot be determined due to insufficient clinical information. low pump flow: data log reviewed: no motor current (mc) spike outside of p-level changes observed. suction alarms occurred started on 28/6 and ended on 29/6 after blood given. no positioning issue or placement signal issue seen. purge spikes seen. the cause of low pump flow issue most likely was patient condition related since suctions resolved after blood administrated.
7428 | 1220648-2025-30401 | 22522683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522683&pc=OZD | 06/25/2025 | 11/05/2025 |  | 07/16/2025 |  |  | 06/30/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1203510 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during use the automated impella controller went into flight mode. it was not connected and sometimes the curves were gone. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | D9: additional information, the device was returned. h6: type of investigation: added codes. this report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed, the data logs were reviewed and verified the reported placement signal issue. after replacing microsd card with a known-good one, the device worked to specification. intermittent losses of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. a review of the device history record (dhr) for the suspect device and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. the root cause is related to the software. all pertinent information available to abiomed has been submitted at this time. b5: this complaint is not reportable for the flight mode issue, this issue is reportable for the placement signal issue. d6a and d7b: these dates are incorrect and have been omitted. the controller was not implanted. h6: medical device problem code was corrected to device sensing issue (2917). h6: investigation findings and conclusions: recoded for the results of the placement signal investigation. h10: corrected to provide information regarding the investigation of the placement signal issue.
7429 | 1220648-2025-30402 | 22522901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522901&pc=OZD | 04/30/2024 | 07/16/2025 |  | 07/16/2025 |  |  | 04/30/2024 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1448350 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella support, a member of the medical team accidentally broke the dust cover of the automated impella controller while walking. there were no patient harm as a result of the broken dust cover. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. dust cover was confirmed to be broken upon return of the console. the broken dust cover was use related.
7430 | 1220648-2025-30403 | 22522952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22522952&pc=OZD | 03/15/2024 | 07/16/2025 |  | 07/16/2025 |  |  | 05/20/2024 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in shock vitals due to heart failure associated with thyroid crisis was implanted with an impella cp device and placed on extracorporeal membrane oxygenation for mechanical circulatory support. during support, purge flow abnormalities had been reported: ¿although there was no problem with the purge pressure, the purge flow rate fluctuated between 0ml/hr and 12ml/hr, which was considered abnormal. after replacing the purge housing with a new one, it started operating normally, so it was thought that there was a problem with the purge housing¿. this complaint was related to a clinical procedure. there were no adverse events caused by the reported automated impella controller issue. the patient recovered heart function and was successfully weaned and removed from impella cp support. | [5 empty] | The investigation has been completed. no product was returned. console logs from the reported day of event are consistent with complaint and show some abnormal purge flow behavior (sudden purge flow drops to 0), while purge cassette was used. also numerous ¿imc-ahp error¿ messages were recorded in imc logs, suggesting possible purge cassette communication issues, or incorrect purge cassette installation. investigation into the purge cassette concluded the issue to be most likely use-related. console ic6642 had been requested but had not been returned therefore its possible contribution to the purge issues could not be fully investigated. the cause of the issue was not determined since product (console) was not returned.
7431 | 1220648-2025-30404 | 22523118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22523118&pc=OZD | 06/23/2024 | 07/16/2025 |  | 07/16/2025 |  |  | 06/23/2024 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support the automated impella controller (aic) experienced an error as no curves or values were displayed. the aic was exchanged to resolve the issue. there was no reported patient harm. | [5 empty] | The device was returned for analysis and the device logs were also reviewed. the root cause of the controller error and loss of placement signal was an optical bench firmware communication protocol anomaly, resulting in the misalignment of data communication packets received. the aic's software operated as designed. a functional check was performed before returning the device to service.
7432 | 1220648-2025-30405 | 22523317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22523317&pc=OZD | 09/09/2024 | 07/16/2025 |  | 07/16/2025 |  |  | 07/02/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507302 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient presented for coronary artery bypass grafting. impella 5.5 was inserted without complication, flows were gradually increased while coming off bypass. multiple blood products were given during procedure. epicardial v wires and chest tubes x4 were left in place prior to chest closure. 3-point fixation was applied in operating room. later it was reported from impact study that the patient had acute kidney injury that is likely related to the bleeding that was associated with the impella (possible) and the surgery (definite). the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7433 | 1220648-2025-30406 | 22523527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22523527&pc=OZD | 07/02/2025 | 07/16/2025 |  | 07/16/2025 |  |  | 07/02/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594690 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced vascular complications that resulted in hemorrhage, cardiac arrest, and death. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7434 | 1220648-2025-30407 | 22523567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22523567&pc=OZD | 07/02/2025 | 09/23/2025 |  | 07/16/2025 |  |  | 07/02/2025 | 07/16/2025 |  |  | 07/31/2025 |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431652 |  | 1000361 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  |  |  | Death | Yes | No | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring repositioning of the pump. additionally, the patient also had a pulmonary thrombus. it is unknown if the thrombus was impella related. furthermore, the purge cassette had a leak. the cassette was replaced and no further issues were reported. impella support continues. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The product was not returned for investigation. a data log was not provided or retrieved from remote server. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. the root cause of the thrombus was unable to be determined due to insufficient clinical details. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30407. b5 on the initial report should not have included purge cassette leak as this not reportable 1220648-2025-30407. b5 the patient outcome of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30407. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-30407. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30407. h6 medical device problem code was erroneously entered on the initial manufacturer device report number 1220648-2025-30407.
7435 | 1220648-2025-30408 | 22523708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22523708&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/16/2025 |  |  | 07/02/2025 | 07/16/2025 |  |  |  |  | 07/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363622 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that once connected to an impella pump, the automated impella controller (aic) displayed a "controller error, switch to backup controller" error message. the aic was replaced and there was no patient harm reported. | [5 empty] | The investigation for the aic controller error has been completed. console logs from the reported day of event are consistent with complaint and show controller error alarm (#1023, loss of communication with pbm1), while pump was running, approx. 2 hours from case start. pump was disconnected and console was shut down. console logs also show multiple pbm1 communication issues since console start, not manifested in any alarms. upon next console power-cycle (presumably for testing in a non-clinical setting) five days later, console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the console was returned for evaluation and during evaluation, the issue was reproduced, as the console presented with system self-check failure, due to loss of communication with pbm1. inspection of the console¿s electronics revealed corrosion on the pbm. this issue, manifested in clinical and test setting, was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. device testing also revealed frequent interruptions of data streaming, due to corrupt microsd card in (b)(4) (unrelated to this complaint). the root cause of the controller error alarm was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. h6 impact code changed to 4629.h8 changed to reuse. | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7436 | 1220648-2025-30326 | 22505554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505554&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/15/2025 |  |  | 06/26/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603430 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 70 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Atrial Fibrillation | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced an arrhythmic event, hematuria and low pump flow leading to replacement of the device. the patient underwent a second left heart catheterization (lhc) as the initial lhc with intra-aortic balloon pump placement report was not sent with transfer work up. thrombus was noted in the left anterior descending artery and potentially in the circumflex artery. the decision was made to place an impella cp with removal of the intra-aortic balloon pump. the following day, the patient was in atrial fibrillation and urine output was yellow. however, the urine output was previously red in the morning's report. vasopressin was started and discussion was made regarding options of optimizing the impella support and weaning drips. impella cp mcs continued. however, the following afternoon, the patient was being transferred from the bed to a stretcher via emergency medical services for transfer. at this time, the impella cp flows decreased, impella in ventricle alarm appeared and there was also a spike in motor current with a dampened waveform. the purge flow and purge pressure remained unchanged. the patient was on heparin purge solution, heparin drip, and cangrelor drip. an echocardiogram at bedside was completed, and the impella cp was readjusted. alarms did not resolve. the patient was subsequently taken to the operating room for a new impella cp placement via the left femoral artery. the patient was reported to be in stable condition following the event. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of low pump flow, optical signal issue, vf/vt/af/at, and hemolysis has been completed. low pump flow: returned product was investigated, and the only visual abnormality was a severe kink in the cannula. however, it is believed that this likely occurred during the return process because the pigtail end of the cannula was stuck inside the knot of the plastic bag. additionally, there was no report from the field of a cannula kink upon explant. there was no biomaterial on returned product, which was to be expected, considering that pump flows returned to specification before explant. the pump was connected to a console and run in a bucket, where low flows did not reproduce. based on the signatures noted in data logs, the root cause of the low pump flows was most likely biomaterial ingestion. optical signal issue: returned product was investigated and there were no relevant visual abnormalities. the pump was connected to a console and all 5s parameters were within specification. additionally, the pump passed the laser continuity test. finally, the pump was run in a bucket for a period of time, and all optical signal metrics were stable during this run. based on the signatures seen in data logs and all optical signal metrics being within specification during the case and on returned product, the root cause of the optical signal issue was most likely a false alarm due to biomaterial. vf/vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. hemolysis: there was no biomaterial or damage to the pump of relevance. as mentioned earlier, there was a kink in the cannula, however, that kink likely occurred during product return because the cannula was returned caught in the knot of the plastic bag, and clinical did not report noting a cannula kink on explant. finally, the pump was hemolysis tested and passed with an average mih of 9.21. since clinical details were limited, there were no abnormalities noted in data logs, and returned product passed hemolysis testing, the root cause of the hemolysis could not be determined. thrombus: a thrombus failure mode was added because the investigation into low pump flows and optical signal issue found that they were due to biomaterial ingestion. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7437 | 1220648-2025-30327 | 22505641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505641&pc=OZD | 06/27/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666814 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock (cgs) was implanted with an impella cp for mechanical circulatory support. the complaint reported that a patient presented for primary percutaneous coronary intervention yesterday and was explanted in the cardiac catheterization lab. the patient went to the floor and decompensated overnight. the impella cp was implanted due to low cardiac output/cardiac index and cgs. during support it was noted that the patient had multiple rounds of ventricular tachycardia (vt) requiring shocks. ultimately, the family withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7438 | 1220648-2025-30328 | 22505705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505705&pc=OZD | 06/28/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/28/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599692 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 101 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced malpositioning of the pump and oozing at the access site. pressure was held on the access site and the patient was administered lidocaine and epinephrine. the pump was successfully repositioned. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. no product was returned for evaluation. access site bleeding - the cause of the access site bleeding could not be definitively determined as limited clinical details were provided. positioning issues: the cause of the positioning issues was most likely due to patient movement as pump was found out of position after patient transferred hospitals per clinical details.
7439 | 1220648-2025-30329 | 22505792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505792&pc=OZD | 06/27/2025 | 09/23/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644983 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during a subsequent impella cp exchange, vascular access was lost. manual compression and surgical repair were required to achieve hemostasis, and the re-implantation of the impella cp was aborted. further evaluation with echocardiography and right heart catheterization revealed normalized left ventricular function and reduced left heart filling pressures. based on these findings, the clinical team elected not to re-attempt impella cp placement and instead continued right ventricular support alone using the existing impella rp flex. | [5 empty] | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be patient condition because the insertion was aborted since the pump was unable to pass through vasculature. vascular damage - peripheral vessel: access to the vessel was lost during removal of the .035 j wire and manual compression and surgical repairs were performed to achieve hemostasis. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention.. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | After further evaluation the device may have caused or contributed to a serious injury (vascular damage - peripheral vessel). the following fields were corrected based on this information manufacturer device report number 1220648-2025-30329. b1 adverse event was inadvertently omitted on the initial manufacturer device report. b2 was inadvertently omitted on the initial manufacturer device report. h1 type of event was erroneously selected on the initial manufacturer device report. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report. h6 health effect - impact code 4624 was inadvertently omitted on the initial manufacturer device report.
7440 | 1220648-2025-30330 | 22505860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505860&pc=OZD | 06/27/2025 | 08/25/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662796 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 111 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced significant oozing from multiple sites on the body, including all puncture sites. in response, the heparin dose was reduced, and three units of blood products were administered. the patient survived the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of the access site bleeding and vascular damage issue was most likely patient condition related since the patient on impella 5.5 support developed bleeding from all puncture sites. h6 type of investigation updated with additional information. h.6. investigation findings updated with additional information. h.6. investigation conclusions updated with additional information. h.10. additional manufacture narrative updated with additional information.
7441 | 1220648-2025-30331 | 22505892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505892&pc=OZD | 06/26/2025 | 08/25/2025 |  | 07/15/2025 |  |  | 06/26/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646046 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and after the implant, the patient had ventricular fibrillation (vf). three shocks were given to the patient and support continued. additionally, when the peel away was pulled part way out of the femoral, the sheath started to split. this resulted in significant bleeding from the access site. the repositioning sheath was placed and pressure was held. a femstop was then placed. support continued. patient survived. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, clinical details note that the patient vfib arrested shortly after cp placement and was given three shocks. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. clinical details note that the medical doctor (md) started splitting the peel-away sheath before the sheath was completely out of the femoral artery, which resulted in significant bleeding at the groin site. after repositioning sheath was placed, pressure was held on groin site for at least 30 minutes before transferring patient to icu. the cause of the sheath split issue was not determined since insufficient clinical details and product not returned. md began to remove sheath from patient but sheath peeled apart prior to being completely out of the femoral artery. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.5 unknown.
7442 | 1220648-2025-30332 | 22505960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22505960&pc=OZD | 06/26/2025 | 09/09/2025 |  | 07/15/2025 |  |  | 06/26/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655753 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed to support percutaneous coronary intervention (pci) procedure. during pci the impella suddenly ran at only 0.9 liters. after pci the impella was explanted. patient survived. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. data log review: data logs show motor current, motor speed, and pump flow level drop, which was followed by suction alarms. no sharp spikes in motor current, but an overall decrease in value. p-level auto-adjusted for the drops at time of suction alarm. flows did not resolve. no variability in purge pressure. low pump flow: the product was returned for investigation and biomaterial was observed around the impeller. data logs show motor current, motor speed, and pump flow level drop followed by suction alarms. p-level auto-adjusted for the drops at time of suction alarm. the cause of the low pump flow was most likely biomaterial ingestion, as physical product investigation showed biomaterial around the impeller.
7443 | 1220648-2025-30323 | 22506019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506019&pc=OZD | 05/05/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/28/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660300 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female | 59 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right femoral artery access site occlusion with a "probable" relationship to the impella device used: patient presenting for worsening right lower extremity pain on ambulation. patient denied any rest pain but said that after walks even 10m there is severe right lower extremity pain, the patient underwent successful right femoral thrombectomy and endarterectomy bovine pericardial patch. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7444 | 1220648-2025-30333 | 22506062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506062&pc=OZD | 06/27/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  | 03/31/2025 |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327783 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of an impella cp device was unsuccessful for a patient in acute myocardial infarction/cardiogenic shock requiring mechanical circulatory support. the physician attempted more than once to load the impella cp device on a 0.035 guidewire not realizing a 0.018 guidewire was available in the package for use and eventually aborted the implant. thereafter, the patient was treated with mediation and transported alive, sedated and ventilated to another hospital. it was noted that there was harm to the patient. | [5 empty] | Additional information noted the implant attempt should have been made the different guidewire and lack of experience, not adhering to the implantation protocol as trained was the reason for the event. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7445 | 1220648-2025-30295 | 22506119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506119&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/15/2025 |  |  | 06/26/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622842 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female | 69 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after one day of impella cp support, the patient began to develop symptoms of hemolysis with hematuria. flows were adjusted and the pump was repositioned. however, the issue worsened. in addition, the patient also then lost pulses below their knee in the impella leg. due to the noted conditions, the decision was made to remove the pump and escalate to impella 5.5 support. both the ischemia and hematuria resolved following impella cp explant. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and limb ischemia could not be determined.
7446 | 1220648-2025-30334 | 22506122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506122&pc=OZD | 07/02/2025 | 09/08/2025 |  | 07/15/2025 |  |  | 07/03/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026633475 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that coronary artery bypass grafting was performed after intra-aortic balloon pump support, and extracorporeal membrane oxygenation/impella was performed due to postoperative cardiopulmonary withdrawal difficulties. it was reported that the support was being reduced in preparation for the weaning, however, the pump was accidentally removed when changing the patient's position. as the patient still required support, the condition was urgently escalated to 5.5. patient survived | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues have been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most like use related as the pump was accidentally pulled out when changing position.
7447 | 1220648-2025-30335 | 22506224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506224&pc=OZD | 06/27/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685147 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a couple hours after the implant it was noted the patient was in atrial fibrillation. further, the team was unable to doppler pulses; the patient had low pulsatility. positioning was check and noted the impella free floating in mid left ventricular, free of the mitral valve. the patient was prepped and transferred to another facility and cannulated for extracorporeal membrane oxygenation (ecmo) upon arrival. hemodynamics stabilized upon initiation; however, the patient was severely hypoperfused on the impella cp prior to cannulation with mean arterial pressure of 40 on impella p-8 support level. the left femoral impella cp access was utilized for ecmo arterial cannula due to the occluded right superficial femoral artery stent and distal perfusion sheath placed to perfused to the left leg around the ecmo cannula. the impella cp device was explanted and escalation to an impella 5.5 was discussed as a de-escalation strategy from va ecmo if the patient recovered. the patient was noted to have survived the impella cp device explant; however, the status of the patient following the event was unknown. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7448 | 1220648-2025-30336 | 22506249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506249&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689149 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention (pci)_was implanted with an impella cp device for mechanical circulatory support. after the pci to the left anterior descending, circumflex and right coronary arteries, the physician made the decision to explant the impella cp device. the first perclose did not take; therefore, the physician aborted the perclose and placed the short peel-away sheath. thereafter the patient was taken to recovery until activated clotting time was appropriate. it was noted approximately one hour later, the sheath was removed, manual pressure was applied, and the site was stable. there was no harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of removal issues has been completed. no product was returned for evaluation. clinical details noted the perclose was attempted to be used after the pump was explanted and failed, the physician opted to insert a short 14fr introducer and wait until activated clotting time was appropriate before removing. after an hour, the introducer was removed and hemostasis was achieved with manual pressure. no problem was noted with the impella device per the physician. the cause of the removal issue was most likely a use related issue as the perclose sutures were not applied per instructions for use indicating that perclose should be applied prior to impella insertion.
7449 | 1220648-2025-30337 | 22506325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506325&pc=PYX | 06/15/2025 | 08/13/2025 |  | 07/15/2025 |  |  | 06/15/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025487332 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in cardiogenic shock from acute myocardial infarction was eventually placed on bipella (impella 5.5 and impella rp flex) for mechanical circulatory support (mcs) and experienced many complications. one in particular, hemolysis, requiring intervention was associated with the impella rp flex. the patient was initially implanted with an impella cp for mcs. after approximately eight hours, the patient was escalated to an impella 5.5 for continuation in therapy. however, the first impella 5.5 implanted malfunctioned and therefore was explanted and replaced. the replacement impella 5.5 was positioned under transesophageal echocardiographic guidance and ramped achieving a flow of 4.3 liters per minute. the patient was sent to the unit intubated and on impella 5.5 mcs. the following morning, the patient remained intubated. however, did not require sedation at this time. an echocardiogram was completed, and ejection fraction was noted as 15-20% with right ventricular systolic function reduced. the plan was to take the patient to the catheterization lab for left and right heart catheterization. overnight, the patient had required additional blood products, volume resuscitation, and multiple amps of bicarbonate. the patient was in atrial fibrillation with rapid ventricular response and sustained ventricular tachycardia. the cardiologist gave multiple amiodarone boluses with amiodarone drip, and attempted beta blockers intravenous. however, rate control was unable to be achieved. the physician also attempted to cardiovert the patient two times. the patient was eventually cardioverted back into sinus rhythm and an impella rp flex was implanted (patient now on bipella support). it was noted that since placing the impella rp flex and cardioverting, the patient has been able to wean pressor requirements. the patient was continued on bipella support per the noted plan. at this time, the patient was sedated along with intubation. the patient was noted to move all extremities with purposeful movement noted to bilateral upper extremities. throughout the day, the patient was noted to have ectopy and later that evening had two runs of ventricular tachycardia and was able to be shocked to normal sinus rhythm each time. the impella 5.5 measured 4.4 centimeters from the aortic valve annulus. the left ventricle chamber appeared narrow, and therefore, the patient was repositioned in the bed to ensure no contact with the wall. the inlet tip appeared close to the lateral wall but touching at the time of the echocardiogram. when the patient was flat on his back, the impella 5.5 inlet was in mid-ventricular space. the echocardiogram technician had completed the exam when the patient went into ventricular tachycardia and then ventricular fibrillation. the patient was shocked and the patient remained in ventricular fibrillation despite the second shock. medication and fluids were administered and the patient was shocked back to sinus rhythm. the impella 5.5 was slightly rotated and locked at 34.5 centimeters marking, measuring 4.2 centimeters from the aortic valve annulus. the patient was in sinus rhythm with no further runs of ventricular tachycardia. however, ectopy was still noted at times. the following day, it was noted the patient¿s heart rate was irregular. there was ectopy at times. there were no more ventricular tachycardia and ventricular fibrillation issues after the event overnight. urine was green in the foley tubing/bag and urine output has been less than 30 milliliters per hour to no urine for the last five hours. bumex drip was started overnight. creatinine continued increasing and a discussion was made for possibly starting continuous renal replacement therapy (which was started). labs were drawn and were hemolyzed. the impella 5.5 and impella rp flex were weaned by one performance level. labs were drawn again. the next day, the patient remained on mechanical ventilation and sedated with dialysis running. dressings for both the impella 5.5 and impella rp flex were noted clean, dry, and intact. a chest x-ray showed the impella rp flex was in the proper position above the pulmonic valve. labs came back hemolyzed. overnight, no ectopy and one unit of blood and 500 of albumin were given. echocardiogram this day in the morning showed left ventricular ejection fraction of around 10%. the plan of care was to continue the patient on bipella (impella 5.5 and impella rp flex) support. however subsequently, the patient¿s family withdrew care. the impella 5.5 and impella rp flex devices were both turned off along with intravenous drips and mechanical ventilation. the patient thereafter expired. | [5 empty] | Hemolysis: no product was returned for analysis. hemolyzed labs were reported. clinical confirmed good pump positioning. the data logs show several motor current spikes before reported hemolysis. there was a motor current spike while the pump was running at p8 with speed deviation and shift in placement signal. the pattern identified is most likely as a result of biomaterial interaction with the impeller. there were also 'impella stopped - motor current high' alarms and a small rise in purge pressure with a drop in purge flows. placement signal trended upward for the remainder of support. the root cause of the hemolysis was most likely as a result of biomaterial as evidenced by the ingestion pattern observed in the data logs thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. temporary pump stop: no product was returned for analysis. the logs show several motor current spikes with 'impella stopped - motor current high' alarms. before 'impella stopped - motor current high' alarms there was a motor current spike with speed deviation characteristic of an ingestion. placement signal trended upward for the remainder of support; there was also a small rise in purge pressure with a drop in purge flows. the root cause of the temporary pump stop was most likely as a result of biomaterial as evidenced by the ingestion pattern observed in the data logs and multiple motor current spikes. | Medical safety officer reviewed the event and determined there is not enough evidence to exclude the bipella impella 5.5 (serial number: (b)(6) as an associated factor in the patient's outcome. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. the impella rp flex is one of three reports associated with the patient's implant experience. refer to the following: manufacturing report number: 1220648-2025-29934 for the report on the initially implanted impella 5.5 device serial number: (b)(6) that malfunctioned. refer to manufacturing report number: 1220648-2025-30262 for the report on the bipella impella 5.5 device serial number: (b)(6) (associated with the demise outcome).
7450 | 1220648-2025-30338 | 22506492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506492&pc=OZD | 06/29/2025 | 08/25/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684319 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 failed as the pump kinked in response to the patient anatomy. the second impella 5.5 was attempted but also failed to advance into position. pump delivery was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of delivery issue is determined to be patient condition because it was difficult to pass the left subclavian stenosis and the cannula got kinked during implantation. later they tried with a different pump and had the same difficulty while inserting and the case was aborted. the root cause of cannula kink was determined to be patient condition because pump was unable to pass past laca for left subclavian stenosis and cannula was kinked during proximal to outflow cage during the process. h6 medical device problem code inadvertently omitted on the initial manufacturer device report number.
7451 | 1220648-2025-30339 | 22506534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506534&pc=OZD | 06/27/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532998 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced oozing from the access site. blood products were administered and the site was redressed. following clinical deterioration, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7452 | 1220648-2025-30341 | 22506669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506669&pc=OZD | 06/29/2025 | 08/25/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655002 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 112 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump stopped unexpectedly roughly 14 hours into patient support. a high motor current was noted leading up to the failure. the pump was subsequently explanted and va ecmo was on standby for support if needed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation has been completed. per the clinical investigation, after 14 hours on support, pump stop occurred with noted high mc. the pump was not returned. data log review showed a mc spike occur 2 hours prior to pump stop. following the mc spike, mc trended upward with flows and ps trending down. the cause of the pump stop was most likely biomaterial due to data log review showing an mc spike flowed with increasing mc without p-level changes along with lowering flows and ps.
7453 | 1220648-2025-30342 | 22506705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506705&pc=PYX | 06/29/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026651003 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella rp flex support and during support, flows dramatically decreased and the pump was thought to have a clot. at this time, the patient was turned and a ¿placement signal not reliable¿ alarm appeared. labs and imaging were sent in and the patient was found to have hemolysis. the pump was explanted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of hemolysis, low pump flow, thrombus, and optical signal issue has been completed. the impella device was not returned for evaluation. hemolysis: data log review showed suction occur early on in the case. motor current (mc) spike greater than 1600 ma occurs at 10:40 on 26 jun 25. also, there was again a small mc spike and drop in flows. prior to the mc spike, mc was shifting down while placement signal (ps) trending up on 29 jun 25 relative to hemolysis event reported. this failure type indicates biomaterial ingestion. the cause of the hemolysis was most likely biomaterial ingestion based on log analysis and clinical review. low pump flow: this failure type indicates biomaterial ingestion. the cause of the low pump flow was most likely biomaterial ingestion based on log analysis and clinical review. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. optical signal issue: the cause of the optical signal issue was most likely biomaterial interaction with sensor based on log analysis and clinical review.
7454 | 1220648-2025-30343 | 22506832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506832&pc=OZD | 06/29/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685030 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that following implant, the impella cp pump became stuck in the sub valvular apparatus. multiple attempts to reposition the device were unsuccessful and the staff opted to continue with support despite the pigtail being stuck. the pump was weaned overnight and removed the following day. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7455 | 1220648-2025-30344 | 22506874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506874&pc=OZD | 06/27/2025 | 08/25/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with acute myocardial infarction and cardiogenic shock was implanted with a second impella cp for mechanical circulatory support. during support, the patient remained in refractory shock and ventricular fibrillation arrest requiring cardiopulmonary resuscitation for approximately 45 minutes. given the ongoing hemodynamic instability, the physician decided to initiate bipella support. percutaneous coronary intervention was performed, and the patient remained stable throughout the procedure. in the intensive care unit, despite repositioning under echocardiographic guidance, the patient continued to exhibit low flow. a chest x-ray revealed a kink in the proximal cannula near the outflow. further evaluation with echocardiography and right heart catheterization showed that the left ventricle had normal function and that left heart pressures were decreased. based on these findings, the physician decided to explant the impella cp and continue with right ventricular support only. the patient survived the event. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. clinical reported low pump flows which were sustained after repositioning efforts. echo cxr revealed a kink in the proximal cannula near the outflow. data logs were reviewed, no mc spike observed. suction occurred above p2. no placement signal issue or purge spike seen. drop in flow observed directly after case initiated at p-9. the cause of the low pump flow issue was most likely due to the cannula kink as evidenced by clinical confirmation of kink in the cannula seen via echo and data logs confirming low flow and suction. pump was placed in a hurry during active cpr. the cause of cannula kink most likely was use related due to improper technique. h6 type of investigation updated with additional information. h6 investigation findings updated with additional information. h6 investigation conclusions updated with additional information. h10 additional manufacture narrative updated with additional information. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7456 | 1220648-2025-30345 | 22506956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22506956&pc=OZD | 06/29/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a cp pump was placed but the access site was oozing and the team infused 2 units of blood and placed a femostop. the next day the cp was explanted and escalated to the 5.5 pump. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7457 | 1220648-2025-30346 | 22507062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507062&pc=OZD | 06/29/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674782 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that following impella cp implant, the staff was unable to find distal pulse on the impella side. it was noted that both extremities became cool to the touch coinciding with some discoloration. plans were made to initiate heparin and to place a perfusion catheter to address the condition. the attempt to place a perfusion cannula ultimately failed, however, blood flow to the impella leg was later confirmed via ultrasound. after three days of support, the pump was explanted. following explant, a hematoma developed at the access site. manual pressure, one perclose, angioseal, and femostop were utilized to achieve hemostasis. the following day, the patient showed right sided hemiparesis and global aphasia. a ct finding discovered a prior occipital infarct of unknown timing. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7458 | 1220648-2025-30347 | 22507301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507301&pc=OZD | 06/28/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/28/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665453 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a hematoma at their impella access site during support. the pump was removed in response to the bleed and hemostasis was achieved. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7459 | 1220648-2025-30348 | 22507449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507449&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614416 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was turned on in preparation for impella placement when "system self-check failed, changed console immediately" occurred. the aic was exchanged and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: on the initial attempt to boot the console it did not detect the service dongle, thus preventing the ability to enable ethernet. the console automatically rebooted once and then booted to the blue ¿system self-check failure¿ screen. the console was opened and it was discovered that the cable connecting the service dongle connector to the 4-in-1 was disconnected. it was reconnected and the console was then successfully booted into service mode allowing for download of the logs. the root cause of the aic issue was a defective pbm board. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7460 | 1220648-2025-30349 | 22507485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507485&pc=OZD | 06/27/2025 | 09/17/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100209 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an error message when booting the automated impella controller, switched to backup console, no further problems or errors. there was no patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. based on the data log and device analysis the cause of the "system self-check error" was due to a malfunction with the purge unit driver. d9 device returned to manufacturer date added h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-30349.
7461 | 1220648-2025-30350 | 22507511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507511&pc=PYX | 06/28/2025 | 09/10/2025 |  | 07/15/2025 |  |  | 06/28/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502670 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 88 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon set-up and priming of the impella rp flex, the pump issued an ¿optical sensor system error¿ screen. the pump was exchanged for a new one. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d4 serial number was erroneously entered on the initial manufacturer device report number. d4 lot number was erroneously entered on the initial manufacturer device report number. d4 udi was erroneously entered on the initial manufacturer device report number. d9 device available for evaluation was updated with additional information. d9 device returned to manufacture was updated with additional information. d9 date device returned to manufacturer was updated with additional information. g4 pma/510k number updated as it was erroneously entered on the initial manufacturer device report number. h3 device evaluated by manufacture was updated with additional information. h5 labeled for single use was updated as it was erroneously entered on the initial manufacturer device report number. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation was completed. clinical stated that upon set-up and priming sequence for rp flex, pump issued a ¿optical sensor system error¿ screen. a second pump was used and implanted without issue. data log review showed psnr (opm invalid bench signal) alarm was triggered which is potentially related to the optical sensor. logs returned (5 rt and 5lt) with some running no more than 2 seconds with abnormal 5s parameters indicate an issue with how the pump is seated in the aic. rt log shows snr under 15, contrast under 16%, gain nearing 2.6, and lamp at 100%, all variables that match known air gap pattern. the pump was returned and inspected. no physical abnormalities were noted with the pump. psnr was unable to be reproduced. no abnormalities found with the optical 5s parameters. the red lumen was still intact as the pump was never implanted. uson testing was completed where external pressures of 50, 100, 150, 200mmhg yielded cvp readings of 60, 112, 165, and 219, respectively. the cause of the optical signal issue was most likely an air gap based on data log review and the returned pump having no abnormalities.
7462 | 1220648-2025-30351 | 22507564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507564&pc=OZD | 07/06/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 07/06/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660924 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient experienced an episode of ventricular tachycardia while maintaining a palpable pulse and remaining alert. the patient received one ampule of intravenous lidocaine and a single defibrillation shock, which successfully restored normal sinus rhythm. an intravenous amiodarone infusion was initiated. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7463 | 1220648-2025-30352 | 22507697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507697&pc=OZD | 06/26/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/26/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603483 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support had mild ischemia in lower extremity. the patient had rhabdomyolosis from lower leg ischemia. computed tomography angiography showed some distal flow down leg. however, a distal perfusion sheath was placed. the impella cp was successfully weaned and explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7464 | 1035166-2025-00045 | 22507904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507904&pc=DYB | 04/27/2025 | 10/15/2025 |  | 07/15/2025 |  |  | 06/24/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | INTRODUCER, CATHETER |  |  | ABIOMED 23F INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | DP-21849 |  | 0052-3006 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | NA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Fluid/Blood Leak | [6 empty] | Cardiogenic Shock | Ischemia Stroke | [11 empty] | Death | [5 empty] | FLOLAN. | LVOPHED/NOREPINEPHRINE. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We have been notified of a complaint involving an introducer (material#: 0052-3006, lot#: dp-21849). it was reported that the axillary sheath for 5.5 was damaged out of the package, found during the procedure. once placed into the 10mm platinum graft and two graft locks placed there was significant bleeding from perforation points down the sides of the hub. new axillary sheath placed; issue resolved. the patient received 3 units of prbc (packed red blood cells). additionally, the patients' outputs dropped overnight, and suction alarm occurred on p8. lv (left ventricle) seemed to be sucked down. down to p7 and flolan started to help right to left flow. (flolan is a brand name for epoprostenol, to treat pulmonary hypertension, acts as vasodilator). patient recovered and was eventually weaned off levo (levophed/norepinephrine, vasoconstrictor). the patient had a right mca (middle cerebral artery) stroke. care was ultimately withdrawn, and the patient expired. no medical intervention or additional surgical procedure was required. there was no procedure delay, the device was not used in a tortuous path when the event occurred. no specific technique was used for insertion; all best practices were followed. the customer reported, patient expired on (b)(6) 2025. the physician reported the cause of death, as care was withdrawn. yes, the instructions for use were followed before the procedure. it was reported the axillary sheath was damaged out of the package as being "split not at tip". we assume it was damaged out of package. the perforation for the peel-away portion was leaking without any significant force. after two graft locks were placed at the appropriate location, clamp was released to de-air sheath; this is when leaking occurred and significant bleeding was noted at the perforations for peel-away around the hub. operator placed thumbs over sides until new axillary sheath was placed. no excessive force was used. this additional device was used in the procedure impella 5. with smartassist s2 set, us (1000100), lot#: 2025501622, sn#: (b)(6). the 0052-3006, lot#: dp-21849 is available for evaluation. the lot number of the repositioning axillary sheath for 5.5 is unknown. no pictures, images, or videos available of the malfunction. the device is available for evaluation. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 1: b4,d4,g3,g6,h2,h3,h6 and h11. correction : d4 lot number corrected. one 23f abiomed axillary introducer sheath was returned under rma# 1202112728 without the dilator. there were no other accessories. blood was found on and inside the sheath. according to the event description summary, the axillary sheath was damaged out of the package as being "split not at tip". during the decontamination process, the sheath leaked immediately from a small crack in the hub when water was flushed through the sideport tubing assembly with a syringe. returned device analysis revealed the sheath leaked through a hairline crack in the hub directly under the hemostatic valve. the origin of the crack in the hub is not known; it may have been caused by the user encountering resistance during the procedure or by factors related to the patient's anatomy. a leak test is performed by qa 100% before shipping to the customer. in conclusion, the review of the dhr did not identify any manufacturing anomalies of this type and passed all final testing prior to shipment. the stated failure of the returned product was confirmed by our investigation; however, the root cause of the failure cannot be determined. per procedure aabiomed introducer sheath in-process and final inspection visual inspection - sample size: 100% with the naked eye at a distance of 12" to 18", verify the assembled sheath matches the drawing. ensure parts are free of kinks, cracks, splits, sinks, excessive flash, loose/embedded fm, or any other damages. verify split caps are properly placed and secured onto sheath hub and free of cracks/ damages or excessive adhesive vacuum leak test - sample size: 100% perform pressure and vacuum leak test per procedure . per ifu : never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. corrections are not required because the investigation did not conclude that the product or process did not meet specification at the time of shipment. the investigation concluded that the product met specification at the time of shipmen. integer inc /oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may have not been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, integer.inc/oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
7465 | 1220648-2025-30353 | 22507908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22507908&pc=PYX | 06/27/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026652254 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia Stroke | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed an access site bleed following removal of the introducer and advancement of the repositioning sheath. two purse string sutures were added to achieve hemostasis at the site. roughly five days into support, the patient suffered an ischemic stroke. support continued for two more days, at which point the decision was made to withdraw care and that patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ | The investigation of access site bleeding and stroke/cva has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by placing 2 additional purse string sutures. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details.
7466 | 1220648-2025-30354 | 22508044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22508044&pc=OZD | 07/07/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 07/07/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683591 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on impella 5.5 support for mechanical circulatory support. during support, the patient experienced mild oozing at the access site. bivalirudin was withheld due to bleeding concerns. pressure dressings were applied, and platelets were transfused in the setting of chronic thrombocytopenia. the hematoma was monitored, and the insertion site was redressed. improvement was noted in site appearance with decreased oozing. bivalirudin was subsequently restarted after clinical stabilization of the surgical site. the patient remained on impella support. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7467 | 1220648-2025-30356 | 22508289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22508289&pc=OZD | 06/28/2025 | 08/13/2025 |  | 07/15/2025 |  |  | 06/28/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683632 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and lost pulses in the left lower extremity. the impella cp was removed and replaced. no patient harm was reported. | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. there is not enough evidence to exclude the impella cp pump 1 as an associated factor to the patient's outcome. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30356. b5 mso statement the first impella cp started a cascade of events, for further events related to manufacturer device report 1220648-2025-30356, see manufacturer device report 1220648-2025-30359. h1 type of report was erroneously selected on the initial manufacturer device report number. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30356.
7468 | 1220648-2025-30357 | 22508643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22508643&pc=OZD | 06/28/2025 | 08/18/2025 |  | 07/15/2025 |  |  | 06/28/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 82 KG | Non Hispanic | Asian | Death | Yes | Yes | Malposition of Device | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and was at risk for heparin induced thrombocytopenia so platelets were transfused and systemic anticoagulation was on hold. later, the patient was rolled for a bed change. during the movement, the pump moved into the aorta. the pump was repositioned but the patient coded and expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues and thrombocytopenia has been completed. the impella device was not returned for evaluation. positioning issue: the root cause of the positioning issue was most likely due to patient movement as evidenced by clinical detail of pump shifting when patient was rolled over for bed change and 'impella position in aorta' alarm triggering. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details.
7469 | 1220648-2025-30358 | 22508929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22508929&pc=OZD | 06/29/2025 | 08/22/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676229 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 70 KG | Non Hispanic | Asian | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an attempted impella 5.5 implant and during the implant, the pump was attempted in the left axillary artery and dissection occurred. vascular surgery deployed a stent graft. no further interventions and flow was restored to the arm. the staff was unable to place the impella 5.5 due to anatomy. the implant was aborted. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation has been completed. product wasn't returned for investigation, and data logs are not relevant. based on the clinical detail that delivery was aborted due to small tortuous vessels, the root cause of the delivery issue was most likely patient condition. based on the clinical details that the vascular damage occurred during access and the patient had small tortuous vessels, the root cause of the vascular damage was most likely patient condition.
7470 | 1220648-2025-30359 | 22509248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22509248&pc=OZD | 06/29/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683632 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella cp was placed for cardiogenic shock. the impella cp was successfully placed. during support, it was noted that the patient's right upper extremity with poor perfusion/cool - per cardiac surgery fellow brachial pulses were improving. the lower left extremity remained ischemic. the patient was transitioned to do not resuscitate status by the family with discussion of palliative care. the impella cp was removed and a new pump was inserted axillary. the patients outcome is critical. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7471 | 1220648-2025-30361 | 22509624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22509624&pc=OZD | 06/27/2025 | 09/23/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417603 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge clip of an automated impella controller was broken. there was no noted patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) several times. device analysis: the console was examined and a broken purge disc flag was identified. the root cause of this issue was that the purge disc detect 'flag' got stuck due to dextrose deposits. h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-30361. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7472 | 1220648-2025-30360 | 22509905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22509905&pc=OZD | 06/29/2025 | 08/18/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583091 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 95 KG |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump began to encounter high purge pressure during patient support. tissue plasminogen activator was provided in response to the alarms. support was continued until the planned wean and withdrawal the following day. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. high purge pressure: no product was returned for analysis. the data logs show a gradual rise in purge pressure with no prior motor current spikes outside p-level changes. concurrently there was a decrease in purge flow. the logs show a gradual rise in purge pressure from 6/29, at 00:00am, till alarm was triggered at 11:00am. the logs show high purge pressure remained elevated for about 10 hours before there was a decrease to normal levels over about 4 hours. purge flow values also increased as high purge pressure resolved. the root cause of the high purge pressure was most likely due to biomaterial build up as evidenced by the data log analysis revealing a gradual rise in purge pressure.
7473 | 1220648-2025-30362 | 22509915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22509915&pc=OZD | 06/29/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684320 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 84 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 was being inserted when resistance was noted. imaging was reviewed and stenosis was seen in the subclavian. the impella 5.5 placement was aborted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition. re-reviewing imaging, noted subclavian stenosis.
7474 | 1220648-2025-30363 | 22510127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22510127&pc=OZD | 06/27/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630955 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that the impella cp was placed on a acute myocardial infarction and cardiogenic shock patient for hemodynamic support. it was reported that the pump stopped minutes after implantation. suction was present prior to the pump stop. however, a clot was visible after removal. activated clotting time was in target range. the registered nurse reported that clots were also present in the diagnostic catheters. the patient was in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7475 | 1220648-2025-30364 | 22510911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22510911&pc=PYX | 06/29/2025 | 08/28/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025601206 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 114 KG |  | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella rp flex placed. the patient was supported by extracorporeal membrane oxygenation and bipella (impella rp flex and impella 5.5). the team had placed the right sided support but there was motor current and purge pressure rise/spike and a pump stop. the team suspected clot ingestion. the pump was explanted after just under two hours of support. the patient survived. and was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of high purge pressure, temporary pump stop, and thrombus has been completed. the pump was not returned. data log review showed a mc spike occurring just over 1 hour after pump start. the motor current (mc) spike began a trend of increasing purge pressure with decreasing purge flows. the mc spike triggered an alarm #13 for impella stopped - mc high. high purge pressure: the cause of the high purge pressure was most likely biomaterial ingestion due to the mc spike occurring to begin the trend of higher purge pressure. pump stop: the cause of the pump stop was most likely biomaterial due to the mc spike which was immediately followed by a gradual rise in purge pressure. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7476 | 1220648-2025-30365 | 22511266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22511266&pc=OZD | 06/29/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/29/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661913 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the decision for an impella cp was made. during support, it was noted that urine output was red, in addition to the patient being on aggrastat due to thrombus in right coronary artery. also noted was two units of blood were given due to platelet decrease. the impella cp was successfully weaned and explanted. the patient survived. | [5 empty] | The impella pump was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7477 | 1220648-2025-30366 | 22511549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22511549&pc=OZD | 06/30/2025 | 08/26/2025 |  | 07/15/2025 |  |  | 06/30/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 58 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp could not be advanced during delivery due to the patient's existing stenosis and calcification. both the right groin and left groin were attempted, but neither were viable for pump delivery. the implant was aborted as a result of the noted resistance. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition due to stenosis and calcification in right groin impeding the advancement of impella.
7478 | 1220648-2025-30367 | 22511588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22511588&pc=OZD | 06/30/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/30/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662124 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced limb ischemia. the pump was explanted and the patient developed a hematoma. the patient was taken to surgery for surgical closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7479 | 1220648-2025-30368 | 22511790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22511790&pc=OZD | 06/26/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/26/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678051 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 136 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced sleeve damage to the pump at the sheath. additionally, the patient had access site bleeding and positive blood culture results. antibiotics were given and heparin was withheld. the patient was also transfused multiple blood products in response to positive disseminated intravascular coagulation results. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7480 | 1220648-2025-30369 | 22511799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22511799&pc=OZD | 06/30/2025 | 08/27/2025 |  | 07/15/2025 |  |  | 06/30/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was removed due to the patient experiencing thrombus and hemolysis. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis and thrombus has been completed. the impella device was not returned for evaluation. hemolysis: data review: data logs showed pump ran without issue during support. there were suction alarms triggered in early of the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7481 | 1220648-2025-30370 | 22511853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22511853&pc=OZD | 06/30/2025 | 08/29/2025 |  | 07/15/2025 |  |  | 06/30/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 98 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced hemolysis. the p level of the pump was adjusted. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. urine reported as being red and concern for hemolysis on the 2nd day of impella support. there was no information related to pump position. data review: data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
7482 | 1220648-2025-30371 | 22512101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22512101&pc=OZD | 06/30/2025 | 09/03/2025 |  | 07/15/2025 |  |  | 06/30/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694656 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 93 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced red urine. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for hemolysis has been completed. the pump was not returned for investigation. the data log shows several suction alarms during the reported hemolysis issue, with fluctuations in pump flow and placement signal while running on a steady p-level. placement signal and pump flow also showed a positive trend later while on a steady p-level. no motor current spikes or major positioning issues were noted in the data log. according to the clinical notes, low cvp indicates a blood volume issue during the time of hemolysis. the cause of the hemolysis issue was most likely patient condition related, based on data log review.
7483 | 1220648-2025-30372 | 22512311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22512311&pc=OZD | 06/27/2025 | 07/15/2025 |  | 07/15/2025 |  |  | 06/27/2025 | 07/15/2025 |  |  |  |  | 07/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665760 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient developed an intracranial hemorrhage. sodium bicarbonate was infused into the purge solution in an attempt to reduce bleeding. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer as the device remains implanted supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7484 | 1220648-2025-30281 | 22492818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22492818&pc=OZD | 06/17/2025 | 07/15/2025 |  | 07/14/2025 |  |  | 06/17/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597017 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 31 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia. the patient was defibrillated. the pump was found to be mispositioned so it was explanted. upon explant, the patient lost pulses in her leg for which she underwent an embolectomy. the family decided to withdraw care and the patient expired the next day. | [5 empty] | H6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30281. | Medical safety officer review: there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7485 | 1220648-2025-30282 | 22492920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22492920&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/14/2025 |  |  | 06/23/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681937 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2025 |  |  |  |  | It was reported that a patient presented with shortness of breath and chest pain and was admitted and worked up in the emergency room by cardiology for st-elevation myocardial infarction. the patient was then transferred for medical management. the decision was made to place an impella 5.5 with potential of impella rp flex placement. it was noted that the inquire wire were advanced through the graft into the sheath and there was challenge getting through a portion of the vessel that was narrowed and not seen on computed tomography of the chest (prior to device pre assessment). multiple attempts using alternative wires under fluoroscopy were attempted throughout the case. the patient remained stable throughout the case. help was called in to assist with the wire access. the decision was made to place an impella cp via the groin. additionally, after the axillary placement of the pump was aborted, the patient went down for a washout of the axillary pocket which required replacing multiple units of blood products. the patients outcome is unknown. | [5 empty] | The investigation of delivery issues and access site bleeding has been completed. the impella device was not returned for evaluation. delivery issue : the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through the vasculature. access site bleeding : the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2025-30282 in accordance with updated procedures. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7486 | 1220648-2025-30283 | 22492962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22492962&pc=OZD | 06/22/2025 | 11/24/2025 |  | 07/14/2025 |  |  | 06/24/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1669290 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 98 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support and during this support when doing a purge cassette change, the automated impella controller (aic) stopped recognizing the purge cassette and the screen froze. consequently, the aic was exchanged with no harm to the patient noted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. imc logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) several times. device analysis: the console was examined after arriving in fs. a broken purge disc flag was identified. conclusion: the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. es2023-0090 documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h3 updated. d10 concomitant medical products and therapy dates updated. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7487 | 1220648-2025-30284 | 22493016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493016&pc=OZD | 06/25/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/25/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685044 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 111 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and lost doppler distal pulses and his toes were mottled. the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7488 | 1220648-2025-30285 | 22493120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493120&pc=OZD | 06/25/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/25/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683620 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a hematoma at their impella 5.5 access site during support. the patient was taken to the operating room for a washout to treat the condition. support was continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7489 | 1220648-2025-30286 | 22493242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493242&pc=OZD | 06/25/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/25/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667267 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the placement signal of an implanted impella 5.5 pump appeared low during patient support. the placement signal was disabled and support was continued uninterrupted. three days later, the purge pressure increased resulting in a drop in purge flow. the purge pressure measured around 1200 mmhg and a purge pressure high alarm appeared. a lysis protocol was followed to manage the high purge pressure. the patient remains on impella 5.5 support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist, section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7490 | 1220648-2025-30287 | 22493396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493396&pc=OZD | 12/18/2024 | 09/08/2025 |  | 07/14/2025 |  |  | 12/18/2024 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586753 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Cardiac Tamponade | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing for the patient admitted in for a coronary artery bypass grafting. there was a cardiac tamponade and re-thoracotomy performed for coagulation imbalance. the team gave blood for the event and the pump was explanted and patient survived the harm. additionally, there was acute kidney injury. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of ventricle perforation and renal failure has been completed. the impella device was not returned for evaluation. ventricle perforation: the cause of the ventricle perforation was not determined since product was not returned and insufficient clinical details. renal failure: the cause of the renal failure was not determined since product was not returned and insufficient clinical details.
7491 | 1220648-2025-30288 | 22493411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493411&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/14/2025 |  |  | 06/25/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677301 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 70 KG | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support and developed hematoma and bleeding from the right groin. the impella access site was on the left side. the right side had previous iliac stenting. heparin was withheld and blood products were transfused. the patient also showed signs of ischemia to bilateral lower extremities and right hand had rising lactate. the patient was made comfort care and later expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel, limb ischemia, and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the patient is having bleeding from multiple sites. the root cause of access site bleeding is determined to be patient condition because the patient is bleeding from multiple sites. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
7492 | 1220648-2025-30289 | 22493528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493528&pc=OZD | 06/24/2025 | 09/09/2025 |  | 07/14/2025 |  |  | 06/24/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648163 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp slipped just a little too deep into the ventricle and placement signal low alarm occurred. the alarms were rectified. additionally, the cp was not able to go higher than p5 without having suction alarms. support continued until successful explanation and the patient remained in stable condition. | [5 empty] | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. low pump flow: data log reviewed: no motor current (mc) spike observed. suctions occurred on (b)(6) above there were 3 triggers of intermittent suction and 2 triggers of continuous suction. continuous suction particularly suggests that mc values are low during both phases of the cardiac cycle. clinical stated, compliant reported about persisting alarms "placement signal low". the evening of (b)(6), there was a significant drop in placement signal (ps) minimum values close to 0 mmhg. for this reason, the "placement signal low" alarms began to trigger. per rep email, the pump slipped just a little too deep into the ventricle at that point. the alarms were rectified after a short time and did not return. ps waveform was found to be ventricular in the data logs at the time of the issue and once the ps returned to normal, the suction and placement signal low alarms resolved. the cause of the suction was most likely improper positioning since the pump was reported to slip deep. positioning issues: the cause of the positioning issue cannot be determined due to insufficient clinical data provided, unknown why device slipped deep. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7493 | 1220648-2025-30290 | 22493555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493555&pc=OZD | 06/26/2025 | 09/10/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677671 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male | 100 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed oozing from their repositioning sheath connection following impella cp implant. it was noted that the blue suture hub was a little too deep in the groin. the staff was advised to replace the pump and the device was removed after eight hours of support. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. access site bleeding: clinical reported blood oozing out between the blue suture hub and re-access port, as well as blue suture hub being positioned slightly deep in the groin. the pump was returned for investigation and no abnormalities were observed. the cause of the access site bleeding was most likely use related, as clinical reports the blue suture hub being too deep in the groin. the blue butterfly (suture hub) is not a hemostatic seal and not intended to be inserted into the skin. d9 device returned to manufacturer date added.
7494 | 1220648-2025-30291 | 22493698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493698&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681939 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was attempted multiple times to be implanted but the procedure was aborted due to the patient¿s vascular anatomy. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. | The investigation for delivery issue has been completed. the device was not returned for evaluation. nor were data logs provided for review. therefore the root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to patient's anatomy..
7495 | 1220648-2025-30292 | 22493900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493900&pc=OZD | 06/24/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/24/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654363 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female | 72 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an elderly patient transferred from community to tertiary center with acute myocardial infarction and cardiogenic shock. the impella cp pump was placed but there was ventricular tachycardia on day after implant treated by defibrillation. pump remained on for 3 days and then expired on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7496 | 1220648-2025-30293 | 22493940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22493940&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684269 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility stated that a decision was made to use an impella cp to support for percutaneous coronary intervention and recovery in the intensive care unit. during use in the catheterization lab, there was a "placement signal unreliable¿ alarm which lost placement signal halfway through case. shockwave was used at the time of placement signal loss. the decision was made to remove the impella cp and the case was successfully completed with a new impella cp. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for optical signal issue has been completed. there were no data logs returned. the clinical stated that there was a loss of placement signal halfway through the case during the time of shockwave. the returned product was able to reproduce ps unreliable alarms and the 5s parameters were revealing a damaged sensor. gain 1.37, light 2.75, contrast 5.8%, snr 12, and cavity length 32768. the laser continuity test did not reveal any breaks in the fiber line as light was able to exit the sensor face. the root cause of the optical signal issue is most likely due to a damaged sensor due to shockwave but the distance between the senor and shockwave is unknown, based on clinical details and returned product analysis.
7497 | 1220648-2025-30294 | 22494005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494005&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/14/2025 |  |  | 06/25/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658063 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed for higher level of care transfer. it was noted that while impella was being prepped/primed the registered nurse inadvertently started the pump outside of the body. staff report that the impella was not over the wire when pump was started. they disconnected the pump, and restarted the aic. the pump was then inserted pp, wire was removed, and pump was started in auto mode. additionally, point of care ultrasound was performed at bedside and cp measured at 5cm. impella pulled back 3cm. it was also noted that the staff was doing an extubation trial and the patient began moving his extremities, subsequently pulling the device across the aortic valve. md attempted to reposition the device, however, was unsuccessful. decision was made to remove impella. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.5 unknown. | The investigation of positioning issues has been completed. pump was not returned for investigation data log review was not required for this case run. as per the clinical details, the patient started moving his extremities, which resulted in the device being pulled across the aortic valve. the cause of the positioning issue was most likely patient movement, based on given clinical details.
7498 | 1220648-2025-30296 | 22494145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494145&pc=OZD | 06/25/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/25/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638883 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during impella access there was bleeding, ischemia and dissection. it was noted that there were unable to repair in cath lab, so patient going to operating room for vascular repair. the patient was weaned successfully and cp removed. contralateral access gained due to inadequate right femoral artery hemostasis. angiogram showed right common femoral artery occlusion. patient taken to intensive care unit post procedure with plan to go to operating room for surgical repair of right femoral artery. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: he instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7499 | 1220648-2025-30297 | 22494277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494277&pc=OZD | 06/18/2025 | 10/24/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680964 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 110 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 accidentally kicked the impella power plug and it became unplugged from the automated impella controller. the nurse quickly plugged the pump back in and support continued. | [5 empty] | D6b "if explanted, give date" corrected. | The investigation for temporary pump stop has been completed. the device nor data logs were returned for evaluation. clinical details noted that while patient was rolling in the bed, they accidentally kicked the impella plug and the pump plug became disconnected from the aic. pump was able to be reconnected quickly and started back up at previous p-level without issue. the root cause of the temporary pump stop was most likely due to patient movement based on clinical details of patient rolling and accidentally kicking aic plug out of console. clinical details noted that "pump in aorta" alarms were occurring while on support. additionally, ao & lv waveforms were lost then returned without troubleshooting. flow and motor current were not affected. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7500 | 1220648-2025-30298 | 22494375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494375&pc=OZD | 06/26/2025 | 12/04/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomedical engineer and clinical representative reported a console with a controller error. the automated impella controller was to be replaced and the instructions for use followed by the use of a backup console. no noted harm or injury as this occurred outside of patient use in a training. | [5 empty] | Data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. additionally, no controller error alarm was seen in the imc logs. device analysis: console ic4102 was sent back to field service for further investigation. the issue was not reproduced. further, visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel, have caused the issue with the pbm. the pbm was replaced with a known good one. subsequently, the test pump and purge cassette was run successfully without any issues. no controller error alarms or related issue were seen. root cause: no issues were found with the aic related to the reported fm of "aic - controller error" alarm. the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7501 | 1220648-2025-30299 | 22494447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494447&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546041 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp could not be completed due to the patient's anatomy. the catheter could not pass the right iliac. the patient survived. | [5 empty] | The investigation for delivery issue has been completed. the device was not returned for evaluation nor were data logs provided for review. therefore, the root cause of the delivery issue was most likely patient condition related since the cp was unable to pass through a patients right illiac artey. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
7502 | 1220648-2025-30300 | 22494598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494598&pc=OZD | 06/26/2025 | 11/25/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1620355 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 52 KG | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) had a controller error and a battery failure error during patient support. the aic was replaced. | [5 empty] | The investigation has been completed. data analysis: ¿ the reported complaint case was initiated on (b)(6) 2025, involving pump sn: (b)(6). ¿ on (b)(6) 2025, during support, the imc log process likely crashed, as no imc log entries were recorded after 17:03:57. however, the lt log indicates that the pump remained under support during this period. note: since this support initiation, the rt log files were likely being zipped approximately every 1 hour and 18 minutes. before the imc log crash, the last rtlogger file record was at 15:46:12, shows that the imc log process likely crashed just prior to the next expected rtlogger file recording around 17:03:57. on the complaint date, (b)(6) 2025, at 10:31:31, the imc log resumed recording for a brief period before stopping again after 13 minutes. the rt log ended 3 minutes after the imc log stopped, indicating an unexpected console shutdown. this unexpected controller shutdown was misreported as a ¿controller error¿. subsequently, the pump was transferred to a backup console ic7179. on (b)(6) 2025, after ic9319 was manually rebooted, the console displayed ¿unexpected controller shutdown ¿ alarm,¿ event #603. device analysis: the console was tested upon arrival at fs. upon power cycling the console 10 times, the 'unexpected controller shutdown' could not be reproduced. root cause: the root cause of the unexpected controller shutdown was not determined due to the inability to reproduce it. there were no instances of a ¿controller error¿ or ¿battery communication failure¿ in the logs. the unexpected controller shutdown was misreported as a ¿controller error¿ and ¿battery communication failure. d10 concomitant medical products and therapy dates updated. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7503 | 1220648-2025-30301 | 22494962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494962&pc=OZD | 06/26/2025 | 11/25/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1874262 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 77 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support (mcs) as a bridge to durable ventricular assist device or orthopedic heart transplant. approximately 15 days into the support, the nurse unplugged the automated impella controller (aic) for patient ambulation and the aic generated critical ¿battery temp high¿ alarm. the nurse restored alternating current (ac) power to the aic resolving alarm. the nurse was advised by customer support to continue monitoring and change the aic if the alarm returns. subsequent discussion was made with the team, and it was decided to bring the backup aic to the bedside. the patient returned to bed. the original/complaint aic was again separated from ac power with no battery alarms noted for two to three minutes. the back up aic was then removed outside the patient room. the patient was transferred or placed on the original aic with no harm noted to the patient as a result of this event. the patient continued on impella mcs. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. data analysis: console logs from the complaint date revealed the ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc. battery and power data embedded in ltlogs do not show any temperature rise or elevated battery temperatures, as ltlog data is recorded every 60s. device analysis: the console¿s batteries and power management system were tested, and no issues were observed. root cause: the battery level low alarm was due to a momentary communication glitch between the batteries and pbm, where a single sample from battery data (in this case value of battery temperature) was corrupt. d9 updated. h3 updated. d10 concomitant medical products and therapy dates updated.
7504 | 1220648-2025-30302 | 22494967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22494967&pc=OZD | 06/24/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/24/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686539 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 72 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp for the support of the patient admitted in with acute myocardial infarction and cardiogenic shock. the patient had right femoral arterial tract bleed requiring transfusions overnight and in the morning. fem stop was placed and the doctor placed a stitch bedside and bleeding stopped. later it was reported that care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7505 | 1220648-2025-30303 | 22495172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495172&pc=OZD | 06/27/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during preparation of an automated impella controller, a controller error was encountered. consequently, the aic was replaced with no noted harm to the patient. no further details were provided. | [5 empty] | Theconsole was received from the customer and an investigation regarding the returned device has been completed. console logs were partially consistent with what was reported in the complaint. instead of a controller error, instances of controller failure (epm1 sw corruption) alarms were observed in the logs. this is likely what the clinical staff meant to report as there were no controller errors observed in the logs. console was tested and the controller failure was able to be reproduced while testing on an engineering lab bench. the issue resolved after troubleshooting with a known good impellatronic board. : the root cause of the controller failure was a corrupted epm firmware on the impellatronic.
7506 | 1220648-2025-30304 | 22495180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495180&pc=OZD | 06/24/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/24/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490979 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 130 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient presented for a planned intra-aortic balloon pump placement for induction of anesthesia and then to change to 5.5 support prior to coming off pump. it was reported that the patient had several episodes of non perfusing beats causing mean arterial pressure to drop and complete reliance on pump. during support it was also reported that the patient was stable however had suction alarms related to volume. two units packed red blood cells were given as well as albumin. in addition to when the patient moved to chair had a brief run of non perfusing ventricular tachycardia with complaints of dizziness only. the patient also went into atrial fibrillation and was on iv amio. it was also noted that the patient went into atrial fibrillation with rapid ventricular response requiring cardizem and amiodarone. the 5.5 was successfully weaned and explanted in operating room without issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7507 | 1220648-2025-30305 | 22495395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495395&pc=OZD | 06/27/2025 | 09/10/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624105 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support (mcs) and six hours after the implant, high purge pressure and low purge flow was encountered. purge cassette change did not solve problem. lyse protocol was sent out on behalf of the medical office and started, given overnight. subsequently, the purge pressure settled down and purge flow went. all metrics were back in range. it was noted there was no harm to the patient as a result of this event. the patient continued on impella mcs for approximately 10 days before being bridge to a left ventricular assist device with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. insufficient data logs returned as time of reported high purge pressure. clinical details noted that 6 hours into support, high purge pressure with low purge flows. purge cassette did not resolve issue, tissue plasminogen activator (tpa) was added to purge overnight and resolved issue. the cause of the high purge pressure could not be definitively determined as no product and insufficient data logs were returned for evaluation.
7508 | 1220648-2025-30306 | 22495591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495591&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Difficult to Setup or Prepare | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during preparation for impella cp implant, the scrub prematurely pulled out the easyguide lumen. a new pump was subsequently prepped for implant. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the delivery issue has been completed. the device was not returned for evaluation, nor were data logs provided for review. therefore the root cause of delivery issue was most likely use issue since during preparation for impella cp implant, the scrub prematurely pulled out the red lumen.
7509 | 1220648-2025-30307 | 22495628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495628&pc=OZD | 06/27/2025 | 08/27/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654279 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy in cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support (mcs). post implant, same day, the distal part of anticontamination sleeve tore and was fixed with tegaderm. there was no report of harm to the patient as a result of this event. the patient continued on impella 5.5 mcs for approximately seven days before being successfully weaned. the impella 5.5 was explanted and a left ventricular assist device was placed with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. it was reported that the sleeve tore on the first day of support due to excessive force. product was not returned for investigation and data logs are not relevant. based on the clinical detail that the sleeve tore due to excessive force, the root cause of the product damage was determined to most likely be a use issue.
7510 | 1220648-2025-30308 | 22495719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495719&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612814 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the patient¿s heart transplant and explant of impella 5.5, a pseudo aneurysm was noted to the right axillary artery which ruptured, accompanied by a dissection of the artery. vascular surgery repaired site. also, a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding due to rupture of right axillary pseudoaneurysm with a "definite" relationship to the impella device used. patient survived. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (unites states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation)." section: warnings & cautions: cautions: "physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta." caution: "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance." | The investigation for the artery dissection and access site bleeding has been completed. product was returned for evaluation. upon evaluation, the pump was found to be cut. the cannula had a small portion of the catheter returned along with part of the repo unit with a piece of the graft attached. no kinks noted in the cannula. the clinical states that during explant of the 5.5, there was a pseudo aneurysm which ruptured and caused a bleeding and a dissection. due to lack of clinical details provided regarding the explant and insufficient product returned, the root cause of the artery dissection cannot be determined. the root cause of the access site bleeding cannot be determined.
7511 | 1220648-2025-30309 | 22495731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495731&pc=OZD | 06/30/2025 | 09/25/2025 |  | 07/14/2025 |  |  | 07/07/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674266 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the anchoring ring of the impella could not be closed completely and the shaft was not able to be secured in place during impella cp management. the suspected defect was dealt with by applying a tape. there was no patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the product was returned. when the pump was being pulled out of the box, the anticontamination sleeve became ripped off of the ring. there were no damages noted on the product as when the anchor was tight, it sat flush. the guidewire functioned within specification as it was able to lock and unlock with ease and move along the catheter with no issues. under the guidewire there was slight catheter deformation that can occur when the anchor is tightened too tight. there were no issues with rotating it. there was no problem found with the device in regard to the positioning issues.
7512 | 1220648-2025-30310 | 22495895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495895&pc=OZD | 07/05/2025 | 09/10/2025 |  | 07/14/2025 |  |  | 07/07/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622111 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a high purge pressure alarm appeared during impella 5.5 support. it was noted that the purge pressure was over 1000 mmhg at the time of the alarm. the doctor expressed concerns that the pump may stop based on the provided readings and opted to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump and purge cassette were returned for investigation. the returned pump was visually inspected and no abnormalities were observed. the pump was primed without issue and the high purge pressure was unable to be reproduced since purge values were within specification. data log showed on (b)(6) 2025 at 10:34:17 several temporary high motor current pump stops occurred with a sudden rise in purge pressure to 1050 mmhg indicating a biomaterial ingestion event. the pump was eventually able to be restarted but purge pressure remained high until (b)(6) 2025 at 13:46:46 when p-level was reduced to 0 which resolved high purge pressure with values returning to 600 mmhg. data logs from a day prior to the high purge pressure event also showed a long cassette change and a long bag change both exceeding the recommended 2 minutes resulting in brief losses of purge to the pump. purge heparin was also noted to be decreased from 20 iu/ml to 0 iu/ml on 7/4 and it is unclear if bicarb was added to the purge. the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review. temporary pump stop: failure mode temporary pump stop was added based on the data log review findings. the cause of the temporary pump stop was most likely biomaterial ingestion based on the data log review. thrombus: thrombus failure mode was added based on the data log and high purge pressure investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added.
7513 | 1220648-2025-30311 | 22495982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22495982&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681804 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella cp implant and after the pump was implanted, a kink was visible and there were frequent suction alarms. no force was applied on insertion, the kink was present due to patient anatomy. support continued and patient was eventually weaned. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation for low pump flow and product damage ( cannula kink) has been completed. a kinked cannula was observed on the returned product. the pump was free of any biomaterial, and no other physical damage was observed. the pump ran as expected in a bucket of water. the data log shows low pump flow with active suction alarms from the start of the case, without any noted motor current spikes. the cause of the low pump flow issue was a kinked cannula. the cause of the product damage related to the cannula kink issue was most likely due to patient conditions related to difficult anatomy, based on the provided clinical details.
7514 | 1220648-2025-30312 | 22496188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22496188&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a physician elected to place a patient on impella cp support prior to percutaneous coronary intervention. the first attempt to place an impella cp via the right femoral artery was unsuccessful having felt great resistance in attempting to cross over the highly calcified aortic arch. upon removing and inspecting the impella cp, the surgeon noticed that the abiomed 0.018¿ wire was severely kinked. the medical doctor elected to place a fresh 0.018¿ v-18 wire into the left ventricle (lv) and make a new attempt at placing the patient on impella support. the physician successfully placed the v-18 wire into the lv without complication. however, once again when inserting the impella cp over the wire, the impella again met resistance upon trying to cross over the aortic arch. after a few more attempts to advance the impella, the physician then elected to abort the impella insertion and place the patient on intra-aortic balloon pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (guidewire kink) has been completed. product was not returned for investigation. data log review was not required for this case. the cause of the delivery issue was most likely patient condition related to calcified aortic arch. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided.
7515 | 1220648-2025-30313 | 22496202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22496202&pc=OZD | 06/26/2025 | 08/22/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684265 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a impella cp placed pre-percutaneous coronary intervention. however it was noted that during impella access there was bleeding after peel away sheath removed and repo sheath placed. forward tension sutures used and angle match done. ct surgery consulted for additional help and placed multiple mattress sutures and pressure dressing was applied. additionally, there were suction alarms at p-9, dropped to p-7 and volume given. the cp was successfully weaned and explanted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | During the implant procure, arterial bleeding from impella access site despite attempting mattress sutures, angle matching, and forward tension sutures. bleeding was seen after peel-away was removed and repositioning sheath placed. activated clotting time (act) was 227 at time of bleed and noted to be 326 at case start. the pump was not returned. data logs were not recovered. access site bleeding - major: the cause of the access site bleeding - major was most likely high patient act values. low pump flow: the cause of the low pump flow was most likely patient condition related due to lower p-level and blood volume given after suction alarms with no further mention of suction events.
7516 | 1220648-2025-30314 | 22496450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22496450&pc=OZD | 06/24/2025 | 11/03/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653653 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 110 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem |  |  |  |  | Stroke/CVA | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information was received that the patient encountered stroke during support. | The complainant reported that two days into impella 5.5 support, a placement signal not reliable alarm appeared. the team was made aware of how to handle the alarm should it recur. later that day, flows on the pump were lower than expected. after multiple alarms, the pump unexpectedly stopped. the pump was removed as a result of the noted stoppage. the patient remained supported via extracorporeal membrane oxygenation. |  |  |  |  | B5 updated with additional information. h6 health effect clinical code 1770 stroke was added due to additional information. h6 health effect impact code 4614 serious injury added due to additional information. | The investigation of low pump flow, optical signal issue, and pump stop has been completed. optical signal issue: clinical details report that the psnr alarm triggered on (b)(6) 2025. data logs were reviewed, and the alarm was confirmed. around 06:30 on (b)(6) 2025, there was a sudden drop in snr and gain, lamp remained at 100%, contrast began to oscillate between +/- 3200, and light dropped. the optical 5s did not recover for the remainder of the case. returned product had all 5s parameters still out of specification when connected to a console, and a controller failure alarm triggered. the pump failed the laser continuity test with light leaking from the optical fiber. the leak was coming from within the red pump plug on the patient cable side where the optical fiber inserts into the ferrule. additionally, the wiring coming out of the patient cable appeared to have migrated into the red pump plug, over the near post, causing a sharp angle between the optical fiber and ferrule connection. the data logs and returned product are consistent with a torqued optical fiber. based on the pattern seen in data log review and abnormal migration of wiring/optical fiber within the red pump plug, the root cause of the optical signal issue was determined to be a damaged fiber due to manufacturing operator error. low pump flow: data logs were reviewed and there were several abnormalities. pump flows were variable and often low throughout the entire case. there were 95 triggers of alarm #33 and 92 triggers of #35 (position unknown). additionally there were 11 suction alarms, and 13 preventing retrograde flow algorithm notifications. there were no abnormal motor current spikes. the only spikes were when the pump was started at case start and restarted after being turned off on 24/jun/2025. purge pressure was stable throughout the case. impella position unknown alarm #33 triggers when both imod (mc modulation) and ps pulsatility are below their respective threshold for the p-level. this suggests that the pump and patient both lack pulsatility. similarly, the suction alarms and triggers of the retrograde flow algorithm indicate that mc was low, particularly mcmin. this suggests limited volume flowing through the cannula. returned product was found to have significant amounts of biomaterial in the outflow cage/wrapped around the impeller after soaking. when the pump was attempted to be run in a bucket, it initially experienced high mc pump stops, then low pump flows, then saturated flows. however, when the biomaterial was fully removed, the pump flows were normal for p-9. it is unclear whether or not the biomaterial was ingested during the case in the field or not because data logs do not suggest that it was, but how it was found to be sitting in the outflow cage and cannula is abnormal if it were to be picked up pump explant. there is potential that further data logs at the end of the case were missing. since the characteristics seen on data logs do not align with the significant amount of biomaterial on returned product, the root cause of the low pump flows could not be determined. pump stop: data logs were reviewed and there were no pump stop alarms noted at any point in the case. additionally, there were no abnormal motor current spikes during the case. toward the end of the case, the preventing retrograde flow and reverse flow algorithm notifications did trigger. additionally, the final real time log showed the pump flows being calculated as 0 l/min even though the pump was running at a ms equivalent to p-8. these factors could be what the user saw and believed to be a pump stop, however, this could not be confirmed. returned product passed electrical testing of the motor cables, confirming there was no electrical damage. the pump was found to have significant amounts of biomaterial in the outflow cage/wrapped around the impeller after soaking. when the pump was attempted to be run in a bucket, it initially experienced high mc pump stops, then low pump flows, then saturated flows. however, when the biomaterial was completely removed, the pump flows were normal for p-9. returned product having significant biomaterial in the outflow cage and around the impeller is indicative of a biomaterial ingestion, however, there were no pump stop alarms or signatures of ingestion in the available logs. for these reasons, the root cause of the pump stop could not be determined. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
7517 | 1220648-2025-30315 | 22496540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22496540&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669884 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support during a bridge to transplant, and after approximately two days after start of the support, a purge system blocked alarm appeared coinciding with a purge pressure measurement of over 1000mmhg. there were no kinks were noted in the setup, and replacing the purge cassette failed to resolve the issue. tissue plasminogen activator (alteplase) was subsequently placed in the purge for two hours, and the purge pressure and flow returned to normal. there was no harm to the patient as a result of this event. the patient continued on the impella 5.5 support for approximately 10 days before receiving a heart transplant and the impella 5.5 device successfully explanted with a survived outcome. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was returned for evaluation. high purge pressure (hpp): data log review showed approximately 16 minutes of pp rise. shortly after this rise, there was a long cassette change. prior to the rise, there were no clear trends/spikes in motor current (mc) besides slight changes in amplitude. the hpp was sustained for about 2 hours before temporarily recovering to normal pressure. the pp began to rise again over 20 minutes with another long cassette change within that time. hpp was sustained for about another hour before resolving in a 15 minute pp decrease. there were no motor current spikes prior to either pp rise. the data logs were not characteristic of a kinked purge line. the pump was returned with the catheter cut. no biomaterial was found in or surrounding the purge gap. there were a few bends and kinks in the catheter but purge pressure testing could not be completed since the pump was returned cut at the catheter. the cause of the high purge pressure was most likely biomaterial as there was an initial 16 minute rise in pp and no motor current spikes based on data logs but the type of biomaterial is unknown since there were fluctuations in pp that did not clearly match an ingestion or buildup pattern. | The impella device was received from the customer. upon completion of the investigation, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7518 | 1220648-2025-30316 | 22496822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22496822&pc=OZD | 06/27/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585548 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 77 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. at the receiving facility, the patient remained on support. however, the impella cp had been left in place via the retained peel-away introducer sheath. the limb associated with the access site became ischemic. the patient subsequently required limb perfusion intervention and fasciotomy for compartment syndrome. after 69 hours of support, the impella cp was explanted. surgical repair of the access vessel was performed to achieve hemostasis. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7519 | 1220648-2025-30317 | 22497185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22497185&pc=OZD | 06/25/2025 | 10/13/2025 |  | 07/14/2025 |  |  | 07/01/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420044 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a console located in a storage room was exposed to water when the sprinkler system was activated. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) ¿ damage before therapy has been completed. the impella device was returned for evaluation. data analysis: not applicable, as the console was not in use on the complaint occurrence date. device analysis: this console was tested after arriving and, upon visual inspection and running the test cassette and pump no abnormalities were found. root cause: the root cause of water exposure to the aic is a use issue. d9 device returned to manufacturer date added b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-30317. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7520 | 1220648-2025-30318 | 22497999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22497999&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026688391 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient underwent an attempted implantation of an impella 5.5 for mechanical circulatory support. multiple attempts to advance the impella 5.5 through the graft anastomosis were unsuccessful despite repositioning and changes in arm angle. the pump was inserted but could not navigate the turn at the graft site to improve vessel alignment. as a result, the procedure was aborted and an impella cp was placed via the right femoral artery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since the pump inserted but would not make turn from graft anastomosis improve vessel despite multiple attempts.
7521 | 1220648-2025-30319 | 22499134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22499134&pc=OZD | 06/26/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/26/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665841 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 87 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient became septic during impella 5.5 support. the origin of the sepsis could not be determined. it was additionally reported that the patient experienced a gastrointestinal bleed that was believed to be unrelated to the pump. support was continued uninterrupted. after thirteen days of support, the patient's underlying condition did not improve and therefore care was withdrawn. the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components . ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7522 | 1220648-2025-30320 | 22500282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22500282&pc=OZD | 06/23/2025 | 08/12/2025 |  | 07/14/2025 |  |  | 06/23/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670768 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | Asystole | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support experienced ventricular tachycardia which progressed to ventricular fibrillation and asystole. the patient had spontaneous return of spontaneous circulation. the patient was then put on higher vasoactive support. however, after 6 days of support the patient expired. | [5 empty] | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrythmia could not be determined. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients."
7523 | 1220648-2025-30321 | 22500340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22500340&pc=OZD | 06/28/2025 | 08/26/2025 |  | 07/14/2025 |  |  | 06/28/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666089 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to advance beyond the aortic arch during attempted delivery. multiple wires were utilized, however the pump was still unable to overcome the stenosis at the anastomosis of the ascending aorta. the implant was aborted as a result of the noted issue. | [5 empty] | A4 and a5, race, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since it appears that there is a stenosis at the anastomosis of the ascending aorta which led to cp pump get hung up on the aortic arch.
7524 | 1220648-2025-30322 | 22500375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22500375&pc=OZD | 06/27/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084303 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Loss of Data | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During investigation of the console, the investigation team found a software issue of automated impella controller loss of opm data. to resolve the issue, the console was repaired before sending back to the customer. there was no patient harm. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. real time log data indicated that those alarms were related to momentary loss of opm data (all signals from optical bench interpreted as -16384 or 0xc000, due to communication crc error). the log data also showed continuously failing icb handshakes indicating possible rlm malfunction. rlm journal showed frequent unexpected reboots. the console was tested and rlm was able to connect to the engineering test wifi but data streaming was not working, and its backstrap cable was electrically compromised (which was contributing to frequent reboots of rlm). after replacing backstrap cable, data streaming issue would only intermittently resolve. analysis of usb traffic from aic to rlm using protocol analyzer confirmed intermittent communication, even after successful icb handshake, suggesting that breakdown of communication might be not between aic and rlm firewall but firewall and som. there was no evidence of rlm som microsd card corruption. console alarms not being logged on impellaconnect portal were due to data streaming loss, caused by failure of rl03514 and its backstrap cable. the cause of the aic issue was a software issue.
7525 | 1220648-2025-30325 | 22500711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22500711&pc=OZD | 06/28/2025 | 07/14/2025 |  | 07/14/2025 |  |  | 06/28/2025 | 07/14/2025 |  |  |  |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026692014 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had minimal distal pulses to the right lower extremity. the pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7526 | 1220648-2025-30324 | 22500741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22500741&pc=OZD | 06/27/2025 | 08/26/2025 |  | 07/14/2025 |  |  | 06/27/2025 | 07/14/2025 |  |  | 07/31/2025 |  | 07/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388207 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while advancing an impella cp during delivery, the inserted guidewire pulled back across the aortic valve. the devices were removed and access was reobtained with a new .018 guidewire. the impella cp implant was successful upon reinsertion. however, a placement signal not reliable alarm appeared on startup. the waveform was said to be dashed out at the time of the alarm. no intervention was required and support was continued uninterrupted. there was no patient harm. | [5 empty] | The investigation of optical signal issue and delivery issues has been completed. the impella device was not returned for evaluation. pump was not returned for investigation. data log shows the placement signal not reliable (psnr) alarm after pump was inserted into body. the trend on 5s optical parameter were match the known trend of the optical fiber damage during psnr event. placement signal was not seen recover during whole case run. the cause of the optical signal issue was most likely a damaged optical fiber line, based on the data log review. however, the exact location of the damaged fiber line was unknown without a returned product investigation. the cause of the delivery issue was not determined since product was not returned and sufficient clinical details was not provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7527 | 1220648-2025-30231 | 22477922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22477922&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657765 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient arrived via helicopter. on transfer to a bed from the stretcher, impella cp had an in aorta alarm occur. transesophageal echocardiogram and transthoracic echocardiogram confirmed the pigtail was out of the left ventricle and sitting in right coronary cusp. hemolysis was also noted. urine was red and the patient was clearing lactate. therefore, the pump was explanted and replaced with a new impella cp smartassist. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely improper positioning of the device, based on data log review and provided clinical details. the cause of the positioning issue was most likely patient movement, based on provided clinical details. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7528 | 1220648-2025-30232 | 22478008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478008&pc=OZD | 06/25/2025 | 09/09/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626740 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella 5.5 triggered a high purge pressure alarm during patient support. it was noted that the purge pressure had reached 1068 mmhg at the time of the red alarm. thhe next day, lysis protocol was applied in response to the alarm resulting in an increase in purge flow and decrease in purge pressure. support was continued uninterrupted. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella 5.5 returned for evaluation. upon visual inspection, minimal biomaterial present around part of outflow cage. pump was connected to console and purged for approximately 10 minutes with no high purge pressure alarms triggering; pump was then run at p-9 and no alarms were triggered. some biomaterial was kicked off the pump when the pump was turned on. data logs were reviewed and the purge pressure rise was observed to be approximately 1 hour over two real time data logs. no motor current spikes were present at start of purge pressure rise. clinical details noted that purge pressure high alarms were triggered. tissue plasminogen activator (tpa) was added to purge and alarms were able to resolve. the cause of the high purge pressure was most likely due to biomaterial buildup based on returned data log analysis of long purge pressure rise and clinical details noting that tpa resolved the issue.
7529 | 1220648-2025-30233 | 22478087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478087&pc=PYX | 06/25/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561621 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male | 110 KG | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was having an impella rp flex implanted and during the implant, 0.027" guidewire kinked. a new guidewire was used successfully. additionally, the patient had suspected hemolysis. the p level was lowered and support was continued. the patient survived the procedure. | [5 empty] | The investigation of product damage (guidewire kink) and hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: data review: data logs showed pump ran without issue during support. there were high pressure purge alarms triggered near the end of the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. guidewire kink: the root cause of guidewire kink was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
7530 | 1220648-2025-30234 | 22478181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478181&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672710 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 95 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision was made to place an impella cp for early stages of cardiogenic shock and perform primary percutaneous coronary intervention. it was noted that the patient was a little restless and a small hematoma was noted around the impella sheath after the patient arrived to the intensive care unit. the patient was progressing well, tolerated the wean, and the decision was made to remove impella cp instead of keeping it in overnight due to the hematoma and concern of it worsening after starting heparin infusion. the impella cp was successfully weaned and explanted post percutaneous coronary intervention. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient movement because the patient was little restless and hematoma was noted around the sheath after the patient arrived in icu. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7531 | 1220648-2025-30235 | 22478231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478231&pc=OZD | 06/25/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648412 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male | 120 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support developed compartment syndrome on the leg where the impella access site is located. intervention was done for the patient¿s leg to have perfusion again. additionally, suction alarms occurred and two units of blood products were administered. the patient¿s leg has optimal blood flow. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7532 | 1220648-2025-30236 | 22478353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478353&pc=OZD | 06/25/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683495 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient developed hemolysis during impella cp support. the support level was reduced and the condition resolved. support was continued uninterrupted and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7533 | 1220648-2025-30237 | 22478523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478523&pc=OZD | 06/21/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/21/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679731 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a impella 5.5 placement, in hopes to optimize the patient while performing a high-risk percutaneous coronary intervention. during support it was reported that the patient kept going in and out of atrial fibrillation, when this happens patient became hypotensive, amiodarone bolus was given. the patient was also transfused platelets due to platelets dropping. the patient also experienced some suction issues which they gave some volume to resolve. in addition, the patient continued to have orally bleeding when heparin was restarted and at time required to go up on impella flow. heparin was stopped and continue to slow wean. there was also swelling to right upper extremities and it was noted that they plan to remove impella. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7534 | 1220648-2025-30238 | 22478561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478561&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682001 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 99 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had access site bleeding. blood products were transfused but the patient's hemodynamics declined and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7535 | 1220648-2025-30240 | 22478739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478739&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671894 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed a clot in the pericardial sac space. the patient's family declined surgery and care was withdrawn. the patient subsequently passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | A2, a4, and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7536 | 1220648-2025-30241 | 22478953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22478953&pc=OZD | 06/23/2025 | 07/29/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661144 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 119 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp exhibited access site bleeding and hematoma, leading to an acute vasovagal episode. manual pressure was applied and one unit of packed red blood cells was transfused to obtain hemostasis. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely patient movement since patient sat up in bed which led to oozing and hematoma at impella site. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-30241. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-30241. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7537 | 1220648-2025-30242 | 22479043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479043&pc=OZD | 01/16/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 support passed away from a stroke following impella explant. it was believed that a thrombus dislodged during explant, resulting in the noted stroke and resulting death. no further details were available from the site. | [5 empty] | Further information noted the physician does not remember exactly when the event occurred (best estimate provided) and certain details were not known. patient- and product-specific information and implant/explant dates are unavailable, not known; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7538 | 1220648-2025-30244 | 22479213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479213&pc=OZD | 06/24/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650847 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Aortic Valve Insufficiency/ Regurgitation | Mitral Valve Insufficiency/ Regurgitation | Tricuspid Valve Insufficiency/ Regurgitation | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, day 2, transesophageal showed moderate to severe aortic insufficiency, moderate mitral valve regurgitation and moderate tricuspid. the impella cp device was weaned from performance level p-7 to p-4. the patient was able to be managed without increase in drips. this same day, when discussing distal pulses, the nurse reported faint pulses. the intensivist during rounds assessed and noted left foot was cold to touch (leg with impella cp device) compared to right. all distal pulses in left foot were absent. the intensivist noted the patient's history of severe peripheral vascular disease and established a goal for the patient to wean off the impella cp device with no plans to escalate to an impella 5.5. the impella cp device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿
7539 | 1220648-2025-30245 | 22479324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479324&pc=OZD | 06/24/2025 | 08/22/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received from long-term outcome and quality indicator (loqi) impella registry that following impella assisted percutaneous coronary intervention (pci) procedure the transthoracic echocardiogram (tte) showed impella was too deep. impella was repositioned. hemodynamics immediately improved and impella was turned to p3. | The complainant reported, a patient undergoing high-risk percutaneous coronary intervention implanted with an impella cp device for mechanical circulatory support (mcs) experienced arrhythmic events requiring intervention; however, the patient would subsequently expired. the patient was admitted with complaint of shortness of breath for one week and went into pulseless electrical activity. left heart catheterization revealed tricuspid valve disease. the patient was a surgical turndown and was planned for hrpci. the impella cp device was implanted ahead of the pci that included drug-eluting stents to the left main, left anterior descending, diagonal and circumflex arteries. during the procedure, the patient experienced low blood pressure. levophed and milrinone were started with the decision to leave the impella cp implanted. the patient was sent to the unit on impella mcs. overnight, the patient went into ventricular tachycardia and was cardioverted to normal sinus rhythm. the impella cp device performance level was decreased to p-3 with plans to explant. the patient's native cardiac output was gradually improving, and the patient was weaning towards explant of impella cp device; however, the patient was noted to be in atrial fibrillation with rapid ventricular response. a bolus of amiodarone was given, and drip started. additionally, the patient was noted to have a moderate sized nosebleed and was given afrin. an otolaryngologist was consulted. earlier the next morning, the patient went into ventricular tachycardia, cardioverted three times and started on a lidocaine drip. the patient was sedated with propofol and intubated, which appeared to break the ventricular tachycardia. however, shortly later in this morning, the patient went back into ventricular tachycardia and pulseless electrical activity. cardiopulmonary resuscitation was performed for 30 minutes. the patient received multiple shocks, multiple doses of epinephrine and started on epinephrine drip and given multiple doses of amiodarone and lidocaine. the patient's hemoglobin had dropped to 8.5 from 10.2 status post nosebleed. with development of incessant ventricular arrhythmias despite extensive efforts at resuscitation, it could not be controlled and ultimately the patient expired in the midst of prolonged advanced cardiopulmonary life support attempt. |  |  |  |  | B1 updated to reflect both adverse event and product problem based on the additional information received from the clinical site. b5 updated with additional information received from the clinical site. h6 additional failure mode added based on the additional information received from the clinical site. the impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. g2 report source; study was missing from manufacturer device report 1220648-2025-30245. | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of positioning issue and vt/vf/at/af was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
7540 | 1220648-2025-30247 | 22479437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479437&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660237 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, it was noted that the impella cable was twisted and that there was reduced purge flow and no other actions were done. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and four days into patient support, the impella 5.5 access site was red, swollen, tender, and warm to the touch. the cite was being monitored closely for an infection while antibiotics were initiated and anti-coagulation was paused. support continued with the implanted pump following the intervention while patient await heart transplant. approximately 11 days later the impella 5.5 device was weaned and explanted as patient was bridged to a heart transplant that was noted to be successful. |  |  |  |  | The investigation of infection/sepsis and high purge pressure has been completed. the impella device was not returned for evaluation. high purge pressure (hpp): data logs were reviewed and only one "purge pressure high" alarm was triggered during case. real time (rt) log at time of hpp alarm shows sudden increase in purge pressure with drop in purge flow and stalling of purge flow ticks for approximately 5 minutes before suddenly resolving. clinical details noted that aic cable was twisted/kinked and causing low purge flow. cable was un-kinked/twisted and issue resolved. the cause of the high purge pressure was most likely due to a kinked purge line along the console cable based on clinical details and pattern of sudden increase then resolving of purge pressure observed in data logs. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ a3 patent gender was erroneously selected on the initial manufacturer device report number 1220648-2025-30247. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30247. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30247. h6 medical device problem code 1491 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30247. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30247. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7541 | 1220648-2025-30248 | 22479481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479481&pc=OZD | 07/03/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 07/03/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676886 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 96 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support during coronary intervention. a hematoma was noted around the arteriotomy sheath and manual pressure was applied to achieve hemostasis. accumulated blood was expressed and the hematoma softened. the patient's activated clotting time at the end of the case was 400 seconds and the integrilin infusion was discontinued. the patient was reported to be stable post procedure. | [5 empty] | The investigation for access site bleeding has been completed. clinical reported bleeding around the right arteriotomy sheath after the insertion of the guidewire repositioning unit. the pump was returned for investigation and no kinks or abnormalities were observed. the cause of the access site bleeding is most likely due to anticoagulation management, as the act was reported 400 seconds at the time of the noted hematoma. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7542 | 1220648-2025-30249 | 22479544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479544&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685637 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after being placed in surgical mode during a scheduled coronary artery bypass graft procedure, the support level of the implanted impella cp pump could not be raised above performance level p-2 without encountering suction. after being unable to raise flows to the desired level, a transesophageal echocardiogram was utilized to visualize what appeared to be a clot on the inlet cage of the pump. due to the patient's hemodynamic stability, the decision was made to explant the pump. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. low pump flow: the cause of low flow issue most likely was biomaterial ingestion as motor current (mc) spike observed in log review and transesophageal echocardiogram (tee) revealed what appeared to be clot on inlet cage of impella. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. g1 manufacturing site name, address, country, and fax number was erroneously entered on the initial manufacturer device report number 1220648-2025-30249. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7543 | 1220648-2025-30250 | 22479734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479734&pc=OZD | 06/24/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565511 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported, a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and after the implant, poor distal perfusion was noted on the impella extremity. external right to left fem-to-fem bypass placed with improved circulation to right lower extremity, and the impella mcs continued until two days later. the impella cp device was successfully weaned and explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7544 | 1220648-2025-30251 | 22479756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479756&pc=OZD | 06/29/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 07/03/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026652776 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced an uncontrolled bleed during impella cp support. the patient subsequently passed away on support due to multiple organ failure. as a result of hemorrhagic shock. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7545 | 1220648-2025-30252 | 22479829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479829&pc=OZD | 06/13/2025 | 08/11/2025 |  | 07/11/2025 |  |  | 06/13/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026638593 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented for decompensated ischemic cardiomyopathy, femoral intra-aortic balloon pump (iabp) initially placed then patient put on impella 5.5 support to continue with heart failure workup. the impella 5.5 placed under fluoroscopy per instructions for use recommendations. iabp decreased frequency and confirmed position. iabp explanted with manual pressure held, patient extubated and transferred to intensive care unit in stable condition. during support the nurse reports patient had been having intermittent suction alarms with movement that were self resolving. additionally, it was noted that the medical doctor reported the patient had been outside therapeutic range on partial thromboplastin time, intermittently the last several days, reports pump not performing appropriately, patient decompensated and needed to be cannulated for hemodynamic support. the impella 5.5 was replaced with another 5.5 and the procedure was completed successfully. there was no patient harm. | [5 empty] | The investigation for the low pump flow has been completed. after 42 days on 5.5 support, pt experienced intermittent suction alarms. the physician confirmed correct positioning and suction alarms self resolved. data log review showed suction events continued throughout the case and flows were intermittently lower than expected at p-7, dipping below the expected 3.9-4.5l/min. pump experiences a brief ps rail to 3000 mmhg with no impact to 5s parameters after 15 seconds of lowered snr and contrast before recovering. one instance of purge system blocked occurs but resolves in under 1 minute. the pump was returned and inspected where biomaterial was found wrapped around the impellor blade and in the outlet cage. the cause of the low pump flow was most likely biomaterial due to the biomaterial found on the returned pump wrapped around the impellor blade and in the outlet cage. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7546 | 1220648-2025-30253 | 22479887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479887&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539843 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 111 KG |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp and experienced access site bleeding. heparin was removed from the purge and pressure was held. care was withdrawn and the patient expired. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient movement because the bleeding was noted right after the patient transfer and hemostasis was achieved by pressure dressing. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7547 | 1220648-2025-30254 | 22479897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479897&pc=OZD | 06/21/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/21/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677279 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 121 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support developed bleeding and heparin was stopped. it was further reported that the plan was to explant the impella but when p level was reduced to p2, the patient went into ventricular tachycardia and received amio. the patient care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7548 | 1220648-2025-30255 | 22479948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22479948&pc=OZD | 06/25/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665853 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that cp was placed to support the patient with acute myocardial infarction complicated by cardiogenic shock (ami/cgs) and in need of percutaneous coronary intervention(pci). the pump was placed but there was a hematoma that developed at the site. the team did contralateral leg ballooning for hemostasis and gave rbcs. the patient survived the event/harm. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7549 | 1220648-2025-30256 | 22480056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22480056&pc=OZD | 06/25/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680700 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female | 56 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced a massive bleed from their chest tube site during impella cp support. the patient was emergently taken for a coiling of the intracostal artery to manage the bleed. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7550 | 1220648-2025-30257 | 22480145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22480145&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673904 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and lost doppler pulses. the medical team monitored the patient until she was successfully weaned and the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7551 | 1220648-2025-30258 | 22480388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22480388&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674576 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and developed ischemia as the peel away sheath was not removed. the antegrade sheath was placed to perfuse the limb. the patient was successfully weaned from the pump and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7552 | 1220648-2025-30259 | 22480775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22480775&pc=OZD | 06/25/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674560 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 78 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced significant pain to the right groin. the decision to do an ao runoff from the contra lateral side was made. the runoff revealed extravasation from the site of initial stick. the decision was made to place a covered stent. bleeding was controlled. the patient received fluid bolus. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7553 | 1220648-2025-30260 | 22481015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22481015&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671912 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 93 KG |  | White | Injury | Yes | No | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and during support, the device began showing ¿impella in unknown position¿ with dissociation of systolic waveforms. there were also signs of hemolysis with gross hematuria, elevated lactate dehydrogenase and drop in hemaglobin. the device positioning was checked multiple times and found to be in good position. the staff determined the pump should be removed. prior to removal, the automated impella controller showed flow saturation. the device was removed and a large piece of biomaterial was noted to be occluding one of the outlet cages. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7554 | 1220648-2025-30261 | 22481209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22481209&pc=OZD | 06/25/2025 | 12/04/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1130896 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death |  |  |  |  | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella cp was placed and there was a possible pump stop with a controller failure alarm. the aic was replaced. care was withdrawn and the patient expired. | [5 empty] | Data analysis: the case associated with the reported complaint was initiated on 24 june 2025. in step 3 of the case start wizard, after the initial connection of pump sn: (b)(6), it was disconnected and reconnected within a few seconds. within 1 minute and 12 seconds of the pump being started, the console displayed ¿controller error (opm communication stopped) [optical pump connected] ¿ alarm,¿ event #1043. this confirms the reported failure mode. lt logs confirmed that all signals received from the optical bench (snr, contrast, gain) were reported as +16384 values, with the ps at 3000. after that, no more samples were recorded. after the controller error, support continued for 6 minutes and 23 seconds until the pump was manually turned off by setting the p-level to p0. there was one occurrence of an impella stopped alarm, event #115, which most likely indicates the pump was disconnected. no impella stopped alarm was observed prior to the controller error. the absence of snr samples impacted the console¿s ability to recognize the pump disconnection, which prevented the user from shutting down the console in the normal way. the console was then forcefully shut down by pressing the power button for more than 60 seconds. the pump sn: (b)(6) was moved to the backup console ic9142, and support was provided for 6 days and 14 hours (until (b)(6) 2025). device analysis: this console was tested after arriving at fs. the reported failure mode ¿ controller error ¿ could not be reproduced. root cause: the root cause of the controller error was due to an aic software issue. this was entered into jama with defect/ bug id gid-dft-2927090. the root cause of the inability to shut down the console was the absence of snr samples, which was caused by an aic software issue. no impella stopped alarm was observed prior to the controller error. the impella stopped alarm that occurred after the controller error was most likely misreported as a pump stop, caused by the pump disconnection. there was no log evidence to confirm the 'controller failure (red alarm).' the 'controller error (yellow alarm)' was misreported as 'controller failure (red alarm)¿. d10 concomitant medical products and therapy dates updated. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7555 | 1220648-2025-30262 | 22481350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22481350&pc=OZD | 06/14/2025 | 08/11/2025 |  | 07/11/2025 |  |  | 06/14/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597647 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 89 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Atrial Fibrillation | Hemolysis | Tachycardia | Ventricular Fibrillation | Ectopic Heartbeat | [8 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from acute myocardial infarction was implanted with an impella cp for mechanical circulatory support (mcs). after approximately eight hours, the patient was escalated to an impella 5.5 for continuation in therapy. however, the first impella 5.5 implanted malfunctioned and therefore was explanted and replaced. the replacement impella 5.5 was positioned under transesophageal echocardiographic guidance and ramped achieving a flow of 4.3 liters per minute. the patient was sent to the unit intubated and on impella 5.5 mcs. the following morning, the patient remained intubated. however, did not require sedation at this time. an echocardiogram was completed, and ejection fraction was noted as 15-20% with right ventricular systolic function reduced. the plan was to take the patient to the catheterization lab for left and right heart catheterization. overnight the patient had required additional blood products, volume resuscitation, and multiple amps of bicarbonate. the patient was in atrial fibrillation with rapid ventricular response and sustained ventricular tachycardia. the cardiologist gave multiple amiodarone boluses with amiodarone drip, and attempted beta blockers intravenous. however, rate control was unable to be achieved. the physician also attempted to cardiovert the patient two times. the patient was eventually cardioverted back into sinus rhythm and an impella rp flex was implanted (patient now on bipella support). it was noted that since placing the impella rp flex and cardioverting, the patient has been able to wean pressor requirements. the patient was continued on bipella support per the noted plan. at this time, the patient was sedated along with intubation. the patient was noted to move all extremities with purposeful movement noted to bilateral upper extremities. throughout the day, the patient was noted to have ectopy and later that evening had two runs of ventricular tachycardia and was able to be shocked to normal sinus rhythm each time. the impella 5.5 measured 4.4 centimeters from the aortic valve annulus. the left ventricle chamber appeared narrow, and therefore, the patient was repositioned in the bed to ensure no contact with the wall. the inlet tip appeared close to the lateral wall but touching at the time of the echocardiogram. when the patient was flat on his back, the impella 5.5 inlet was in mid-ventricular space. the echocardiogram technician had completed the exam when the patient went into ventricular tachycardia and then ventricular fibrillation. the patient was shocked, and the patient remained in ventricular fibrillation despite the second shock. medication and fluids were administered, and the patient was shocked back to sinus rhythm. the impella 5.5 was slightly rotated and locked at 34.5 centimeters marking, measuring 4.2 centimeters from the aortic valve annulus. the patient was in sinus rhythm with no further runs of ventricular tachycardia. however, ectopy was still noted at times. the following day, it was noted the patient¿s heart rate was irregular. there was ectopy at times. there were no more ventricular tachycardia and ventricular fibrillation issues after the event overnight. urine was green in the foley tubing/bag and urine output has been less than 30 milliliters per hour to no urine for the last five hours. bumex drip was started overnight. creatinine continued increasing and a discussion was made for possibly starting continuous renal replacement therapy (which was started). labs were drawn and were hemolyzed. the impella 5.5 and impella rp flex were weaned by one performance level. labs were drawn again. the next day, the patient remained on mechanical ventilation and sedated with dialysis running. dressings for both impella 5.5 and impella rp flex were noted clean, dry, and intact. a chest x-ray showed the impella rp flex was in the proper position above the pulmonic valve. labs came back hemolyzed. overnight, no ectopy was and one unit of blood and 500 of albumin were given. echocardiogram this day in the morning showed left ventricular ejection fraction of around 10%. the plan of care was to continue the patient on bipella (impella 5.5 and impella rp flex) support. however subsequently, the patient¿s family withdrew care. the impella 5.5 and impella rp flex devices were both turned off along with intravenous drips and mechanical ventilation. the patient thereafter expired. | [5 empty] | Medical safety officer reviewed and determined there is not enough evidence to exclude impella 5.5 (serial number (b)(6)) as an associated factor in the patient's outcome. the is one of two reports associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-29934 for the report on the initially implanted impella 5.5 device serial number (b)(6). the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis, ectopy, and positioning issues has been completed. clinical reports runs of vt (ventricular tachycardia), ectopy, atrial fibrillation (af), and ventricular fibrillation (vf) throughout the case. pump appeared to be touching lateral wall of lv during some echos and while the patient was in different positions in bed. amio boluses, beta blockers, and magnesium were given, cardioversion was attempted, and patient was shocked multiple times throughout the case. data logs are not applicable and product was not returned. the root cause of the vt, ectopy, af, and vf was unable to be determined due to insufficient clinical details. clinical details note that the physician checked pump positioning via echo due to runs of vt. the 5.5 measured 4.4 cm from av annulus. the lv chamber appeared narrow and therefore, patient was repositioned in bed to ensure no contact with wall. inlet tip appeared close to lateral wall but was touching at the time of echo. when patient is flat on back, the inlet is mid-ventricular. the 5.5 was slightly rotated and measured 4.2 cm from av annulus. another echo the next day showed inlet measuring 4.0 cm from av annulus. data log review showed no major positioning issues except position unknown alarms during suction. the root cause of the positioning issues was most likely patient condition related as the clinical noted that the patient's lv was narrow and the patient's position in bed affected the pump's proximity to the lv wall. clinical details state that labs were hemolyzed and it is unknown if the hemolysis resolved. eventually care was withdrawn. data log review showed suction with position unknown alarms. there was also a slight downward trend in ps. the root cause of the hemolysis was not determined since product was not returned and it is unknown if/when the hemolysis resolved.
7556 | 1220648-2025-30263 | 22481406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22481406&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680113 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female | 173 KG | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a disruption in the integrity of the repo-sleeve. tegaderm was used to seal the sleeve. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7557 | 1220648-2025-30264 | 22481678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22481678&pc=OZD | 06/23/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585976 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella 5.5 support developed gastrointestinal bleed. therefore, systemic heparin was withheld. additionally, the patient became tachycardiac. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 and a5, ethnicity, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7558 | 1220648-2025-30265 | 22481759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22481759&pc=OZD | 06/20/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620658 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male | 75 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced an ischemic stroke during impella 5.5 support. no noted clots were observed via echo near or on the implanted impella following the event. support continued uninterrupted. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7559 | 1220648-2025-30266 | 22482006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22482006&pc=OZD | 06/23/2025 | 08/26/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679253 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 69 KG |  |  | Injury | Yes | Yes | Increase in Pressure | Material Puncture/Hole | Pressure Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced bleeding and hemodynamic instability. a washout was performed. during the washout there was high purge pressure alarm and blood was noted inside the purge sidearm. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of product damage (hole in catheter), high purge pressure, and vascular damage - peripheral vessel has been completed. high purge pressure: the logs showed a stepwise 3-minute increase in purge pressure, followed by a high purge pressure (hpp) alarm. purge flow dropped to 0 during that time, and the purge trend gradually recovered in 15 minutes. no motor current spikes were reported in the log. the real time log was not available on constellation during the time of the issue and could not be retrieved from the remote server. according to the clinical notes, an acute drop in purge flow (pf) and high purge pressure occurred, which was resolved after the nurse traced the line. the cause of the high purge pressure issue was most likely a kinked purge line, based on data log review. the product damage (purge tubing kink) failure mode was added as an investigated failure mode based on the high purge pressure and data log investigation. product damage (purge tubing kink): the cause of the product damage (purge tubing kink) issue was not determined since product was not returned. product damage (hole in catheter): the cause of the product damage (hole in catheter) issue was not determined since product was not returned for investigation. vascular damage: the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ h6 medical device problem code 3012 was erroneously entered on the initial manufacturer device report number. h6 medical device problem code 1504 and component code 3038 were inadvertently omitted on the initial manufacturer device report number. h10 instructions for use were incomplete on the initial manufacturer device report number. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7560 | 1220648-2025-30267 | 22482041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22482041&pc=OZD | 07/03/2025 | 09/17/2025 |  | 07/11/2025 |  |  | 07/04/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638601 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp implantation, when attempting to remove the .018 guidewire, there was resistance in removing the wire. the impella cp came out along with the guidewire into the aorta. the impella cp became bent and kinked at the cannula and the shaft near the motor. a snare catheter was inserted from the upper arm and the impella cp was stretched while pulling the impella pigtail under the clavicle. the .018-inch guidewire still could not be removed from the impella cp. after straightening, the impella cp was removed at the same time as the wire. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. only cp pump was returned for review. visual inspection found multiple kinks on can. no other issues were found on the rest of the pump. delivery issues: guidewire was not returned for review. the cause of delivery issue was unable to be determined as limited clinical details were provided and no guidewire was returned for review. cannula kink: the cause of the cannula kink was most likely an interaction with the guidewire per clinical report. d9 device returned to manufacturer date added.
7561 | 1220648-2025-30268 | 22482084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22482084&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026693880 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient had a cp pump placed for the support of a high risk percutaneous coronary intervention (hrpci) when delivery failed. despite all troubleshooting measures, wire exchanges, shockwave and ballooning, the pump could not traverse the anatomy and deliver to the lv. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since inability of pump to overcome vascular resistance despite angioplasty and shockwave observed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7562 | 1220648-2025-30269 | 22482311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22482311&pc=OZD | 06/25/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, there was access site bleeding with a hematoma. the pump was explanted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
7563 | 1220648-2025-30270 | 22482348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22482348&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  | 08/31/2025 |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388663 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted in the patient pointed toward the mitral valve. the pump was removed and reimplanted successfully. later, the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of positioning issue most likely was wireless insertion due to improper technique since the impella device is not indicated for wireless insertion. b1 and h1 were erroneously selected on the initial manufacturer device report number 1220648-2025-30270. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7564 | 1220648-2025-30271 | 22483345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22483345&pc=OZD | 06/23/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/23/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026683573 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 59 KG | Non Hispanic |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella 5.5 was implanted in a patient and was bent in the wrong direction. the physician attempted to reposition the pump and a left ventricle perforation was noted. the ventricle was repaired. later, the patient exhibited bleeding from the chest tubes and a washout was performed. five units of packed red blood cells and three units of fresh frozen plasma were planned to be given. two days later, the plan was for the patient to go to the operating room for a washout. the following day, the patient was taken back to the operating room to fix the left ventricle perforation that was not properly fixed. ultimately, care was withdrawn that day and the patient expired. | [5 empty] | The investigation for the ventricle perforation and the positioning issue has been completed. clinical details report that the pump was noted to be angled in the wrong direction, so they attempted to rotate/reposition. during this event, the ventricle was perforated. the pump was put in surgical mode and the perforation was repaired. it was also reported that the pump was initially positioned shallow intentionally because the patient had a small lv. data logs were reviewed and just after the pump was started, the position in aorta alarm triggered for ~20 seconds. two minutes later, the pump was put in surgical mode. this position in aorta alarm likely occurred while the physician was attempting to reposition/manipulate the pump, based on how data logs align with the clinical timeline reported. product was not returned for investigation. based on the clinical detail that the pump was intentionally positioned shallow which likely then made it more difficult to reposition the pump, the most likely root cause of the ventricle perforation is patient condition, for the positioning issues, clinical details report that the pump was intentionally positioned shallow due to the patient's anatomy (small lv). shortly after initial positioning, it was noted that the pump was angled in the improper direction, so the physician attempted to reposition. during this repositioning attempt, the ventricle was perforated, and after the perforation was repaired, the physician elected to intentionally position the pump shallow again due to the patient's anatomy. data logs were reviewed and just after the pump was started, the position in aorta alarm triggered for ~20 seconds. two minutes later, the pump was put in surgical mode. this position in aorta alarm likely occurred while the physician was attempting to reposition/manipulate the pump, based on how data logs align with the clinical timeline reported. product was not returned for investigation. based on the clinical details provided, there was difficulty initially positioning the pump and subsequently repositioning the pump due to the patient's small lv. therefore, the root cause of the positioning issues was most likely patient condition. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7565 | 1220648-2025-30272 | 22484059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484059&pc=OZD | 06/24/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681800 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Hypovolemic Shock | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient had an impella cp implant, prior to the implant and the patient had very small vessels with both legs very cool and clammy with no distal pulses found by doppler. the staff proceeded with the impella implant and a distal perfusion catheter was placed after the implant. it was further reported that the patient got hypotensive overnight and team believed she was hypovolemic therefore packed red blood cells/platelets/fresh frozen plasma were given. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7566 | 1220648-2025-30274 | 22484247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484247&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679382 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Fracture | [5 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced suction alarms. the pump was repositioned and a bolus of albumin was given to resolve the alarms. the patient also experienced ventricular tachycardia, received defibrillation, transitioned to atrial fibrillation, and was then treated with lidocaine and amiodarone. due to signs of anemia the patient was transfused platelets, fresh frozen plasma, and packed red blood cells. upon explantation the pump broke. all pieces were removed from the patient. the patient survived. | [5 empty] | The investigation of low pump flow, positioning issues, vascular damage - peripheral vessel, vf/vt/af/at, and product damage (detached inflow/outflow ¿ great vessel has been completed. damage (detached inflow/outflow - great vessel): there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. due to a lack of sufficient evidence during the removal of the 5.5 and lack of product returned, the root cause of the product damage (detached inflow/outflow) cannot be determined. vt/vf: on the 27th, the patient had runs of vt. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. positioning issue: data logs show "impella position unknown" alarms at the beginning of the case. due to lack of clinical details provided, it is unknown why the pump was deep so the root cause of the positioning issues cannot be determined. low pump flow: the data logs show that there was reduced pulstaility in the motor current (mc) and placement signal (ps), and the suction alarms were accurately triggered as the patient was in diastolic suction throughout support. there were no motor current spikes or any indication of an ingestion related event. based on the clinical details and data log analysis, the root cause of the low pump flow is most likely due to the patients condition as they were having volume related issues throughout support. vascular damage - peripheral vessel: the clinical details state there was bleeding on the 26th in which 3 units of prbc's were given. due to lack of clinical details stated, the root cause of the vascular damage - peripheral vessel cannot be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7567 | 1220648-2025-30275 | 22484291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484291&pc=OZD | 06/24/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688964 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that bleeding was noted while dressing the access site following impella cp implant. the staff tightened the perclose, added a suture, and applied manual pressure but the bleeding worsened. vascular was consulted and the decision was made to cutdown the right femoral access site and place a new pump via the left femoral side. hemostasis was achieved following pump removal. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7568 | 1220648-2025-30273 | 22484316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484316&pc=OZD | 07/01/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 07/01/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637975 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with advanced heart failure was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient developed a hematoma at the surgical site. heparin was withheld due to the hematoma. the patient was taken to the operating room for hematoma evacuation and a washout. the patient remained on impella support. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7569 | 1220648-2025-30276 | 22484389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484389&pc=OZD | 06/25/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/25/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678520 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 117 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella 5.5 support showed a thalamic infarct on head ct. the patient underwent an embolectomy. the pump was positioned in the papillary muscle and the patient underwent a washout and repositioning of the pump. the family decided to withdraw care and the patient expired. | [5 empty] | The investigation of stroke/cva, positioning issues, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: & impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7570 | 1220648-2025-30277 | 22484505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484505&pc=OZD | 06/24/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659975 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 88 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient had an attempted impella 5.5 implant and the pump was unable to advance into the left ventricle (lv). the doctor removed the impella 5.5 and was going to reattempt access using the impella wire. during this time the patient's extracorporeal membraneoxygenation (ecmo) flows were not sustaining greater than one liter. echocardiogram showed no volume in the heart. the physician discussed options with perfusion and anesthesia regarding additional volume and volume requirements. mean arterial pressure had dropped significantly and ecmo flows remained low, despite volume and pressers. blood pressure continued to drop. the patient expired at this time. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7571 | 1220648-2025-30278 | 22484534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484534&pc=OZD | 06/18/2025 | 10/31/2025 |  | 07/11/2025 |  |  | 06/26/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143527 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal of an implanted impella cp was intermittent during patient support. no intervention was required and support continued uninterrupted. upon further engineer investigation, it was believed that the optical signal issue may be related to a loss of opm data associated with the automated impella controller. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. analysis: cp pump (b)(6) was connected to ic2714 on (b)(6). the rt logs show raw optical sensor fast toggled between ~80mmhg and -1638.4mmhg for the duration of the case. the lt logs for the whole case show that at the beginning of the case, contrast was switching up to 5000% as placement signal intermittently was at 3000mmhg. this behavior was repeated on the next pump that was run on this console. the console was not returned for investigation. root cause: the cause of the optical signal issue was not determined since the console was not returned but was most likely due to the console as the behavior was repeated across pumps. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. | A4 is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7572 | 1220648-2025-30279 | 22484591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484591&pc=PYX | 06/24/2025 | 08/21/2025 |  | 07/11/2025 |  |  | 06/24/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025599007 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female | 121 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis during impella rp flex support. positioning was confirmed to be correct and the condition was verified via measured hemoglobin levels. blood products were given to treat the condition and support was continued uninterrupted. the patients outcome is unknown. | [5 empty] | The investigation for hemolysis has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product other than clamp marks on the cannula. no further investigation was performed on returned product due to severe cannula damage. data logs show no abnormalities related to motor current spikes, positioning issues, or placement signal trends during the entire case run. as per the clinical report, on the (b)(6), hemolysis was reported after lvad implant and showed some resolution the next day; however, it continued until the end of support. multiple blood products were infused during the reported hemolysis event. the cause of the hemolysis issue was not determined since the provided clinical details and data log were not conclusive without performing the hemolysis test. the pump could not be sent for hemolysis testing due to severe damage to the cannula, which could falsify the results. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7573 | 1220648-2025-30280 | 22484930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22484930&pc=OZD | 07/01/2025 | 07/11/2025 |  | 07/11/2025 |  |  | 07/01/2025 | 07/11/2025 |  |  |  |  | 07/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026683582 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient had intermittent episodes of cardiac arrest prior to impella placement. the patient was transferred to the intensive care unit in critical condition. he remained in persistent ventricular tachycardia (vt) and ventricular fibrillation (vf). although he was shocked and briefly converted, he repeatedly reverted to vt/vf. following clinical deterioration, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7574 | 1220648-2025-30181 | 22466775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22466775&pc=PYX | 06/23/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026639548 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 183 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that three days into impella rp flex support, placement signal readings were noting irregularities with the pump coinciding with low pump flows and high motor current. shortly after, a pump stop occurred for which attempts to restart the pump were not performed. plans were subsequently made to remove the device. no patient harm was reported. | [5 empty] | The investigation of pump stop has been completed. pump stop: upon visual inspection, biomaterial present throughout pump cannula and in both inflow and outflow cage. pump was soaked in tergazym and attempted to be run. pump was able to turn on and run at p-9 for approximately 20 minutes. motor current at start of pump run was low and slightly variable with some spikes in motor current before pump was able to gradually increase to stable values. no pump stop alarm was triggered during pump run in the lab. data logs were reviewed and the logs showed a decrease in pulsatility on the morning on the last day of support with no change in p-level. later on the same day, a trend up in motor current was observed with large variability in placement signal and decrease in pump flows were observed. spike in motor current occurred at time of pump stop alarm. clinical details noted a trend in motor current and pump flows prior to pump stop occurring. pump was not attempted to be restarted and was explanted. the root cause of the pump stop was most likely due to biomaterial buildup within the pump based on returned product showing evidence of biomaterial around inflow cage and data log analysis with trend up in motor current prior to pump stop. thrombus: based on the returned product investigation, an additional failure mode of "thrombus" was added to the investigation. biomaterial was present within the cannula of the pump on returned product. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
7575 | 1220648-2025-30182 | 22466819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22466819&pc=OZD | 06/23/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687162 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp and experienced ectopy. the pump was repositioned to resolve the issue. the patient also experienced bleeding from the groin site and suture. one unit of packed red blood cells was given and a femstop was placed to obtain hemostasis. the patient was successfully weaned and the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7576 | 1220648-2025-30183 | 22466886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22466886&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  | 06/30/2025 |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360207 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 76 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during attempted implantation of an impella cp, the catheter would not track over the aortic arch due to heavy calcification and tortuosity. after multiple unsuccessful attempts, implantation was aborted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the impella catheter wouldn¿t track over the aortic arch due to heavy calcification, tortuosity. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7577 | 1220648-2025-30184 | 22466948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22466948&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313338 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) had a controller error alarm flash and the patient lost placement signals during each one. this issue occurred at least twice. the aic was replaced to continue patient support. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) and optical signal issue has been completed. data analysis: on the complaint date (b)(6) 2025, the console displayed ¿controller error (opm comm crc invalid) [optical pump connected] - alarm,¿ event #1045. this confirms the reported failure mode. the real time (rt) logs confirmed that all signals received from optical bench (snr, contrast, lamp, gain) are represented as -16384 values and the ps was 3000 due to the crc error. device analysis: this console was tested after arriving at fs. the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7578 | 1220648-2025-30185 | 22466991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22466991&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686553 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 150 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was unable to advance beyond the central extracorporeal membrane oxygenation (ecmo) cannula in the ascending aorta during attempted delivery. after multiple failed attempts, the decision was made to abort the implant and leave the patient on venoarterial ecmo and intra aortic-balloon pump support. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was most likely an interaction with another device per clinical that states the pump was unable to pass across the central extracorporeal membrane oxygenation (ecmo) cannula in the ascending aorta. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7579 | 1220648-2025-30186 | 22467140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22467140&pc=OZD | 06/23/2025 | 08/13/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672699 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 99 KG | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Pumping Stopped | Infusion or Flow Problem | Incomplete or Inadequate Priming |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp and experienced air in the purge due to a crack and there was a purge sidearm leak. the patient rolled on the pump causing the pump to stop. a new impella cp was used to continue patient support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | Purge leak pump: no product was returned to confirm leak location on sidearm. the root cause of the purge leak - pump was determined to be due to a damaged sidearm as evidenced by clinical details but sidearm leak location is unknown. pump stop: clinical reported purge pressure high and motor current pump stop after leak. the data logs show a sudden drop in purge pressure and rise in purge flow. this is as a result of the purge leak. moments later there is a sudden rise in purge pressure and an upward trend in motor current likely as a result of blood ingression into motor. purge pressure remains high and there are motor current spikes with speed deviation outside p-level changes characteristic of biomaterial interaction. there are purge system blocked alarms before impella stopped - motor current high alarms. the root cause of the pump stop was most likely due to blood ingress as evidenced by log analysis showing high purge pressure and rise in motor current after purge leak and alignment with clinical timeline priming issue: clinical reported 'air in purge' alarm. the data logs show a sudden drop in purge pressure and rise in purge flow. this is as a result of the purge leak. the 'air in purge' alarm triggers before a 'purge system open' alarm. the alarm resolves after about 45 minutes as clinical mention purge sidearm bypass was attempted. logs show a purge cassette change initiation that was not completed. limited clinical information related to other troubleshooting attempts related to purge and priming were mentioned. the root cause of the priming issue was not determined as limited clinical information was provided and no product was returned. saturated flow: clinical reported high motor current and flow saturation. the data logs confirm flow saturation with pump flowing at about 4.0l/min at p7. flow saturation is likely as a result of biomaterial which interacted with pump as a result of lack of purge fluid flowing through purge gap due to sidearm damage. continuous high purge pressure supports this. the logs also show motor current spikes with speed deviations outside of p-level changes characteristic of ingestion. the root cause of the saturated flow was most likely biomaterial as evidenced by the data log analysis showing high purge pressure and motor current spikes and the aligning clinical information provided thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7580 | 1220648-2025-30187 | 22467244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22467244&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597837 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was placed to support a patient post supraventricular tachycardia and atrial fibrillation. the pump was placed and was supporting when the pump and the patient were transferred from community to a tertiary center. the console had power off issues but this did not cause any harm. upon explant of the impella cp, the access site had issues achieving hemostasis. multiple closure devices and another sheath had to be placed to achieve hemostasis, as well as vascular surgeon closure. the patient survived and is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7581 | 1220648-2025-30188 | 22467319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22467319&pc=OZD | 06/23/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681635 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed hemolysis during impella cp support. the patient's urine was pink/dark amber so the pump support level was reduced and the pump was repositioned. support continued with the implanted pump following the actions taken to manage the conditions. the patient was successfully weaned and the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7582 | 1220648-2025-30189 | 22467379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22467379&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654225 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was being prepped and a purge pressure low alarm appeared. leaking was observed inside the purge cover from the luer lock connecting the purge reservoir and the purge filter. a new pump was primed and used successfully. no patient harm was reported. | [5 empty] | The investigation of purge leak ¿ pump has been completed. priming of the pump, purge pressure (pp) low alarm occurred. purge leak noted inside purge cover. data logs were not recovered. the pump was returned and inspected. priming the pump was able to recreate the purge leak at the filter vent holes on the sidearm filter where purge fluid was able to leak out of both holes. the cause of the purge leak was sidearm filter damage where purge fluid leaked out for the filter vent holes. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
7583 | 1220648-2025-30190 | 22467939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22467939&pc=OZD | 06/19/2025 | 08/26/2025 |  | 07/10/2025 |  |  | 06/19/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675082 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp placement was aborted due to the patient¿s anatomy. patient's outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following : ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7584 | 1220648-2025-30191 | 22468306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22468306&pc=OZD | 06/19/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/19/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682088 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported patient had a cp pump placed to support the high risk cardiac procedure patient. the pump was placed and allowed for the angioplasty and was explanted at just over 1 hour of support. the placement was complicated by plaque in the native anatomy. the vessel required vascular repair and stented with a self-expanding stent. flow was resolved and patient survived the harm. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7585 | 1220648-2025-30192 | 22468387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22468387&pc=OZD | 06/20/2025 | 08/16/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662127 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Defective Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the pump kept jumping deep and jumping back out of the aorta. the impella was left deep. additionally, during support, the patient had bleeding at the access site. the staff did not know if the sheath was cracked or just leaking. one unit of blood products were given to the patient. the pump was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding, introducer damage ¿ mechanical, and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: the clinical states that the pump started with out issue but the impella kept jumping deep and back out of the aorta. it was noted that the percloses failed. the data logs show "placement signal (ps)low" alarms with slight variations in flow, motor current (mc) and ps during the time of the reported positioning issues. there were no "impella position unknown" or "impella position in aorta," alarms seen. due to lack of clinical details around the securing of the pump and what was occurring during that timeframe, the root cause of the positioning issues cannot be determined. introducer damage: it was stated that it was unknown if the sheath was leaking or if the groin was bleeding around the sheath. however, it states that the 2 percloses failed. due to lack of clinical details and no product returned, the root cause of the introducer damage - mechanical cannot fully be determined as damage cannot be noted. access site bleeding: based on the clinical details stating that there was bleeding and the perclose devices failed, the root cause of the access site bleeding is most likely due to use. d2 common device name was erroneously entered on the initial manufacturer device report number 1220648-2025-30192. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30192. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7586 | 1220648-2025-30193 | 22468642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22468642&pc=OZD | 06/20/2025 | 08/07/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647300 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support the patient with a st-segment elevation myocardial infarction. the cp was escalation from an intra-aortic balloon pump. the support was not complicated but, the patient pulled back the pump and the team lost support in the left ventricle. the pump was removed and replaced due to the malpositioning. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the product was returned and no physical abnormalities were found during visual inspection. the root cause of positioning issue was most likely patient movement related since the patient was disoriented and pulled his right fem impella triggering impella in aorta alarm. h.10 additional manufacturer narrative was incorrect on the initial manufacturer device report number 1220648-2025-30193. it stated the device was not returned and it was.
7587 | 1220648-2025-30194 | 22468786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22468786&pc=OZD | 07/02/2025 | 08/29/2025 |  | 07/10/2025 |  |  | 07/02/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026694360 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that that the impella cp was inserted without any issues but once turned on the motor current was flat, and flows remained very low. the provider stated that the impella catheter was kinked. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow and the product damage has been completed. clinical details statse that when impella was turned on, flows were reported at 0.3/0.5 l/min, and the provider stated that he kinked the cannula. the product was not returned for investigation nor were any data logs. the cause of the low pump flow was most likely due to the cannula kink, which was reported in the clinical details. the cause of how the cannula kink occurred cannot be determined due to lack of clinical details.
7588 | 1220648-2025-30195 | 22468834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22468834&pc=OZD | 06/20/2025 | 08/07/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686401 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 99 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered low pump flows and thrombus. there was a lot of difficulty delivering the pump through the anastamosis and through subclavian stenosis into the ascending aorta. the activated clotting time was 262 at the time of insertion, but the pump was in the graft for an extended amount of time while being manipulated down the artery. the pump was finally delivered into the left ventricle (lv) and in an ideal position in the midventricular space. once the pump was started, however, flows would not go over 2 l/min and it would alarm for suction at p5 and higher. the pump was removed and a clot was visualized on the outlet. another pump was placed and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of thrombus and low pump flow has been completed. the impella 5.5 was returned for evaluation. upon visual inspection, no biomaterial or abnormalities with the pump were noted. the pump was run at p-9 for approximately five minutes and low pump flows did not reproduce. data logs confirm low flows occurring upon pump initiation. no significant motor current spikes were present during pump run to indicate ingestion. clinical details noted that with difficulty during insertion, the pump was left indwelling in the graft for an extended period of time. additionally, clinical details noted that the pump had a clot present at the outlet cage upon explant. the root cause of the low pump flow could not be definitively determined as no abnormalities were noted on returned product and limited clinical details were provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. a.4 weight was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30195. e.1 added fax number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30195. | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7589 | 1220648-2025-30196 | 22469002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469002&pc=OZD | 06/20/2025 | 08/15/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686398 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 99 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented in decompensated heart failure. started on dobutamine, echo showed ejection fraction of 15%, and decision was made to place impella 5.5 as bridge to advanced therapies. however the first 5.5 had flow and suction issues that was due to a thrombus on the outlet. so a second 5.5 was attempted however was unsuccessful at getting through the subclavian stenosis. the only wire used was the impella 0.018 wire. it was noted that the second 5.5 attempted on the patient in the same procedure setting. the physician used an new 0.018 impella wire and had good purchase in the left ventricle. when advancing the 5.5 it would get caught at the subclavian/inominate junction and he could not make the turn into the ascending aorta. no buddy wire, different 0.018 wire, or arm manipulation was attempted. ultimately a cp was placed. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was returned for evaluation. product damage (catheter kink) was added based on the returned product investigation delivery issues and product damage (catheter kink). the clinical stated that "while advancing the 5.5 it would get caught in the subclavian/innominate artery due to arterial stenosis and resistance met. the pump was unable to be advanced past it." the returned product had a kink in the catheter near the motor housing. no other damages were observed. based on the clinical information, the root cause of the delivery issues is most likely due to the patients condition as the pump was unable to advance past the stenosis. the root cause of the product damage (cannula kink) is due to the patients condition as resistance was met causing the pump to not be able to be advanced further. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7590 | 1220648-2025-30197 | 22469003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469003&pc=OZD | 07/02/2025 | 09/03/2025 |  | 07/10/2025 |  |  | 07/03/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during use the automated impella controller (aic) suddenly blacked out and after 10 seconds, restarted. the aic was replaced to successfully complete the case. there was no harm to the patient. | [5 empty] | The console was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. console logs from the reported day of event are consistent with complaint and show that after 28 days of uninterrupted hemodynamic support a software process became unresponsive, and console performed partial reset (as per design). during the restart, pump was paused for less than 30 seconds, after which pump automatically resumed operation. support continued for another 38 hours, after which the pump was manually stopped, presumably to be transferred to a backup console. aic logs also indicated multiple interruptions of data streaming. console was tested, but the issue was not reproduced, and no unresponsive sw processes were observed. there is no evidence that would indicate any relation between unresponsive process and interruptions of data streaming. unexpected reboots of the rlm were most likely caused by som microsd card corruption, however root cause of its corruption could not be determined.
7591 | 1220648-2025-30198 | 22469086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469086&pc=OZD | 06/20/2025 | 08/15/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686506 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance during delivery due to a large plaque burden in the subclavian artery. the implant was aborted and an impella cp was placed for planned support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, 5.5 device could not be advanced past plaque burden in subclavian artery. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿.
7592 | 1220648-2025-30199 | 22469297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469297&pc=OZD | 06/24/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/24/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 107 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported, a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. following the placement of an impella cp repositioning sheath, the patient developed an active arterial bleed and hematoma at the access site. angle matching was performed, manual pressure was held, but the condition persisted. the impella cp was therefore explanted in response to the uncontrolled bleeding. hemostasis was achieved with two perclose, and hematoma resolved. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7593 | 1220648-2025-30200 | 22469396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469396&pc=OZD | 06/19/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655216 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced ventricular tachycardia requiring defibrillation and intubation. the patient was positive for heparin induced thrombocytopenia and was treated with bivalirudin. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7594 | 1220648-2025-30201 | 22469480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469480&pc=OZD | 06/20/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/20/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661029 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male | 92 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5 the patient experienced a gastrointestinal bleed. systemic heparin was withheld. impella support continued and no patient harm was reported. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7595 | 1220648-2025-30202 | 22469577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469577&pc=OZD | 06/23/2025 | 09/09/2025 |  | 07/10/2025 |  |  | 06/24/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645470 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and approximately three days into the support, the motor current repeatedly increased, and the pump stopped. the purge flow decreased, and the purge pressure increased. after several unsuccessful restart attempts, the decision was made to explant the device which was completed. after removal, manual compression was applied and the patient was reported to be in stable condition. | [5 empty] | The investigation of pump stop, saturated flow, and high purge pressure has been completed. the impella device was not returned for evaluation. data log shows the motor current gradually increasing to 1100 followed by pump stop. multiple unsuccessful attempts were made to restart the pump, with reported pump stop and motor current high alarm. purge pressure gradually increased with fluctuating trends before the motor current started rising. purge pressure never reached the trigger for high purge pressure but remained high during the pump stop event. additionally, pump flow gradually became saturated during the upward motor current trend. no sudden spike in motor current was reported during the case, which rules out an ingestion event. the cause of the pump stop issue was most likely biomaterial buildup, based on data log review. the cause of the saturated pump flow issue was most likely biomaterial buildup, based on data log review. the cause of the high issue was most likely biomaterial buildup, based on data log review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ mpella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7596 | 1220648-2025-30203 | 22469732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469732&pc=OZD | 06/21/2025 | 08/16/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657728 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 91 KG |  |  | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia requiring defibrillation. due to symptoms of low pump flow, a chest tube was placed, blood products were transfused, and systemic heparin was withheld. the patient was successfully weaned and the pump was explanted. | [5 empty] | The investigation of vascular damage - peripheral vessel, low pump flow, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. vascular damage: product was not returned for investigation. since insufficient clinical details were provided, the root cause of the vascular damage could not be determined. low pump flow: product was not returned for investigation. clinical details provided that the patient had a surgical bleed that aligned with worsened hemodynamics and suction alarms. hemodynamics improved when fluid was drained and replaced. suction and position unknown alarms in data log review aligned with the provided timeline. for these reasons, the root cause of the low pump flows was most likely patient condition. vf/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30203. h6 medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2025-30203. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7597 | 1220648-2025-30204 | 22469733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469733&pc=OZD | 06/23/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674726 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced ventricular tachycardia. overnight, the patient was placed on extracorporeal membrane oxygenation (ecmo) for the worsening arrhythmias, atrial fibrillation with rapid ventricular response and instability despite the impella cp in place. the patient continued on impella mcs for approximately 11 days. the report noted the patient improved through ecmo decannulation followed by impella cp weaning and explant a week later with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7598 | 1220648-2025-30205 | 22469772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22469772&pc=OZD | 06/21/2025 | 07/14/2025 |  | 07/10/2025 |  |  | 06/21/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1838437 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | Defective Device | Connection Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after ambulating a patient around the unit, the utilized automated impella controller would not resume charging once plugged back into an outlet. multiple outlets were tried, but the issue persisted. upon further analysis, it was noted that the plug appeared to have separated from the cord itself. the console was swapped in response to the noted issue. no patient harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: the case start on 11th april 2025, the ac power was connected, as ac power alarms was cleared multiple times through out the case. however, on 21st june 2025, the "ac power disconnected - alarm" (event # 109) was triggered and could not be cleared until the console was shut down. additionally, the console was turned on multiple times and "ac power disconnected - alarm" (event # 109) was seen. the lt log indicates that the batteries were actively charging until the occurrence of event # 109. see lt_power.jpg. device analysis: the console (b)(6) was returned to fs for further investigation. upon inspecting the aic, an loose neutral (white) wire was discovered inside the ac cord plug. further examination of the plug revealed that the wire was broken, and was completely detached from the terminal. see exposedwire.jpg, and loosewire.jpg. it appears that the wire covering or outer sheath was improperly cut, as per the vendor wiring instructions. clamp marks are visible on the other sheath but was very light and not properly torqued, which likely contributed to the wire becoming loose and detaching from the plug. see broken_white.png, length.png, and wiring instruction interpower.pdf. additionally, the batteries were checked, and no charging issues were found. the console was also tested with a known good pump and purge cassette, and no issues were observed. note: the supplier quality team was made aware of the complaint and the investigation findings. see email.pdf root cause: the reported complaint about the console not resuming charging charge even connected to the wall outlets was due to the neutral wire inside the ac cord plug was broken and had become disconnected from the terminal. corrective action: events of this type will be monitored for trends. pia-2025-0000337 was previously opened in response to several aic ac power cord manufacturing defects of the same type. escalation should not be reinitiated based on this case alone. before returning the console to the customer, the following actions are instructed to perform: replace ac power cord set (6000-0171) due to defective. perform section 18, 19, and 20 from preventive maintenance procedure (0042-7309). perform a full functional check on the console and optical system (0042-7316).
7599 | 1220648-2025-30206 | 22470203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470203&pc=OZD | 06/21/2025 | 08/04/2025 |  | 07/10/2025 |  |  | 06/21/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665339 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female | 67 KG |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that nineteen days into impella 5.5 support, the purge flow dropped while the purge pressure increased to 1000 mmhg. lysis therapy was discussed but ultimately declined due to the patient's existing cerebral hemorrhage. information pertaining to the bleed and it's origin was not provided. the staff was advised to replace the pump and decisions regarding to further treatment were pending. | [5 empty] | Upon further review, manufacturer device report number 1220648-2025-30206 was found to be a duplication of manufacturer device report number 1220648-2025-29645. all further investigation findings and additional information will be submitted on manufacturer device report number 1220648-2025-29645. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7600 | 1220648-2025-30207 | 22470405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470405&pc=OZD | 06/22/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661000 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 12 days into impella 5.5 support, a bleed developed at the access site. sandbags were placed to apply pressure and heparin was temporarily halted to manage the condition. support was continued uninterrupted following the intervention and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anti-coagulation management because hemostasis was achieved by reducing heparin.
7601 | 1220648-2025-30208 | 22470418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470418&pc=OZD | 06/21/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653620 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 91 KG | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced multiple complications during impella 5.5 support. it was documented that two days into support, the patient experienced a gastrointestinal bleed. the patient remained on heparin despite the noted condition. the following day, a ct scan revealed that the patient also had multiple head bleeds. at that point in time, anticoagulation was halted. additional intervention was not noted. after four days of support, the sterile sleeve became loose, at which point the rn was able to clean and resecure the component. support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7602 | 1220648-2025-30209 | 22470478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470478&pc=OZD | 06/13/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/13/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668090 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured nonsustained ventricular tachycardia during impella 5.5 support. the pump was repositioned in response to the condition and support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7603 | 1220648-2025-30210 | 22470578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470578&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681702 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 patient experienced ventricular tachycardia during ambulation two days into support. one shock was given to manage the condition and the pump was repositioned under echocardiogram accordingly. the patient remains on support. | [5 empty] | The impella device was not received from the customer as it continues to support the patient, therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7604 | 1220648-2025-30211 | 22470604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470604&pc=OZD | 06/16/2025 | 10/03/2025 |  | 07/10/2025 |  |  | 06/16/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666571 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 64 KG | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | Pumping Stopped |  |  |  |  | Unspecified Vascular Problem | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with ami/cgs was admitted. lesion for left main and underwent percutaneous coronary intervention (pci) with intra-aortic balloon pump (iabp) assistance. there was hemodynamic disruption during pci and extracorporeal membrane oxygenation (ecmo) was added. additionally, there was difficulty withdrawing from ecmo plus iabp, and the patient was transported from iabp on monday morning for the purpose of impella upgrade and an impella cp was placed. impella cp pump 1: during the transfer to the intensive care unit on a bed, tension was applied to the tip, causing it to slip out to the aorta. when tension was applied, the connection part of the sterile sleeve was damaged. the device was replaced (same day as implanted). impella cp pump 2 was implanted and 20 days after, the pump stopped. the pump was restarted and ran at low flow rate (p-2) for a while. per the physician it did not affect hemodynamics significantly. however, there was a possibility that the patient would still become unstable without mechanical circulatory support. the patient was escalated to an impella 5.5. the patient expired 28 days later on the impella 5.5 noted from multiple organ failure. | It was reported that while priming an impella cp a blockage appeared. the purge cassette was replaced and the issue resolved. then, a purge cassette leak occurred. again, the cassette was replaced and the issue resolved. later, the pump stopped. the pump was restarted and low pump flows were observed. the pump was explanted and replaced with an impella 5.5. |  |  |  |  | The investigation of temporary pump stop and low pump flow has been completed. only the purge cassette was returned for investigation. temporary pump stop: after 18 days on cp support, pump stop occurred. the cause of the temporary pump stop was not determined since the pump was not returned for review. the temporary pump stop occurred beyond the 8 day impella cp indication for use per design trace matrix. low pump flow: the failure mode was changed from low pump flow to saturated flow since the pump flow was high at 3.0 l/min at the respective p-level p2. the cause of the saturated flow was not determined since the pump was not returned for review. the low pump flow occurred beyond the 8 day impella cp indication for use per design trace matrix. medical safety review of the event was completed. impella cp pump 1, the event describes a malpositioning of the device and anti-contamination sleeve damage. pump 1 was therefore replaced. this malfunction is not attributable to the demise outcome. however, it is a product malfunction and should be reported. impella cp pump 2 involves a malfunction with pump stop that was restarted with inadequate flow. hemodynamics were not affected significantly per the reporting from the physician. the pump was replaced. the patient expired approximately 28 days later. after being upgraded to the 5.5, an attempt to stabilize the patient's hemodynamics were made but there was little improvement; the patient died of multiple organ failure. in this case, the patient's underlying condition likely contributed most to an outcome of death (cardiogenic shock stage e). however, there is not enough evidence to exclude the impella cp pump 2 as an associated factor given the pump stop malfunction in this complex case. additionally, this device malfunction if occurring in another patient was repeated (as in the case above) could result in significant hemodynamic decompensation and therefore, should be reported as a death type of reportable event. impella 5.5: no complaint ¿, the upgraded to 5.5 and tried to stabilize his hemodynamics, but there was little improvement, and he died of multiple organ failure.¿ this event story involves two product complaints: impella cp sn (b)(6) pump 1 (25-45645-1): malfunction (mdr 1220648-2025-29850). impella cp sn (b)(6) pump 2 (25-45618-1): death (mdr1220648-2025-30211). instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped : ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30211. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30211. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30211. d8/d9 should have been selected as no as the pump was not returned, only the cassette h1 type of reportable event was erroneously selected on the initial manufacturer device report number 1220648-2025-30211. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-30211, as the pump was not received for investigation h6 health effect - clinical code 4582 and medical device problem code 1250 was erroneously entered on the initial manufacturer device report number 1220648-2025-30211. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30211. h10 incorrect instructions for use included on the initial manufacturer device report number 1220648-2025-30211. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
7605 | 1220648-2025-30212 | 22470684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470684&pc=OZD | 06/21/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/21/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674587 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported impella cp placed for support for non-st- segment elevation myocardial infarction. it was noted the patient had no pulses in impella leg and extremity, knee slightly mottled. no ultrasound done yet, md seemed to think patient very clamped down at this time as there was apparently flow in the leg when first inserted. the patient was taken to the operating room for washout and chest closure. again, noted that still no pulses on impella extremity however had slow cap refill. impella was weaned down to p4. additionally, the patient went into atrial fibrillation with rapid ventricular response. the impella was successfully weaned and explanted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7606 | 1220648-2025-30213 | 22470763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470763&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651752 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 57 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient presented to the emergency department with st-elevation myocardial infarction. the impella cp was primed and placed under imaging guidance. waveforms and flows were appropriate. it was noted that after the peel away sheath was removed and repositioning sheath was in place, the physician placed a forward tension suture and angle matched the catheter. oozing was noted as the site was dressed. activated clotting time was 357 and the decision was made to move the patient to the unit and continue to monitor. upon arrival to the unit, the start of a hematoma was noted. pressure over the impella site was held to mitigate the spread of the hematoma and stop the oozing. this was unsuccessful. dressing was taken down and redressed with new angle, more pressure, quick clot, sandbag and purse string suture were all utilized to help quell the bleeding to no avail. the physician was visibly upset and the decision was made to remove the impella cp due to the bleeding and place an intra-aortic balloon pump. one unit of red blood cells was given. vascular surgery was consulted for potential cut down on the artery. the patient's outcome was critical. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7607 | 1220648-2025-30214 | 22470780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470780&pc=OZD | 06/23/2025 | 07/31/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669023 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support (mcs). three days into the support, a hematoma on the right axillary insertion site was noted. the patient complained of arm pain. heparin was discontinued and the hematoma decreased. one to two days later, heparin was subsequently resumed. the patient continued to have pain that was attributed to the hematoma. the team escalated medications to decrease the pain. day 11 of support, the patient was noted to have a fever, and blood was drawn for labs with work up for sepsis. antibiotics were given. no further information regarding the sepsis work up was noted. following, the patient was reported to be in stable condition and remains on impella 5.5 mcs. | [5 empty] | The investigation of access site bleeding and infection/sepsis has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was most likely patient movement since patient was transferred from outside hospital with hematoma noted on right axillary insertion site. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. d6b explant date added. | Plan of care update noted sepsis work up. follow-up for additional information including results has been requested, is still pending and therefore, unable to rule out. the device remains implanted, supporting the patient and therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7608 | 1220648-2025-30215 | 22470791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470791&pc=OZD | 06/21/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/21/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684889 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, there was clot around the impella. argatroban was added to the purge fluid. support continued. additionally, the bottom end of the anticontamination sleeve was ripped off the locking tuohy. the catheter was covered with tegaderm. support continued. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7609 | 1220648-2025-30216 | 22470836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22470836&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2026657197 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that they encountered high purge pressure and that there was suspicion of heparin-induced thrombocytopenia (hit). the team found a problem with high purge pressure, however, there were no kinks in the system and the purge system was changed without any improvement to the issue. then, tissue plasminogen activator lysis was discussed and instructions were forwarded. a recommendation was also made to change the pump and observe the motor current. the purge system was successfully changed to tpa lyse. however, the motor current began to rise continuously and as a result the decision was made to exchange the pump. additionally, on the day of the explant, the patient received argatrobran instead of heparin due to suspicion of hit. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
7610 | 1220648-2025-30217 | 22471009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22471009&pc=OZD | 06/22/2025 | 12/11/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1847282 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that while a patient was on impella 5.5 support awaiting heart transplant the aic was not streaming. additionally, a nurse reported that they have had multiple battery temperature high alarms when the aic was unplugged from the wall. attempted to replicate without success while on the phone. the aic was exchanged. | [5 empty] | The investigation of aic automated impella controller (aic) - battery charging/other issues has been completed. data analysis: console logs from the complaint date revealed there was only one ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) which triggered when the console was unplugged from ac power that self-resolved in 10 seconds. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc, signifying the battery temperature was not actually high. device analysis: the console¿s batteries and power management system were tested, and no issues were observed. root cause: the battery level low alarm was due to a momentary communication glitch between the batteries and power battery manager (pbm), where a single sample from battery data (in this case value of battery temperature) was corrupt. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-30217. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. a.5 unknown. | Type of report was changed from product malfunction to serious injury.
7611 | 1220648-2025-30218 | 22471011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22471011&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682766 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant presented with worsening shortness of breath and was transferred for further cardiac evaluation. the patient decompensated and the decision for ecpella was made. it was noted that the left groin access was obtained using best practice. an impella cp was inserted and started successfully. extracorporeal membrane oxygenation was cannulated from the right groin. a retrograde sheath was inserted for perfusion. the peel away sheath was removed and the repo sheath was inserted. the left groin started bleeding from both the impella site and the retrograde sheath site. a hematoma was noted. hemostasis was unable to be obtained with manual pressure. the decision was made to remove the impella and the peel away sheath was reinserted. bleeding continued and the decision to bring the patient to the operating room for groin cutdown and repair was made. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7612 | 1220648-2025-30219 | 22471347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22471347&pc=OZD | 06/22/2025 | 09/09/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660167 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported patient had a cp pump placed to support the patient admitted in with acute myocardial infarction complicated by cardiogenic shock. the pump was placed but, there were issues with suction and limb ischemia. suction was remedied by repositioning the pump and giving volume. the ischemic leg was treated by re-perfusion from the extracorporeal membrane oxygenation. after 66 hours of support the team withdrew care and patient expired. | [5 empty] | The investigation of limb ischemia and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data log not available to review. the cause of low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. limb ischemia - reversible: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30219. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7613 | 1220648-2025-30220 | 22471983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22471983&pc=PYX | 06/22/2025 | 10/27/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026619133 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female | 78 KG |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Pericardial Effusion | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered a thrombus. the team had a high suspicion an impella rp clot ingestion. the patient had a large output from the pericardial drain and heparin was paused for three hours. the pump had a sudden increase in pressures. blood pressure dropped and saturated oxygen levels decreased. the decision was made to remove the pump and exchange it with a new one. the patient was noted to be in stable condition. | [5 empty] | Clinical details report a sudden increase in pulmonary artery pressures via both the pump placement signal and swan-ganz catheter on (b)(6) 2025. it was also reported that purge flow was low (6.3 ml/hr). the patient had large bloody output from pericardial drains the previous day, so heparin was put on pause for ~3 hours. there was strong suspicion of a clot, so the pump was explanted and replaced with a new rp flex. thrombus: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. vascular damage - great vessel: clinical details report large bloody output from pericardial drain on (b)(6) 2025. heparin was paused for 3 hours to troubleshoot, but it was not reported if the bleed resolved. it was not reported either where the bleed originated from. no abnormalities found on the returned product. since insufficient clinical details were provided, the root cause of the vascular damage could not be determined. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the thrombus, high purge pressure, and pericardial bleeding has been completed. without sufficient clinical information, the root cause of the pericardial bleeding and thrombus could not be determined. however data logs were reviewed finding the sudden step up in pa pressures was confirmed. at the same instant, there was a spike in motor current to 473 ma with a dampening of motor speed. additionally, the dp ps stepped up similarly to the pa signal. purge pressure remained stable for ~6 more hours before it spiked up to about 750 mmhg in 1 minute. this spike in purge pressure aligned with another spike in the motor current and a ms deviation. purge pressure dropped back down into the 600's shortly after. the rise was not high enough to trigger high purge pressure alarms. product was not returned for investigation. based on the clinical details provided that heparin was paused due to a bleed and the signatures of biomaterial ingestion at the time of the purge pressure spike in data logs. the root cause of the high purge pressure was most likely biomaterial ingestion. additional information: h6: clinical code 3271 added impact code added 4641. the ifu states: impella rp with smartassist system. section: using the automated impella controller with the impella rp with smartassist, system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states), ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
7614 | 1220648-2025-30221 | 22472059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472059&pc=OZD | 06/22/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656051 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 74 year old patient with an impella cp due to acute myocardial infarction and cardiac procedure. the impella was inserted and started, using best practices. during support it was noted that the impella was left shallow across av. flow was optimal for p-level. it was also noted that the patient had sedation turned off and became restless. femstop on groin became mispositioned and hematoma formed. femstop was repositioned in addition to received one unit of red blood cells. additionally, the impella was repositioned under echo guidance due to the pump being caught in papillary muscle. it was also noted that the impella appeared to have been explanted. | [5 empty] | The investigation of positioning issues and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of bleeding issue most likely was due to patient movement since it occurred after patient transferred to icu and malpositioned femstop repositioned and hematoma stable. positioning issues: the cause of positioning issue most likely was use related due to improper technique. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7615 | 1220648-2025-30222 | 22472182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472182&pc=OZD | 06/22/2025 | 08/12/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1687316 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 100 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient post supraventricular tachycardia and atrial fibrillation. the pump was placed and was supporting when the pump and the patient were transferred from community to the tertiary center. the automated impella controller would not turn off and had to be unplugged. the battery was turned off for a hard reset to allow it to turn off. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. the impella device was returned for evaluation. data review: console logs consistent to what was reported in the complaint. device analysis: console was tested after arriving in field service. tested the console by power cycling 10 times but was unable to reproduce the issue. conclusion: the console did not shut down properly, leading to an unexpected controller shutdown alarm upon the next boot. this alarm indicated a forced shutdown during the previous boot, which supports what the clinical staff reported. the cause of the inability to shutdown issue could not be determined due to the inability to reproduce. the console operated as intended while testing on an engineering lab bench. d6a and d6b should have been left blank on manufacturer device report 1220648-2025-30222.
7616 | 1220648-2025-30223 | 22472204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472204&pc=OZD | 06/16/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/16/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606205 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 100 KG |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after explanting an impella cp the patient could not obtain hemostasis. the patient was taken to the operating room for surgical closure. however, after explant they were unable to obtain hemostasis, so the patient was sent to the operating room for surgical intervention. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7617 | 1220648-2025-30224 | 22472360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472360&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649017 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient needing mechanical circulatory support. the decision was made to place impella cp for support. during support the patient was experiencing significant bleeding from the groin. manual pressure was held in an attempt to slow the bleeding. however that was unsuccessful, physician decided to remove impella and placed 2 perclose utilizing double barrel technique and decided to reinsert impella. the impella was flushed, inspected, and reinserted into the left ventricle without issue. peel away sheath left in place given the bleeding patient had been experiencing. there is some concern for distal ischemia. foot appeared to be slightly pale in comparison to left side. however, faint doppler pulse found. it was noted that right lower extremity (rle) had slight color difference but both were cold and mottled. external fem fem bypass placed with improved perfusion to rle. impella has been weaned with continued improved hemodynamics. there were concern for sepsis. pan cultures were sent and broad spectrum antibiotics started. also noted was that cutdown was performed, vessel loops placed, impella weaned and explanted without issue, vessel loops tightened, thrombectomy and right femoral artery repair performed with bovine patch, wound surgically closed. manual pressure held, 50 mg of protamine given. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7618 | 1220648-2025-30225 | 22472375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472375&pc=OZD | 06/15/2025 | 08/11/2025 |  | 07/10/2025 |  |  | 06/15/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608658 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with chest pain and st-segment elevation myocardial infarction was escalated from an impella cp to an impella 5.5 for mechanical circulatory support following a percutaneous transluminal coronary angioplasty of the left anterior descending artery. the clinical staff reported high purge pressure alarms on the impella 5.5. they denied any kinks and had already attempted a purge cassette/tubing change. the doctor was to contact medical affairs for tissue plasminogen activator protocol. the abiomed representative walked the registered nurse through a bag change and disabling audio. a patient update later noted that hemolysis was no longer present. the resolution for the hemolysis is unknown. the patient was successfully weaned from impella 5.5 support. | [5 empty] | The investigation has been completed. the impella device was returned for evaluation. the returned product had no physical abnormalities. biomaterial was observed in the purge gap and wrapped around impeller. the data logs revealed a motor current shift after which there was elevated purge pressure and 'purge pressure - high' and 'purge system blocked' alarms. there was also a sudden drop in purge flow. the pump was run in the lab and high purge pressure did not reproduce after biomaterial cleared. the root cause of the high purge pressure was determined to be biomaterial as evidenced by the observed biomaterial blocking the purge gap, the motor current shift seen in the data logs, and returned product not reproducing issue after biomaterial clearing. regulatory evaluation mentioned 'hemolysis no longer noted' with limited clinical information. the returned product had no physical abnormalities. biomaterial was observed in the purge gap and wrapped around impeller. the data logs reveal a motor current shift with speed deviation after which there was elevated purge pressure and 'purge pressure - high' and 'purge system blocked' alarms. no information on hemolysis onset or resolution was provided. the root cause of the hemolysis was unable to be determined as no information regarding hemolysis onset or resolution was provided thrombus failure mode was added because of biomaterial on returned product. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7619 | 1220648-2025-30226 | 22472606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472606&pc=OZD | 06/22/2025 | 07/10/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620605 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was found to have microvascular decompression. plan for high-risk percutaneous coronary intervention with coronary artery bypass grafting as backup. decision made to use impella cp to support case. it was noted that during support the patients labs were hemolyzing. impella was repositioned under echocardiogram and level decreased. patient awaiting explant in cath lab however was noted that traumatic foley insertion lead to bleeding and now continuous bladder irrigation. systemic anticoagulation was being held. the impella was successfully weaned and explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7620 | 1220648-2025-30227 | 22472634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472634&pc=OZD | 06/21/2025 | 09/24/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665839 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 58 KG | Hispanic |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support. it was reported that the separation of the dilator left the clear threaded locking plastic in place in the orange peel away sheath. the cp was able to be inserted without issue. removal of the peel away was done, but the clear locking plastic was unable to be removed. the only option to remove would have been to remove the impella completely. that was not an option, so the decision was made to advance the repositioning sheath into the clear locking plastic. the repositioning sheath was able to be properly positioned without causing bleeding or discomfort to the patient. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of introducer damage ¿ mechanical has been completed. impella cp and 14fr x 25cm dilator were returned for evaluation. upon visual inspection, the clear locking plastic cap was missing from the dilator and found around the catheter of the impella cp. no cracks or signs of excessive force were observed on the dilator or clear plastic ring. clinical details noted that when removing the dilator from the peel-away introducer, the clear plastic locking ring remained on the introducer and became separated from the dilator. the clear ring remained around the impella catheter after insertion and the repositioning sheath was inserted through the catheter without issue. the root cause of the introducer damage could not be definitively determined as no failures were found on the returned product and the issue was unable to be reproduced. d.9. device returned to manufacturer date added h.6. component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30227.
7621 | 1220648-2025-30228 | 22472914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472914&pc=OZD | 06/22/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677843 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and after the implant, the patient had a brief run of ventricular tachycardia. the patient was shocked and amio was started. additionally, the patient had a confirmed subarachnoid bleed. the bleed was increasing and the impella was later explanted due to the bleeding. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of vf/vt/af/at and cva/stroke has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. health effect - clinical code 1770 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30228. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7622 | 1220648-2025-30229 | 22472970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22472970&pc=OZD | 06/22/2025 | 08/20/2025 |  | 07/10/2025 |  |  | 06/22/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684307 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female | 64 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there were low pump flows with runs of ventricular tachycardia. amiodarone bolus and drip was given as well as one unit of packed red blood cells. support continued. additionally, the patient developed bloody bowel movements with small bowel ischemia. bivalrudin was discontinued and blood products were given. furthermore, a computed tomography was performed revealing right parietal hemorrhage, extending to posterior right temporal lobe with adjacent small areas of acute infract with hemorrhagic conversion. anticoagulation was withheld and support continued. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of vf/vt/af/at, low pump flow, vascular damage - peripheral vessel, and stroke/cva has been completed. no damages or thrombus was noted with the returned product. vt/vf: the patient had runs of vt. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessels: it was stated that the patient had small bowel ischemia. due to a lack of clinical details, the root cause of the vascular damage - peripheral vessels cannot be determined. low pump flow: data logs show slight variations in placement signal and motor current with responding flows. suction alarms are seen throughout and flows are on within specification but on the lower end. the returned product did not have any thrombus or kinks in the cannula. low pump flow was unable to be reproduced. based on the clinical details and returned product analysis, the root cause of the low pump flow is most likely due to the patients condition. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7623 | 1220648-2025-30230 | 22473117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22473117&pc=OZD | 06/23/2025 | 08/21/2025 |  | 07/10/2025 |  |  | 06/23/2025 | 07/10/2025 |  |  |  |  | 07/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671973 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was being escalated from cp to impella 5.5. it was noted that during 5.5 support the pump stopped. attempts made to restart pump with no success. alarm stated pump defective. the 5.5 was explanted and a new cp was placed. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. the product was not returned for investigation. the data log shows a motor current spike to 1200ma 62 hours into support followed by one 6 minutes later, with slight interrupted pump flow and speed deviations. approximately 5 minutes later the pump stops and triggers #119 alarm. attempts to restart the pump were unsuccessful. without product returned, the root cause of the pump stop cannot be determined.
7624 | 1220648-2025-30148 | 22456324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456324&pc=OZD | 06/23/2025 | 12/04/2025 |  | 07/09/2025 |  |  | 06/23/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210718 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during training, the automated impella controller (aic) triggered an alarm, and the aic was unable to shut down. the alarm message displayed defective device. however, the impella pump was still connected. a hard shut down of the aic was performed. there was no patient involvement at the time of the event. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: the console logs from the reported day of the event align with the complaint reported. it was noted that the issue associated with the complaint began on (b)(6) 2025. during this period, the pump was connected and within few seconds an error message was recorded: "eeprom problems while reading calibration data." this triggered the "impella defective (error reading eeprom)" alarm (#3). following this, the pump was disconnected and reconnected multiple times, but the same alarm was repeatedly triggered. ultimately, the aic was forcefully shut down because it was unable to detect that the pump had been disconnected. additionally, the console was powered on multiple times on (b)(6) 2025, yet the alarm persisted and never resolved. device analysis: the console (b)(4) was returned to field service for further investigation. during testing, the console displayed an "impella defective" alarm and the reported error was successfully reproduced. the console was unable to detect the pump and could not shut down properly, necessitating a forced shutdown. further analysis revealed that the impellatronic pca was defective, causing communication issues within the console. after replacing the faulty impellatronic pca with a known good unit, the console was able to detect the pump correctly. the aic was then tested using a known good pump and purge cassette, and no issues were identified. additionally, no power supply interruptions were detected between the pbm and the replaced impellatronic pca. the batteries were also inspected and found to be in good condition, with no issues observed. root cause: the root cause for the impella defective (error reading eeprom) - alarm # 3 was due to an impellatronic pca malfunction. no issues were found with the reported fm of aic - shutdown or restart issue.
7625 | 1220648-2025-30149 | 22456385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456385&pc=OZD | 06/23/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/23/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674900 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 73 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced limb ischemia. prior to the beginning of the percutaneous coronary intervention (pci), the patient became hypotensive and bradycardic. medication was administered and the decision was made to implant an impella cp which was successfully completed. pci was completed to the proximal left anterior descending artery with plans to address the right coronary artery once the patient was more stable. the following day, the patient started complaining of leg pain. pulses were not great since the impella placement and an arterial doppler the day prior showed flow. however, pulses were not currently present (action taken to address the limb ischemia is unnoted). the next day, the patient was taken to catheterization lab, and the physician attempted to intervene the right coronary artery. however, this pci was unsuccessful, and the physician opted for medical management. the patient was successfully weaned from the impella cp and was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7626 | 1220648-2025-30150 | 22456504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456504&pc=OZD | 06/06/2025 | 09/24/2025 |  | 07/09/2025 |  |  | 06/17/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568239 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Death | Yes | Yes | Difficult to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Perforation | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On july 11th, 2025 new information was received stating, that due to increased friction between the impella-catheter and the companion sheath, an uncontrolled movement towards the ventricle occurred, causing the pigtail to puncture the ventricle. the patient expired. | The complainant reported a patient was on impella cp support and during the implant, the ventricle was perforated with the pigtail. the clinical representative was told that the patient was alive and doing well. no further details were provided. |  |  |  |  | The investigation of ventricle perforation and difficulty inserting/removing pump through introducer has been completed. the cause of the ventricle perforation issue was most likely use-related, as uncontrolled movement of the catheter punctured the ventricle. the cause of the difficulty inserting/removing pump through introducer issue was not determined since product was not returned. b1 revised to reflect both adverse event and product problem. b2 death and date of death added. b5 patient outcome added. h1 type of report revised to death. h6 health effect - impact code 1802 added and medical device problem code 2920 added based on addtional information received from the clinical site. h10 instructions for use for difficulty inserting/removing pump through introducer added. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during impella section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves . ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7627 | 1220648-2025-30151 | 22456643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456643&pc=OZD | 06/17/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/17/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660211 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant have an impella 5.5 placed via the left axillary to replace a prior impella 5.5 placed via the right axillary that was explanted with concerns of infection. the new impella 5.5 was supporting while the patient was awaiting transplant. the pump lost placement signal after four days of support but remained on after troubleshooting and disabling the alarms. no patient harm was noted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7628 | 1220648-2025-30152 | 22456715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456715&pc=OZD | 06/17/2025 | 08/27/2025 |  | 07/09/2025 |  |  | 06/17/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673670 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female | 64 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and experienced access site bleeding. bleeding was also noted at the old/explanted extracorporeal membrane oxygenation access site. five bags od blood were administered. partial thromboplastin time (ptt) was noted to be very high at 200 seconds. heparin was paused for two hours and ptt was rechecked and was 120 seconds. heparin was paused again for additional two hours. the patient became septic. additional information was received three days later. the patient was hemodynamically stable and the impella cp was successfully weaned due to improving left ventricle function. the groin side bleeding was resolved before impella weaning with optimizing of the sheath-angle and optimizing of activated clotting time. | [5 empty] | The investigation of access site bleeding and infection/sepsis has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was most likely anticoagulation management since partial thromboplastin time (ptt) was very high and groin side bleeding was resolved before impella weaning with optimizing of the sheath-angle and optimizing of activated clotting time (act). infection/sepsis: the root cause of infection/sepsis could not be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7629 | 1220648-2025-30153 | 22456762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456762&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/09/2025 |  |  | 06/17/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593273 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a st-segment elevation myocardial infarction patient was having a planned impella cp insertion. they began prepping and pump set opened. angiogram revealed artery too small for placement of impella. alternate access sites assessed and still too small of a caliber to place impella. the patient's anatomy will not accommodate placement of impella. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since the artery was too small for placement of impella. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. a1-a5, d6a/d6b, e1 initial reporter name and contact are unknown.
7630 | 1220648-2025-30154 | 22456863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22456863&pc=OZD | 06/18/2025 | 08/15/2025 |  | 07/09/2025 |  |  | 06/19/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026677239 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Unspecified Infection | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on impella 5.5 had placement monitoring disabled on console due to consistent impella in aorta alarms. impella placement confirmed in correct position with echo. in addition the providers questioning hiit. purge solution transitioned to d5w w/bicarb & systemic anticoagulation transitioned to bival. the patients temperature increased over the past couple of days. antibiotics & cooling administered. it was noted that the patient continued to improve clinically. the impella was successfully weaned and explanted. | [5 empty] | The investigation of optical signal issue, thrombocytopenia, and infection/sepsis has been completed. the impella device was not returned for evaluation. optical signal issue: the data log shows several "pump position in aorta" alarms on (b)(6) 2025 without any observed abnormalities on the 5s optical signal parameters. additionally, the rt log was closely reviewed regarding motor current and placement signal, with no abnormalities observed. according to the clinical notes, the patient was on ecpella support, and the position was verified during the false position in aorta alarm. there was insufficient clinical information to support if there was any patient condition information relative to false position alarm. the cause of the optical signal issue was not determined since product was not returned and data log was not sufficient to derive root cause. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. h6 health effect - clinical code 1942, health effect - impact code 4627, and medical device problem code 2993 were erroneously entered on the initial manufacturer device report number 1220648-2025-30154. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: infection/sepsis: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ thrombocytopenia: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7631 | 1220648-2025-30155 | 22457031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457031&pc=OZD | 06/19/2025 | 08/15/2025 |  | 07/09/2025 |  |  | 06/19/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679760 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 65 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was unable to be implanted due to the patient¿s anatomy. additionally, the impella catheter kinked during the attempt of insertion. there was no patient harm reported. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of the delivery issue was most likely patient condition related, impella 5.5 was unable to be implanted due to the patient¿s anatomy. product damage( cannula kink) : the root cause of the product damage issue was most likely patient condition related, the impella catheter kinked during the attempt of insertion due to patient's anatomy.
7632 | 1220648-2025-30156 | 22457089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457089&pc=OZD | 06/17/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/17/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681797 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support and upon removal of the 14 french peel-away sheath, the femoral artery became dissected. the patient was taken to surgery. the pump was explanted and the patient was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7633 | 1220648-2025-30157 | 22457120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457120&pc=OZD | 06/19/2025 | 12/09/2025 |  | 07/09/2025 |  |  | 06/19/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1111148 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during inventory, when switching on the automated impella controller (aic), alarms ¿console error¿ & ¿battery failure¿ were triggered. there was no patient involvement and the product was returned for investigation. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: console logs from the reported day of event are consistent with problem description and show multiple and persistent controller error alarms (#1023, pbm1 communication loss), and occasionally battery failure (#360) and battery level low (#23) alarms. power-cycling of the console did not resolve pbm1 communication issues, which were identified in device logs as early as 2025-05-02, while pump 532908 was used, and evidenced by sau_powermod_1 communication error messages (without alarms). there was one instance of controller failure alarm (#1027, pbm sw corruption), which presumably was falsely triggered due to communication errors. device analysis: the console was returned for investigation. all communication performed as expected and both batteries were charging when plugged into ac power, not reproducing the failure mode from the field. root cause: it was found that the issues were caused by incompletely engaged battery transport switch.
7634 | 1220648-2025-30158 | 22457131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457131&pc=OZD | 06/19/2025 | 08/16/2025 |  | 07/09/2025 |  |  | 06/19/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677841 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was placed and in admission to icu plasma free was elevated with lab drawn. by midnight, urine turned red. plasma free and lactate dehydrogenase continued to rise. the physician believes that the patient was having hemolysis secondary to clot entrapment. the patient was taken back to the operating room the next morning for impella removal. the patient was noted to be in stable condition. | [5 empty] | The investigation of hemolysis and thrombus has been completed. hemolysis: the returned product had no abnormalities. there was no biomaterial observed by the impeller or any damage to impeller blades. the data logs do not show any motor current spikes or continuous suction. clinical reported position and preload were both optimal. the pump passed hemolysis testing. the root cause of the hemolysis was not determined as no hemolysis related abnormalities were seen in the data logs and on the returned product and limited clinical information was provided. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7635 | 1220648-2025-30159 | 22457255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457255&pc=OZD | 06/15/2025 | 08/13/2025 |  | 07/09/2025 |  |  | 06/15/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647332 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction and cardiogenic shock had been implanted with an impella cp for mechanical circulatory support. a bedside echocardiogram revealed severe mitral regurgitation with suspected valve vegetation. the impella cp had been placed into the left ventricle. however, flow initially reached approximately 2.0 liters per minute before dropping to 0.5 liters per minute and could not be increased. this was likely due to some valve debris coming off the mitral valve, which showed vegetations on echocardiogram and was the source of the severe mitral regurgitation. the patient had begun to rapidly decompensate prompting the decision to explant the device. a new impella cp was inserted and the patient survived the procedure and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data log review showed a motor current spike about 5 minutes after pump startup with an immediate drop in pump flow. there was a suction alarm triggered shortly after the drop in pump flow. there was also a slight shift down in placement signal (ps). about a minute after the motor current (mc) spike, there was a loss in pulsatility in pump flow. no product was returned. the root cause of the low pump flow was most likely biomaterial ingestion since there was a mc spike immediately followed by a drop in pump flow seen in data log review. thrombus: the failure mode (fm) thrombus was added since the low pump flow was most likely due to biomaterial ingestion based on a motor current spike seen in the data logs. no product was returned. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7636 | 1220648-2025-30160 | 22457273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457273&pc=OZD | 06/19/2025 | 08/30/2025 |  | 07/09/2025 |  |  | 06/19/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1452168 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller cord reported to be fried by the biomed department of the hospital. the power cord will be returned for investigated and there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - damage before therapy has been completed. data analysis: console logs are not relevant to this complaint. note: the aic was not returned for investigation; only the defective ac cord was sent back. device analysis: the console was not returned to field service; however, the reported ac power cord was sent back for further examination. during visual inspection, it was observed that the live (black) wire within the plug was disconnected. it appears that the wire covering or outer sheath was improperly cut, as per the vendor wiring instructions. also, the clamps were not tight enough, and light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose and detaching from the plug and causing a spark. root cause: the reported cause of the reported issue was due to the loose wire inside the ac cord plug. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7637 | 1220648-2025-30161 | 22457375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457375&pc=OZD | 06/19/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/19/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660114 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient developed a gastrointestinal bleed (gib) with bloody stools. the doctor was unsure of the cause of the gib. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7638 | 1220648-2025-30162 | 22457496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457496&pc=OZD | 06/14/2025 | 08/11/2025 |  | 07/09/2025 |  |  | 06/14/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026685040 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 92 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was initially presented to an outside emergency department with complaints of chest pain. a computed tomography (ct) scan revealed a type a aortic dissection with a dilated aortic root measuring 6.5 centimeters. due to left ventricular distension while on venoarterial extracorporeal membrane oxygenation (va ecmo), the decision was made to place a percutaneous impella cp for left ventricular unloading. during support, the patient developed bleeding around the repositioning sheath. a fem-stop was applied for hemostasis, and one unit of packed red blood cells was transfused. the impella cp also exhibited intermittent suction events which were managed with boluses of 25% albumin and by adjusting the p-level settings to relieve suction. due to persistent hemodynamic instability, the patient was escalated to an impella 5.5. the patient was on impella 5.5 support for 12 days and care was ultimately withdrawn, and the patient expired. | [5 empty] | Medical safety officer review: there is not enough evidence to exclude the impella cp device used as an associated factor in the patient's outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of reported bleeding issue most likely patient condition related since bleeding occurred in multiple sites. low pump flow: the cause of low pump flow issue most likely was patient condition related because impella was used as a vent for ecmo in addition to the bolus of 25% albumin, and adjusting p-level to break suction.
7639 | 1220648-2025-30163 | 22457633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457633&pc=OZD | 06/20/2025 | 08/12/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1654451 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered issues with the automated impella controller (aic). as the team was placing an impella cp and as the nurse went to go insert the purge cassette into the aic, it would not click into place all the way. the team had to switch aic's and it then inserted just fine with successful pump placement. the patient was transferred and the team discussed treatment options. ultimately there was nothing else that could be done for the patient and the family decided to withdraw care. the patient expired and there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the impella device was returned for evaluation. data review: console logs were partially consistent with what was reported in the complaint. according to the imc logs: the user inserted the purge cassette during step 2 of the purge setup wizard. at this time the purge disc was not inserted yet and the purge disc help screen was displayed. the user then removed the purge cassette and inserted the purge disc shortly after. numerous consecutive errors were observed after the purge disc was inserted which is indicative of the user likely having issues re-inserting the cassette into the purge assembly device analysis: console was tested after arriving in field service. the reported purge cassette insertion issue was unable to be reproduced. conclusion: the root cause of the reported purge cassette insertion issue could not be determined due to the inability to reproduce. the console operated as intended while testing on an engineering lab bench with a known good purge cassette kit. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-30163. h8 was erroneously selected on the initial manufacturer device report number 1220648-2025-30163. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7640 | 1220648-2025-30164 | 22457756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457756&pc=OZD | 06/20/2025 | 08/16/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644945 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal on an impella cp pump became unreliable during patient support. the placement monitoring was disabled in response to the noted issue and support was continued uninterrupted. two days later, the patient was made palliative care by their family and passed away on support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and lt log of case showed sudden increase in placement signal out of range, with optical 5s parameters consistent with oscillating contrast. clinical details noted that a sudden placement signal not reliable (psnr) alarm was triggered and no abnormalities with patient were noted as clinical team was in the room and monitoring the patient when it occurred. upon review of returned data logs, it is apparent that the console is the potential cause of the issue. the root cause of the optical signal issue could not be definitively determined as no product was returned for evaluation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7641 | 1220648-2025-30165 | 22457913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457913&pc=OZD | 06/20/2025 | 08/15/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667581 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was unsuccessful due to severe aortic tortuosity in the patient's anatomy. the patient expired. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to severe aortic tortuosity. d6b explant date was erroneously entered on the initial manufacturer device report. | A2, a3, a4, and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7642 | 1220648-2025-30166 | 22457924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22457924&pc=OZD | 06/20/2025 | 08/16/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660262 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female | 124 KG | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | Material Puncture/Hole | Failure to Sense |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed an impella 5.5 to support a coronary artery bypass graft patient. when the support was ongoing and on day 2, the pump was being repositioned and the team in error unplugged the impella 5.5 to relieve the cord torque and in doing this the pump stopped and could not be restarted. the pump was removed and replaced without any noted harm or injury. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop and positioning issues has been completed. pump stop: product investigation showed no actual pump stop relative failure as device was not able to be recognized on console. data logs were reviewed and no true pump stop event was found. one event of pump stop alarm was triggered at end but was noted with relative to pump being disconnected while running. clinical reported patient returned to the or for repositioning of the impella and during the repositioning they unplugged the impella 5.5 to remove torque and the pump stopped. the root cause of pump stop was use issue as the team in error unplugged the 5.5 to relieve the cord torque and led to pump stop and could not be restarted. positioning issue: clinical reported patient returned to the or for repositioning of the impella and during the repositioning they unplugged the impella 5.5 to remove torque and the pump stopped. unknown why or what lead the pump to be repositioned. the root cause of the positioning issue was not determined due to insufficient clinical details. pump not detected: visual inspection revealed no biomaterial or other foreign material inside pump cannula. the electrical resistance test showed no open lines or abnormalities. pump was connected to console but the pump sn was not recognized by the console with no sn displayed on screen. no noted alarms on screen but pump p-level change keypress button was disabled for selection as pump was not recognized. blood in red handle was observed during visual inspection on catheter side. puncture hole marks close to 30 cm marking was found. no other physical abnormalities were found. the root cause of the pump not detected issue was blood ingress due to a hole in the catheter resulting from improper use. product damage(hole in catheter): the root cause of the product damage(hole in catheter) issue was use related due to puncture hole impression observed on catheter which likely occurred during support. d9 device returned to manufacturer date added g1 manufacturing site address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30166.
7643 | 1220648-2025-30167 | 22458022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22458022&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/09/2025 |  |  | 06/23/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon cardiology evaluation, st elevations were no longer present; however, the patient was deemed hemodynamically unstable, and an impella cp device was placed for mechanical circulatory support. cardiothoracic surgery was consulted due to the presence of multivessel coronary artery disease (mvd). during an attempt by the physician to reposition the pump, the cannula became kinked. as a result of this damage, the impella cp was later explanted and replaced. the patient survived the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (cannula kink) and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: the clinical details state that after the pump was placed and started it was too deep so the pump was pulled back but then was deemed too shallow. attempting to readvance it kinked the cannula near the outlet. based on the clinical details the root cause of the positioning issues is most likely due to use as it was placed in an improper position. product damage (cannula kink): the clinical stated that the pump was too shallow so it was readvanced and then the cannula kinked near the outlet. based on the clinical details provided, the root cause of the product damage (cannula kink) is most likely due to use as re-advancing the pump kinked the cannula.
7644 | 1220648-2025-30168 | 22458132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22458132&pc=OZD | 06/20/2025 | 08/15/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679374 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 60 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 failed. after gaining axillary access, the impella 5.5 was attempted to be inserted through the graft into the artery but the artery was smaller than anticipated and did not have much elasticity. the outlet cage was able to be inserted into the artery but when the outlet cage and motor was attempted, it was met with much resistance to the point that the medical doctor was afraid of dissecting the artery and therefore, the insertion was aborted. vessel was dissected and was repaired utilizing multiple patches. a drain was placed inside the pocket of the axillary cut down and closure was made successfully. the patient survived and remained on venoarterial extracorporeal membrane oxygenation support. | [5 empty] | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. the product was returned and there were no damages seen. based on the clinical details, the root cause of the delivery issues was due to the patients condition as the vessels were smaller than 7mm and the pump was unable to pass through the anastomosis. the root cause of the vascular damage - peripheral vessels is due to the patients condition as they had small and thinned walled vessels causing a vessel dissection when trying to pass the pump through. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7645 | 1220648-2025-30169 | 22458673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22458673&pc=OZD | 05/07/2025 | 10/06/2025 |  | 07/09/2025 |  |  | 05/07/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657694 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support was repositioning the impella without medical assistance, and the impella stopped. the patient was able to restart the impella at p9 with no issues. it was further reported that the impella contamination sleeve was broken. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of temporary pump stop and product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. temporary pump stop: based on the data log review, no pump stops were found outside of the impella stopped alarm at the end of the case run when the pump is turned off which is to be expected. purge related alarms were seen but were not sufficient to conclude purge system issues. no motor current spikes were seen outside of p-level changes. no electrical alarms or impella defective alarms were seen. since the product was not returned for investigation, a lab run could not be performed to try to replicate the reported issue. the root cause of the temporary pump stop was not determined due to insufficient information related to the failure from the data logs and unreturned product. sterile sleeve damage: based on the clinical review, it was found that the sterile sleeve was broken at some point during the case. the root cause of the sterile sleeve damage was unable to be determined due to lack of sufficient clinical details and unreturned product.
7646 | 1220648-2025-30170 | 22458991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22458991&pc=OZD | 05/21/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 05/23/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643307 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a clot near the pump inlet on the pigtail. there was no functional impact on the pump resulting from the thrombus. the physician ultimately decided to remove the pump. | [5 empty] | A4 is unknown. no product was returned for evaluation. clinical details noted that a clot was present near the inlet on the pigtail during a routine echo and the pump was explanted. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7647 | 1220648-2025-30172 | 22459072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22459072&pc=OZD | 05/25/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 05/25/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655563 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a groin site hematoma during impella 5.5 support. the hematoma was noted at the intra-aortic balloon pump site. blood products were given and a washout was performed. two days later, the patient suffered atrial fibrillation. dobutamine was paused in response to the condition. after five more days of support, the anticontamination sleeve got caught and was inadvertently torn. the sleeve was cleaned and wrapped in a chlorhexidine gluconate dressing. as support continued, intermittent impella position unknown alarms began to appear. the alarm was eventually disabled with no impact to the patient. impella support continues. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7648 | 1220648-2025-30171 | 22459074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22459074&pc=OZD | 06/20/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677103 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 65 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and developed limb ischemia that was treated with distal perfusion. the patient was escalated to an impella 5.5 and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7649 | 1220648-2025-30173 | 22459329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22459329&pc=OZD | 06/06/2025 | 08/04/2025 |  | 07/09/2025 |  |  | 06/06/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025605491 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 95 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 after care escalation from an impella cp. the patient had been admitted in with cardiomyopathy and cardiogenic shock. on the day after implant, the purge pressure rose prompting the team to give tissue plasminogen into the purge. on day 10, a large hematoma in the right axillary soft tissues was noted and there was a possible pseudoaneurysm. the team treated by infusing blood products. additionally, there was a subarachnoid bleed. this prompted the team to stop heparin and switch to sodium bicarbonate. the pump remained on despite this until the family made the choice to withdraw care. medical safety officer review: multiple pumps which included bipella support with impella rp and impella 5.5 supported the patient. impella cp and impella rp related events are serious injury type of reportable events. the demise outcome is associated with the impella 5.5 as there is not enough evidence to excluded the impella 5.5 device used as an associated factor to the patient's outcome; last device used and the patient had hemolysis, extensive repair and blood transfusion which is more likely to be related to the demise outcome. | [5 empty] | The investigation of access site bleeding, stroke/cva, high purge pressure, hemolysis, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: no product was returned and data logs are not applicable. the root cause of the access site bleeding was not determined since insufficient clinical details were provided and no product was returned. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. high purge pressure (hpp): the data log review did not show any high purge pressure alarms until (b)(6) 2025. on (b)(6) 2025, there was a brief spike in purge pressure that resolved within 4 minutes. a few hours after this spike, there was a shift upward in purge pressure from 400 mmhg to 600 mmhg. there were some long bag/cassette changes throughout the case, but none directly prior to the hpp event. purge pressure was stable from (b)(6) 2025 until the hpp event on (b)(6) 2025. no problem was found with the 5.5 during the time of the complaint based on data log review. hemolysis: data log review showed suction alarms throughout the case, but most present between (b)(6) 2025. there are no major positioning issues based on the data logs. there is a motor current spike on (b)(6) 2025. the placement signal varies throughout the case with no clear trend. no product was returned. the root cause of the hemolysis was not determined since no product was returned and insufficient clinical details were provided. vf/vt/af/at: the root cause of the vf/vt/af/at was unable to be determined due to insufficient clinical details. h6 health effect - clinical code 1770 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30173. | Medical safety officer review: multiple pumps which included bipella support with impella rp and impella 5.5 supported the patient. impella cp and impella rp related events are serious injury type of reportable events. the demise outcome is associated with the impella 5.5 as there is not enough evidence to excluded the impella 5.5 device used as an associated factor to the patient's outcome; last device used and the patient had hemolysis, extensive repair and blood transfusion which is more likely to be related to the demise outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7650 | 1220648-2025-30174 | 22459629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22459629&pc=OZD | 06/07/2025 | 08/05/2025 |  | 07/09/2025 |  |  | 06/07/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658519 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female | 121 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support (mcs) and experienced sustained ventricular tachycardia requiring intervention and at time of removal device cannula separation. the patient presented to the hospital for a right heart catheterization. right heart pressures were elevated, and cardiac index was low. the shock team was elevated for durable left ventricular assist device (dlvad) work up and impella 5.5 device placement. the impella 5.5 device was successfully implanted, guidewire was removed and the impella 5.5 pump was started. the sheath was removed, and the graft was trimmed to size. the hub was advanced and skin was closed with butterfly suturing. a transesophageal echocardiogram for placement measured 6 centimeters and was pulled back to 5.6 centimeters. pumping was noted to be running well with no issues. upon transitioning to the bed, the patient experienced sustained runs of ventricular tachycardia. under echocardiogram for the position, the impella 5.5 device was pulled back to 5.4 centimeters. amiodarone bolus was given and potassium was replaced. the ectopy resolved, and the patient was sent to the unit on impella mcs. on day sixteen, the patient was being bridged to the dlvad and during explant of the impella 5.5 device, while explaining through left ventricular apical core, the cannula detached from the outflow cage. however, the patient was successfully bridged to the dlvad with prophylactic impella rp flex placement for right ventricular support. the patient was noted to be hemodynamically stable following the bridge. | [5 empty] | Additional information provided in h3, h6 and h10. the investigation into the reported issues has been completed. with regard to the reported product damage, there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. this issue has previously been seen during removal. the design testing has stated that the bond is tested to a 30 n specification. additionally, the bond was ripped at the cannula to ph bond and was jagged remaining on the cannula. part of it remained on the motor housing. on the inside of the cannula there was also residual glue marks. there were no kinks seen in the cannula or the catheter. due to a lack of sufficient evidence during the removal of the 5.5 and evidence of sufficient glue at the bond, the root cause of the product damage (detached inflow/outflow) cannot be determined. the root cause of the reported ventricular tachycardia was unable to be determined due to insufficient clinical details. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7651 | 1220648-2025-30175 | 22459859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22459859&pc=OZD | 06/20/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660614 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced two episodes of ventricular tachycardia and was successfully cardioverted. afterwards, the pump had to be slightly repositioned due to aspiration alarms. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7652 | 1220648-2025-30176 | 22460351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22460351&pc=OZD | 06/20/2025 | 07/09/2025 |  | 07/09/2025 |  |  | 06/20/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659439 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 80 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced supraventricular tachycardia during insertion. the patient was cardioverted and then experienced atrial fibrillation following impella position in the left ventricle. the patient was later successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7653 | 1220648-2025-30177 | 22460660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22460660&pc=OZD | 06/21/2025 | 08/19/2025 |  | 07/09/2025 |  |  | 06/21/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658772 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 was unable to be advanced beyond the graft anastomosis during attempted delivery. after multiple failed attempts, the decision was made to abort the implant and implant an impella cp via axillary graft. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to resistance met to at cannula once past the graft anastomosis in the vessel.
7654 | 1220648-2025-30178 | 22460759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22460759&pc=OZD | 06/21/2025 | 08/19/2025 |  | 07/09/2025 |  |  | 06/21/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026689418 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, there were suction alarms. the patient received albumin and blood products. suction alarms continued. additionally, the patient had ventricular arrhythmias on support, the pump was found to be very close to mitral valve leaflets. the staff attempted to reposition the pump and the pigtail was pulled across valve. the staff was unable to reposition. the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. low pump flow: on day 2 of support, suction reported. the log was reviewed and the suction was confirmed. no placement signal trends or motor current spikes observed. a long bag change (3 min) was observed 6 hours into initial support with a brief purge system open alarm. no purge related issues observed thereafter. the cause of the suction was not determined since product was not returned and insufficient clinical details were provided. vf/vt: reported intermittent junctional rhythm changes. the root cause of the intermittent junctional rhythm changes was unable to be determined due to insufficient clinical details. positioning issues: repositioning was attempted due to suction alarms and the pump pulled out of the ventricle during repositioning. the cause of the pump pulling out of the aorta during repositioning was use related.
7655 | 1220648-2025-30179 | 22460972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22460972&pc=OZD | 06/23/2025 | 08/21/2025 |  | 07/09/2025 |  |  | 06/23/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678172 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 59 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Cardiac Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella 5.5, the patient's left ventricle was perforated. implantation of the pump was delayed until the left ventricle was surgically repaired. after the pump was implanted, the patient experienced ventricular tachycardia and was treated with lidocaine. additionally, the sternotomy incision began oozing so systemic heparin was withheld. the patient was successfully weaned and the pump was explanted. | [5 empty] | The investigation for the ventricle perforation, ventricular tachycardia, and sternotomy bleed has been completed. without sufficient clinical details and the device evaluation, the root cause of the vt and the sternotomy bleed could not be determined. however, clinical details provided stated that when trying to further advance the impella, the ventricle was perforation. the physician mentioned that the patient's tissues were frail and in the future would use 0.018 wire for insertion. based on the clinical details, the root cause of the ventricle perforation is most likely due to the patients condition as the patient had frail tissues making them more susceptible to experiencing a perforation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7656 | 1220648-2025-30180 | 22461202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22461202&pc=OZD | 06/23/2025 | 08/20/2025 |  | 07/09/2025 |  |  | 06/23/2025 | 07/09/2025 |  |  |  |  | 07/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678212 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that explantation of an impella 5.5 was complicated as the doctor ran into some resistance while withdrawing the cannula out of the graft. the cannula was inadvertently detached from the outlet cage and motor housing, leaving the cannula in the graft. fortunately, the cannula was still protruding from the graft, and was able to be withdrawn completely without further incident. the pump was successfully explanted and no patient harm was reported. | [5 empty] | The investigation of product damage (detached inflow/outflow ¿ great vessel) has been completed. there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. the design testing has stated that the bond is tested to a 30 n specification per dir550-1.1.7.2. additionally, the bond was ripped at the cannula to ph bond and was jagged remaining on the cannula and the motor housing. on the inside of the cannula there was also significant residual glue marks. there were no kinks seen in the cannula or the catheter. clinical details note there was resistance and the pump got caught on the sternal notch. due to a lack of sufficient evidence during the removal of the 5.5 and evidence of the glue bond, the root cause of the product damage (detached inflow/outflow) cannot be determined. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7657 | 1220648-2025-30103 | 22435261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435261&pc=OZD | 07/01/2025 | 09/22/2025 |  | 07/08/2025 |  |  | 07/02/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661753 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 was surgically removed during the weaning removal, but resistance was encountered during removal. the impella 5.5 was able to be removed, but from the right subclavian artery to the brachiocephalic artery and became dissected. a covered stent was placed. the doctor commented that the dissection was due to the subclavian vascular system diameter which was about six millimeters. the patient survived and "the patient's" outcome was serious. | [5 empty] | The investigation of vascular damage - peripheral vessel and removal issues has been completed. vascular damage - peripheral vessel: based on the clinical details, the root cause of the vascular damage - peripheral vessel is most likely due to the patients condition as it was stated that during removal, resistance was encountered and the artery became dissected as it was a smaller diameter. product damage (detached inflow/outflow - great vessel): there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. it was also noted that there was resistance with explanting causing a vessel dissection as the diameter of the vessel was more narrow. the delo band remained jagged on the cannula with part of it attached to the motor housing indicating it was ripped off. residual glue marks are noted on the inside of the cannula and on the motor housing. there were no kinks seen in the cannula or the catheter. due to a lack of sufficient evidence during the removal of the 5.5 and evidence of the glue bond, the root cause of the product damage (detached inflow/outflow) cannot be determined. d9 device returned to manufacturer date added. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7658 | 1220648-2025-30104 | 22435313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435313&pc=OZD | 06/19/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/19/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026681842 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 59 KG | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Unspecified Infection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella 5.5 access site became infected during patient support. no noted intervention was taken to manage the condition and the catheter was noted to be secured. as support continued, intermittent suction alarms began to appear for the next two days. positioning of the device was not verified and the providers elected to not perform an echocardiogram at that time. support remained ongoing despite the noted issue. the pump was successfully weaned and explanted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7659 | 1220648-2025-30105 | 22435507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435507&pc=OZD | 06/19/2025 | 08/21/2025 |  | 07/08/2025 |  |  | 06/19/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680149 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 96 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiac Arrest | Tachycardia | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella cp post procedure for st-elevation myocardial infarction and had to be left in a deeper position due to hypo-kinesis of left ventricle. the site was appropriately dressed and groin best practices were used. no bleeding or hematoma at the groin site were noted. during support, it was noted that the patients vitals began to crash and code was called. advanced cardiac life support (acls) was provided for a total of five minutes which obtained return of systemic circulation. an echocardiogram showed deep placement. however, the physician was unable to get the catheter in a better position so it is still currently around five centimeters in the ventricle. the patient's feet were noted to be cool bilaterally so bilateral lower extremity (ble) arterial doppler was completed. per the preliminary report, both legs had adequate flow. the patient was still mechanically ventilated and sedated. the patient did suffer a pulseless electrical activity arrest while turning the patient requiring two rounds of cardiopulmonary resuscitation and acls medications. in addition, it was noted that the echocardiogram was completed and the impella cp was pulled back to a new landmark. a left ventricle thrombus was noted on echocardiogram from the previous day. however, it was noted to be larger. the patient had a ventricular tachycardia arrest and the family decided to make the patient do not resuscitate. the patient expired on support. | [5 empty] | The investigation for ventricular tachycardia (vt,) positioning issues, low pump flow, and thrombus has been completed. clinical details state that the patient suffered a pea arrest overnight. patient had vtach arrest and family decided to make patient dnr. the root cause of the pea arrest/vtach arrest was unable to be determined due to insufficient clinical details. the cause of the positioning issues was most likely patient condition related as the clinical details stated the pump was left deep due to hypokinesis of lv. the cause of the low pump flow was most likely patient condition related as clinical details noted rv dysfunction, high svr, hypokinesis of the lv, and giving volume temporarily resolved suction, which corresponded with the trends seen in the data logs. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7660 | 1220648-2025-30106 | 22435576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435576&pc=OZD | 06/19/2025 | 08/07/2025 |  | 07/08/2025 |  |  | 06/19/2025 | 07/08/2025 |  |  | 02/28/2026 |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478367 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female | 89 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, the pump was positioned too deep. while trying to reposition, the pump was pulled back across the aortic valve. the staff was unable to reposition. the pump was explanted and replaced with intra-aortic balloon pump. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical information states that the impella was in too deep earlier and was trying to reposition. the device was pulled back across the aortic valve (aov) while trying to reposition. they were unable to recross the aov. the decision was made to remove and place an intra-aortic balloon pump. the cause of the positioning issue was most likely use related as the pump was pulled back across the valve while repositioning and unable to recross. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7661 | 1220648-2025-30107 | 22435661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435661&pc=OZD | 06/19/2025 | 08/18/2025 |  | 07/08/2025 |  |  | 06/19/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 62 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was placed to support a patient admitted in with cardiogenic shock/acute myocardial infarction. the pump was supporting and percutaneous coronary intervention was performed. the pump was explanted and escalated to an impella 5.5 after hemolysis was observed on the impella cp. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: complaint device not returned for analyzing. clinical stated, on 1st day of impella support hemolysis issue noted witch resolved after impella repositioning. data logs reviewed: no motor current (mc) spike outside of p-level changes observed. quick resolved impella in aorta/ ventricle alarms seen. no suction or placement signal issues occurred. the cause of hemolysis issue most likely was improper positioning due to improper technique.
7662 | 22435747 | 22435747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435747&pc=OZD | 06/23/2025 | 07/02/2025 | 07/02/2025 | 07/08/2025 | 07/02/2025 | 07/08/2025 |  | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 60 YR |  | Female | 65 KG |  |  | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.5 alarm device failed. impella cord was disconnected and reconnected per alarm instructions impella 5.5 device continued to alarm "device failure".
7663 | 1220648-2025-30108 | 22435770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22435770&pc=OZD | 06/09/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/09/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587496 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 support was hit positive and argatroban was administered. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7664 | 1220648-2025-30109 | 22436139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436139&pc=OZD | 06/15/2025 | 11/19/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 was attempted to be primed in the standard fashion. however, once impella was priming, the aic immediately stated that there was a placement signal not reliable alarm. the impella and the aic were exchanged. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console logs for the complaint date were not available. reviewing the lt log data for aic ic10193 shows "placement signal not reliable" alarms (#1036 and #1048) that triggered immediately after the pump was plugged into the console, which aligns with the reported complaint. constellation pump-usage records show the pump was initially connected to aic ic8022, which exhibited the same placement-signal behavior. rt logs displayed unreliable placement signal throughout the pump usage for both aics. device analysis: console ic (b)(6) was not returned for further investigation. root cause: the root cause for the optical signal issue i.e. placement signal not reliable alarms, was not determined because the aic was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
7665 | 1220648-2025-30110 | 22436180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436180&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673546 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient was on bipella support in the icu. the patient was taken to the or for a fasciotomy. vascular still is being consulted as the patient might need additional intervention. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7666 | 1220648-2025-30111 | 22436407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436407&pc=OZD | 06/16/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671569 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ABIOMED GUIDEWIRE 0.018IN, V-18 GUIDEWIRE 0.018IN. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp for mechanical circulatory support (mcs); on this support for approximately four days before being escalated to an impella 5.5 due to ongoing cardiogenic shock. during insertion of the impella 5.5, at the point near the innominate artery, the pump would not advance past despite multiple attempts. the impella 0.018 guidewire was exchanged to a v18 stiffer guidewire. however, impella 5.5 device was still unable to advance past the innominate into the aorta. the decision was made to abort the case. the patient survived the impella cp explant and non-placement of the impella 5.5 device and was noted to be in stable condition following the event. any treatment to the patient thereafter was unnoted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the pump would not advance past with multiple different attempts and was unable to advance past the innominate, into the aorta. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use (ifu): section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿
7667 | 1220648-2025-30112 | 22436429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436429&pc=OZD | 06/16/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Malposition of Device | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device was mechanical circulatory support and experienced malpositioning of the device with hemolysis leading to intervention and device explant. the patient was taken to the catheterization lab for stent placement to the mid circumflex artery. during this procedure, the patient became hypotensive, and the decision was made to place the patient on an impella cp device pump. the impella cp was inserted utilizing micro puncture with good position noted on fluoroscopy. the peel away sheath was removed utilizing best practices and the repositioning sheath was placed with forward tension sutures. angle matched sterile dressing was placed over the impella site with perclose sutures intact in stopcock. approximately 11 hours after implant, an intensive care unit (icu) update noted the patient continued to rest on impella support. the patient was noted to have red urine and hemolyzing labs, with the impella cp measuring deep via echocardiogram. the impact of poor positioning was discussed and the team advised they would reposition the pump. now 14 hours after implant, the customer support team was contacted as the impella had positioning issues and position in aorta alarms. the performance level was reduced to p-2, and the alarm was not silenced. the fellow attempted to reposition the impella cp, was unsuccessful, and the pump migrated to the aorta. the pump was successfully exchanged for a new impella cp using the repositioning sheath re-access port. the patient was reported to be stable following the event. | [5 empty] | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. data log reviewed, no motor current (mc) spike or placement signal issues observed. impella position in aorta alarms seen at p2 which correlated with incident impella cp migrated to aorta when fellow was attempting to reposition. positioning issues: the cause of positioning issue most likely use related since impella cp migrated to aorta when fellow was attempting to reposition. hemolysis: urine was red and labs result hemolyzed, impella measuring deep on echo. the cause of hemolysis issue most likely was improper positioning related due to improper technique. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section a4. patient weight is unknown. as noted in the impella instructions for use (ifu): impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7668 | 1220648-2025-30113 | 22436592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436592&pc=OZD | 11/04/2024 | 07/18/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530258 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A user report was received and indicated a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced complications. the impella cp was implanted via the right femoral to support complex protected percutaneous coronary intervention (pci) with presence of thrombotic material in main stem, ramus interventricularis anterior and circumflex arteries, and the patient experienced recurrent ventricular fibrillation. the impella cp device was removed after approximately two days of support. subsequently, severe ischemia of the right leg in acute thrombosis of the entire right pelvic axis was experienced with the need for surgical thrombectomy of the right pelvic axis and endarterectomy of the femoral fork as well as patch plastic with vascuguard patch the following day. the status of the patient was indicated as unknown. however, the patient was noted to have survived the event and explant. | Additional information was received on july 11, 2025 that the ventricular fibrillation was not due to the impella. |  |  |  |  | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions, ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | B5 ventricular fibrillation was not related to the impella device and was erroneously entered on the initial manufacturer device report number 1220648-2025-30113. h6 health effect: clinical code 2130 was erroneously entered on the initial manufacturer device report number 1220648-2025-30113.
7669 | 1220648-2025-30114 | 22436700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436700&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656764 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 66 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and post implant, bleeding was noticed by the innominate vein which required surgical intervention. the patient decompensated quickly after extracorporeal membrane oxygenation (ecmo) decannulation, prompting an emergent impella 5.5 device placement via the right axillary subclavian artery with successful cut down and graft attachment of an 8mm by 30cm vascutek gelweave. multiple diagnostic catheters and guidewires were used to cross the aortic valve into the left ventricle. the impella was inserted over the v18 0.018inch guidewire and measured 5.1 cm from the aortic valve annulas with inlet pointing slightly posterior. prior to closing the site but post insertion, the surgeons noticed bleeding by the innominate vein that got bigger as it was repaired it. no resistance was noted during impella insertion. the vein was lacerated centrally underneath the right clavicle. however, the axillary artery was not compromised. the physician was unable to get to the site and repair causing an emergent sternotomy and vascular surgery to help disarticulate the right clavicle to repair the innominate vein. the patient was crashed on cardiopulmonary bypass emergently, and the site was repaired successfully. the site was then successfully closed; however, the transesophageal echocardiogram (tee) showed the impella was pointing posteriorly. the patient was treated with blood products as well and sent to the unit hemodynamically stable. latest update indicates the patient continues on the impella 5.5 mcs. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use (ifu): impella 5.5 with smartassist system. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
7670 | 1220648-2025-30115 | 22436747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436747&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672975 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Hospitalization | Death | Life Threatening | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was escalated from an intra-aortic balloon pump (iabp) to an impella 5.5 device mechanical circulatory support and experienced bleeding and arrhythmic events requiring intervention. the impella implant procedure took place in an operating room. the patient was intubated, prepped, and draped. the physician performed a cutdown to the right axillary artery and a10mm vascutek gelweave was anastomosed to the artery. the peel away sheath was advanced into the graft and secured with a graft lock. access to the left ventricle was obtained and the impella 5.5 was inserted over the guidewire into the left ventricle under fluoroscopy and transesophageal echocardiogram guidance. the impella was started at performance level p-2 and ramped up to p-7. the peel away sheath was removed, graft trimmed, and the blue suture hub was advanced and secured to the graft. the yellow pin was removed, and the blue suture hub was secured to the patient. dressing was placed to the impella site and the iabp was removed from the right groin with femstop placement. the patient was transported back to the cardiovascular intensive care unit for continued monitoring. in the evening, it was noted the impella site remained clear, dry, and intact, with no evidence of bleeding or hematoma. the impella three-point fixation was in place. however, the femstop placed to the right groin where the iabp was bleeding. it was noted heparin drip was used and may have contributed to the bleeding. manual pressure was held. one day later, the dressing was changed by the practitioner. while the dressing was clear, dry, and intact, with no evidence of bleeding, a small hematoma was noted at the impella site. now four days post start of the impella 5.5 support, the impella insertion site was noted to have a minimal amount of bloody drainage with stable hematoma. no intervention was planned; the clinical rep confirmed that there was no known intervention for the hematoma. staff monitored; the physician was aware. the next day, the patient was having episodes of ventricular tachycardia storm, despite the automated implantable cardioverter defibrillator (aicd). the pump settings were changed, and the patient was started on heparin to aggressively titrate for therapeutic range, as the patient had not been previously started on heparin prior to hospital transfer. overnight, the next day, the patient had ventricular tachycardia and was given amiodarone bolus. the patient continued to have persistent small runs of ventricular tachycardia. an echocardiogram was performed at bedside and confirmed the pump was in stable position. the patient's belly was distended so the bowel regimen was increased. the patient was noted to have emesis of over one liter with increased lactic and distended belly. a computed tomography demonstrated bowel ischemia. after discussing care and goals with the family, the patient was placed on do not resuscitate status and chose to withdraw care. the patient ultimately expired. | [5 empty] | The patient expiry occurred upon withdrawal of support. however, there is not enough evidence to exclude the impella 5.5 device used as an associated factor in the patient's outcome. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7671 | 1220648-2025-30116 | 22436985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22436985&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635404 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella cp support had oozing at the impella access site and fem stop was placed. bleeding resolved. additionally, the patient was not moving the left side and ct scan revealed large hemorrhagic stroke. inr was 3 and anticoagulation stopped. ct scans showed stable hemorrhagic stroke. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smart assist. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7672 | 1220648-2025-30117 | 22437020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437020&pc=OZD | 06/16/2025 | 10/07/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611447 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Tamponade | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 on a patient for support during high-risk cardiac procedure. it was further reported that the patient had high output from chest tubes after clot was expelled. the physician feels that chest tubes clotted leading to tamponade and was collecting blood in chest. blood products were administered, and heparin was held. patient survived. | The complainant also reported that, in the intensive care unit, the patient lost pulsatility while coughing and noticed on the automated impella controller that the placement signal disappeared. the placement signal did return to impella connect shortly after it disappeared. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue, priming issue, thrombus, and vascular damage. no more issues noted with placement signal. upon review of data logs, it is apparent that the console is the potential cause of the issue. the root cause of the priming issue was most likely use-related since clinical details state that the wire was dropped during wire exchange data logs are not applicable and product was not returned. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the root cause of the vascular damage was not determined since insufficient clinical details were provided and product was not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-30117. d10 concomitant medical products and therapy dates updated.
7673 | 1220648-2025-30118 | 22437216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437216&pc=OZD | 06/16/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1870819 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | Appropriate Term/Code Not Available | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered issues with the automated impella controller (aic). there was a tag on the aic saying ¿broken¿ without any context or further clarification. the clinical engineer asked that it go through the standard maintenance process to ensure its safe. there was no reported patient involvement. | [5 empty] | Data analysis: the logs are littered with constant failed handshakes between the automated impella controller (aic) and rlm. device analysis: the aic was booted up and the logs from the field were reproduced. the rlm was not getting through boot up as it was persistently partition switching due to a corrupt usd card. since the rlm would not successfully boot up, the aic could not successfully handshake. replacing the usd card alleviated the failure. root cause: the rlm's inability to boot up was caused by a corrupt usd card. the cause of the corrupt usd card was not determined. h6 medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2025-30118. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
7674 | 1220648-2025-30119 | 22437274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437274&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661329 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella cp pump stopped unexpectedly during patient support. the pump was removed as a result of the noted failure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7675 | 1220648-2025-30120 | 22437680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437680&pc=OZD | 06/20/2025 | 08/15/2025 |  | 07/08/2025 |  |  | 06/20/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677395 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 87 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a surgical/mitraclip patient placed on impella cp for support. the pump failed to start and was therefore replaced by a new impella cp pump. there was no patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of pump did not start has been completed. the returned product was investigated and there were no visual abnormalities. the pump passed electrical testing of all the motor cable and data communication lines. the pump was successfully recognized by the console and all areas of the eeprom read successfully (per imc logs). the pump was finally run in a bucket for 15 minutes, and pump usage data updated (per imc logs). therefore, the reported issue was unable to be reproduced. since clinical details are limited, available imc logs appear incomplete, and the reported issue did not reproduce, the root cause of the pump did not start could not be determined.
7676 | 1220648-2025-30121 | 22437725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437725&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657463 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that a patient had a rhc (right heart catherization) and was admitted for cardiogenic shock. iabp was placed and was scheduled for 5.5 for advance option work up. however, it was noted that the 5.5 sterility was compromised and the case was aborted. | [5 empty] | The investigation of priming issue has been completed. the impella device was not returned for evaluation. priming issue: clinical details report that the impella 5.5 was dropped out of the sterile field. it was replaced with a new pump. product was not returned for investigation. based on the clinical detail that the pump was dropped out of the sterile field and replaced with a new one, the root cause of the priming issue was determined to most likely be a use issue. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7677 | 1220648-2025-30122 | 22437891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437891&pc=OZD | 03/10/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 03/10/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611547 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in the eu had a cp pump placed for the support of a patient admitted in for a hrpci. the cp was explanted after the pci and the access site was bleeding. the bleed was treated by 2 units of rbcs and proglide perclose x2. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7678 | 1220648-2025-30123 | 22437945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22437945&pc=OZD | 06/15/2025 | 08/08/2025 |  | 07/08/2025 |  |  | 06/15/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655963 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 80 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp and experienced bleeding at the groin site which was resolved with a figure eight dressing. the patient's family decided to withdraw care and the patient expired the following day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since non-abiomed product was used which is not recommended per instructions for use.
7679 | 1220648-2025-30124 | 22438054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438054&pc=OZD | 06/20/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/20/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621750 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was post-op coronary artery bypass graft x 4 and hemodynamically decompensated requiring increased inotropic and pressor support. the decision made to bring the patient to the operating room for extracorporeal membrane oxygenation and impella cp placement. the impella cp was advanced into the left ventricle without complication. the peel-away introducer was removed and the repo sheath was advanced. bleeding was noted and did not occlude the insertion site. the perclose suture was attempted to be tighten, however, it did not help the bleeding. the decision was made to remove the impella cp. a 16 french dry seal sheath was placed and a new impella cp was placed through the same insertion site. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7680 | 1220648-2025-30125 | 22438257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438257&pc=OZD | 06/16/2025 | 08/07/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635341 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 59 KG |  | Asian | Malfunction | No | Yes | Detachment of Device or Device Component | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the pump was attempted to be repositioned and the anticontamination sleeve was broken. the pump was later attempted to be repositioned again and the pump went across the valve. the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage) and positioning issues has been completed. the impella cp was not returned for evaluation. the cause of the positioning issue was most likely use related as the pump was pulled across the valve during repositioning and was unable to recross. the cause of the sterile sleeve damage was most likely use related as it occurred during repositioning of the pump. a.4 added weight as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-30125. b.7 added medical history as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-30125. h.6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30125. code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2025-30125. codes 4627 and 4628 were added.
7681 | 1220648-2025-30126 | 22438269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438269&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609627 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support and developed a hematoma at the groin site. the site was redressed, and manual pressure was held. heparin was on hold until hemoglobin and hematocrit were stable. the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7682 | 1220648-2025-30127 | 22438426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438426&pc=OZD | 06/19/2025 | 08/16/2025 |  | 07/08/2025 |  |  | 06/20/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026686410 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 59 KG | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a placement signal not reliable alarm appeared during impella 5.5 support. despite losing the pulmonary stenosis and left ventricle readings, the motor current and flow rate remained stable. support was continued uninterrupted and the patients outcome was stable. | [5 empty] | The investigation of optical signal issue has been completed. the pump was returned for investigation. the returned pump had no physical damage on it, and all the 5s optical parameters were within the specification range. data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed on 19-jun-2025 at this time, snr drops to zero, lamp maxes out at 100%, light decreases, gain decreases, and contrast oscillates between +/-3,000. the real time (rt) logs were closely examined during the time of the psnr event, and the oscillating contrast trend was seen. the pump was remained in use during psnr. optical signal parameters and placement signal was seen returned after 68 hours of case run. upon review of data logs and returned product investigation, it is apparent that the console is the potential cause of the issue. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7683 | 1220648-2025-30129 | 22438481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438481&pc=OZD | 06/16/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667210 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella cp support and during support, the patient had oozing at the access site. heparin was withheld. support was continued. the pump was successfully weaned and explanted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time (act) values, heparin was stopped to achieve hemostasis as per the clinical description.
7684 | 1220648-2025-30128 | 22438483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438483&pc=OZD | 06/17/2025 | 08/15/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1750720 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had impella support and after the implant, there were placement signal not reliable alarms. the automated impella controller (aic) was unplugged and plugged back in to no avail. the aic was then replaced and alarm resolved. patient later expired. | [5 empty] | Corrected information provided in h6. type of investigation fda code device not returned (4114) was reported in error. the device was received for evaluation. | The patient's death is being reported under complaint number 25-45652-2. data analysis: inspection of the aic logs show the pump was recognized when initially connected and communication between the epc and the optical bench begins. when pump flow is initiated (p-level 9) approximately 14 minutes later, ¿processsampleforopminvalidsignalerror: calculator placement signal 1036¿ appears in the logs, followed quickly by the presence of a ¿placement signal not reliable¿ alarm. flow is stopped and the pump is disconnected and reconnected. communication between the epc and the optical bench resumes. flow is resumed (p-level 9) and is again quickly followed by ¿processsampleforopminvalidsignalerror: calculator placement signal 1036¿ in the logs. a suction alarm is also noted. after the appearance of a second ¿placement signal not reliable¿ alarm, the pump is disconnected and the console is shut down. examination of the rt logs show that when the pump is initially connected, snr values are acceptable at approximately 5000. when pump flow begins, snr levels fall precipitously to values below 500. the aic logs confirm the complaint of unreliable placement signal both when the pump is initially connected, as well as after it is disconnected and reconnected. see below for all attachments. device analysis: 5s parameters and an optical bench fringe pattern were acquired. 5s parameters were all in the acceptable range and the fringe pattern showed an amplitude of 2.6v. the console and a test pump were placed on a bench and allowed to run overnight at p-level 4. no placement signal alarms occurred. the console was opened and the cables and connectors between the optical bench and epc were inspected with no integrity issues found. the console was further tested with two different pumps with a p-level setting of 9. inspection of the newly generated rt logs did not show the same drop in snr at high p-levels that occurred in the initial logs. the complaint of placement signal alarms was never reproduced using the test pumps available. the optical connector on the console was cleaned with a ¿one-click¿ fiber cleaner. root cause: because the complaint was not reproduced, the root cause is undetermined. before returning the console to the customer, the following actions are instructed to be performed on (b)(6) : ¿ replace the pbm (0042-1125) replace the optical bench (0042-3015p) perform sections 18-20, 23 of the pm (0042-7309) perform a full functional check (0042-7316)
7685 | 1220648-2025-30130 | 22438779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438779&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 61 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support and developed thrombocytopenia. the patient was switched to bival. the plan is to escalate to an impella 5.5. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7686 | 1220648-2025-30131 | 22438791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438791&pc=OZD | 06/17/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660980 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Positioning Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had ventricular tachycardia requiring 1 shock at 200j. support continued. additionally, the patient's ldh levels were rising the pump was repositioned and pulled back. the clinical rep confirmed that the patient had no urine output so hemolysis could not be visualized and confirmed. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7687 | 1220648-2025-30132 | 22438970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438970&pc=OZD | 06/16/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679246 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 support and during support, the patient had oozing at the access site after partial thromboplastin time (ptt) (101.3) and anti xa assay (0.51) supratherapeutic. a stitch was placed at the jp drain site after trauma during bed bath. improvement was noticed and heparin was withheld for ptt to come down. support was continued and the patient's outcome was stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anticoagulation management because the partial thromboplastin time (ptt) values were higher than therapeutic values and heparin was stopped and extra suture was place to achieve hemostasis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7688 | 1220648-2025-30134 | 22438999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22438999&pc=PYX | 06/17/2025 | 10/07/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  | 11/30/2025 |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024435395 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient underwent emergent surgery for an aortic dissection. postoperatively, the chest was initially left open due to clinical instability and was closed at a later date. following chest closure, the patient developed a decline in right ventricular (rv) function. as a result, the decision was made to place an rp flex device in the cardiac catheterization lab. during rp support, it was noted that flow at performance level p-6 decreased from 2.7 liters per minute to 1.5 liters per minute, without any associated change in motor current or systemic hemodynamics. the dynamic positioning system (dps) was re-zeroed, which corrected the flow issue. a similar event occurred approximately six hours later, requiring another re-zeroing of the dps. the rp flex device was ultimately weaned and explanted successfully without complication. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. impella rp flex was returned. pump was visually inspected and biomaterial was observed beneath the impeller at inflow. no other physical damages to the sensor were observed. data logs were returned and confirmed issue. the root cause of the placement signal issue was sensor drift with drift reproducing during flow loop test and based on log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. corrections that were made from the initial manufacturer device report number 1220648-2025-30134. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella pump was returned. based on log review, there was no true low flow observed apart from sensor drift causing pump flow drifting down. there was no true low flow issue found during the case. the device functioned/operated to specification.
7689 | 1220648-2025-30135 | 22439177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22439177&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella 5.5 was pointed at the mitral valve. attempts to reposition the pump were unsuccessful. additionally, there were placement signal not reliable alarms. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation of positioning issues and optical signal issue has been completed. the impella device was not returned for evaluation. optical signal issue: the data log did not record any placement signal not reliable (psnr) alarms on the day of the reported issue. there were no issues found with the 5s optical signal. there was no issue found on optical signal during the reported time of psnr event. the device functioned/operated to specification. positioning issue: the cause of the positioning issue was most likely user-related, as the pump was found near the mitral apparatus during 18th day of support, without any reported issues with the catheter anchor on the pump.
7690 | 1220648-2025-30136 | 22439214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22439214&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/16/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671929 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 119 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support gastrointestinal physician was consulted for melena and some blood tinged urine. hemoglobin was relatively stable. systemic heparin was withheld and the purge was changed to bicarbonate. support was continued. the clinical representative confirmed the patient did not have a gastrointestinal bleed but they had a hemorrhoid. patient support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7691 | 1220648-2025-30137 | 22439701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22439701&pc=OZD | 06/17/2025 | 11/11/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680845 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 82 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that resistance was encountered when attempting to implant an impella 5.5 into the subclavian artery and through the innominate. the pump was lubricated and attempted multiple times without success. the patient survived and was prepped for an intra-aortic balloon. | [5 empty] | The root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to tortuous anatomy with calcifications and inability to pass 5.5. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7692 | 1220648-2025-30138 | 22439992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22439992&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616848 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 106 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to a kink in the catheter. a second impella cp was used and implanted successfully. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to use issue because the second pump was successfully inserted without any issues after removing the first pump. product damage(cannula kink): the root cause of product damage is determined to be use issue because the doctor inserted the sheath first before inserting the cannula and removed the impella along with the sheath with difficulty and cannula was kinked during the procedure and the second pump was successfully inserted.
7693 | 1220648-2025-30139 | 22440241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440241&pc=OZD | 05/11/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653486 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male | 91 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device, non-sustained ventricular tachycardia with a "possible" relationship to the impella device, and sepsis with a "possible" relationship to the impella device. the patient was started on lidocaine and amiodarone push and then started on both ggts. the patient converted back to normal sinus rhythm. the patient was started on vancomycin. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7694 | 1220648-2025-30140 | 22440290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440290&pc=OZD | 06/17/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664186 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp had hemolyzed labs and discolored urine. the patient also had a thrombus in the left ventricle. two days after implantation the patient coded and expired. | [5 empty] | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7695 | 1220648-2025-30141 | 22440317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440317&pc=OZD | 06/15/2025 | 08/01/2025 |  | 07/08/2025 |  |  | 06/15/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678123 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella 5.5 support. while the patient was in a sitting position and attempted to stand for ambulation, they accidentally stepped on the white cable. this caused the first fixation point to detach and eventually led to the red plug becoming disconnected from the catheter, resulting in a pump stop. the impella was quickly explanted at bedside with no issues, and the pocket was closed successfully. | [5 empty] | The investigation of pump stop has been completed. the data logs show a stable pump run and then at the end of the case there was a motor current spike to 30000 with 115 alarm. the pump was returned and the catheter was pulled out of the kink protector and all motor cable lines were stretched. there was glue seen on the patient cable. the root cause of the pump stop was most likely due to open electrical lines based on the clinical details and returned product analysis. d9 device returned to manufacturer date added. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7696 | 1220648-2025-30142 | 22440499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440499&pc=OZD | 06/20/2025 | 08/19/2025 |  | 07/08/2025 |  |  | 06/20/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679776 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 65 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they were unable to pass the impella 5.5 through the peel-a-way sheath. it was further reported that after the right axillary eight millimeter graft was placed, the peel-a-way sheath was in place. the doctor attempted multiple times to pass the impella 5.5 through the peel-a way sheath with no success. after trying multiple times, the white catheter behind the motor kinked and they lost all pushability. they said that it felt like the sheath was kinked or narrowed at the distal tip. the impella 5.5 was replaced and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of difficulty inserting/removing pump through introducer and product damage (cannula kink) has been completed. product damage (cannula kink) and product damage (sheath kink) was added based on the investigation. the investigation revealed that the pump could not be inserted due to a sheath kink so that failure mode is just against the sheath. the clinical details state that "after the r axillary 8mm graft was placed and the peel-a-way sheath in place. dr linden attempted multiple times to pass the impella 5.5 through the peel-a way sheath with no success. after trying multiple times, the white catheter behind the motor kinked and she lost all pushability. she said that it felt like the sheath was kinked or narrowed at the distal tip. excessive force was applied. there was no imaging of the artery before, artery was small, dilation was not done and resistance was met. patient had chest trauma. the returned sheath had a large kink. there were no over defined score lines seen or damages to the valve. the pump was noted to have a kink in the cannula and a kink in the catheter near the motor housing. product damage (sheath kink): based on the clinical details, the root cause of the product damage (sheath kink) is most likely due to the patients condition as they used an 8mm graft and the artery was small/diseased. difficulty inserting/removing the pump through the introducer: the pump was attempted to be passed through the introducer but was unable to as there was a kink in the sheath causing kinking of the pump. based on the returned product analysis and clinical details, the root cause of the difficulty inserting/removing the pump through the introducer is most likely due to the patients condition of having diseased/small vessels which caused a kink in the introducer and then excessive force was applied to the pump trying to pass it through. product damage (cannula kink). the pump was attempted to be passed through the introducer but was unable to as there was a kink in the sheath causing kinking of the pump cannula. based on the returned product analysis and clinical details, the root cause of the product damage (cannula kink) is most likely due to the patients condition of having diseased/small vessels which caused a kink in the introducer and then excessive force was applied to the pump trying to pass it through, further kinking the cannula and catheter. | The impella pump and introducer were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
7697 | 1220648-2025-30143 | 22440509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440509&pc=OZD | 10/09/2023 | 07/08/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Difficult to Remove | Malposition of Device | Deformation Due to Compressive Stress | Detachment of Device or Device Component |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The publication entitled: "making matters worse with impella" states: despite repeated attempts at repositioning, flow rates remained < 1 l/min. transesophageal echocardiography and thoracic x-ray confirmed that the catheter had bent in, upon itself, and the pigtail had been caught in the outlet following the multiple attempts at repositioning. the catheter migrated toward the descending aorta and urgent removed was decided. an attempt to unscrew the pigtail through femoral manipulation proved unsuccessful. a 7-fr amplatz super-stiff guidewire (boston scientific, united states) was advanced through the impella loop via left radial access, and traction on the catheter from the femoral access successfully disengaged the pigtail from the outlet. concurrently, a 12 mm × 20 mm en-snare loop (merit medical, united states) was inserted to capture the tip of the pigtail. the guidewire was removed, and traction was applied simultaneously to the impella device and the pigtail tip, which extended the impella and allowed its extraction through its introducer. the impella was removed. | [5 empty] | Patient data was not provided. section a was left blank. device details were not provided. section d4 was left blank. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7698 | 1220648-2025-30144 | 22440560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440560&pc=OZD | 06/13/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655756 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Unknown | 60 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiogenic Shock | [11 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced limb ischemia and expired from prolonged cardiogenic shock. | [5 empty] | A3 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7699 | 1220648-2025-30145 | 22440765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440765&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670089 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 87 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. heparin was withheld to achieve homeostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was determined to be anticoagulation management because hemostasis was achieved by turning off the heparin.
7700 | 1220648-2025-30146 | 22440838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22440838&pc=OZD | 06/17/2025 | 07/08/2025 |  | 07/08/2025 |  |  | 06/17/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660775 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 145 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced slow ventricular tachycardia (vt) and atrial flutter. the team tried shocking the patient and when the patient went out of vt, they were in third degree heart block. the team considered the use of a temporary pacer. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients.".
7701 | 1220648-2025-30147 | 22441246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22441246&pc=OZD | 06/18/2025 | 08/15/2025 |  | 07/08/2025 |  |  | 06/18/2025 | 07/08/2025 |  |  |  |  | 07/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673726 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that a patient presented for scheduled left main percutaneous coronary intervention (pci). pci was attempted without support however the patient decompensated. impella cp was opened and sheath insertion attempted. md was unable to advance sheath due to vascular anatomy. the impella insertion was aborted and iabp was inserted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the introducer sheath cannot be advanced into the patient due to patient's vascular anatomy and the insertion was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7702 | 1220648-2025-30060 | 22410490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410490&pc=OZD | 06/17/2025 | 07/16/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660814 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 60 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. prior to being placed on the impella cp, the patient arrived with complaints of chest pain/pressure ongoing for 4-5 days prior. it was planned for the patient to be placed on an intra-aortic balloon bump (iabp), but the patient immediately pulseless electrical activity (pea) arrested. code was called and the decision was made to place the impella cp. during support, the patient experienced limb ischemia bilaterally. mottling and temperature changes were noted. the decision was made to send the patient for vascular surgery to explant the device and repair. subsequently, it was reported that the patient's labs were worsening. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
7703 | 1220648-2025-30061 | 22410537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410537&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636422 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 99 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp was placed to support a patient admitted in with an acute myocardial infarction/cardiogenic shock and arrest. the pump was placed but there were persistent low flow/suction events and the patient's limbs were ischemic bilaterally. there was bleeding from the groin that prompted treating with an extra suture, femstop, and transfusion of two units of packed red blood cells. despite all measures, the patient expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of low pump flow, limb ischemia, and access site bleeding has been completed. the impella device was not returned for evaluation. low pump flow: the data log shows several suction alarms throughout the case run, with an observed downward trend in placement signal and pump flow. pump flow was seen gradually decreasing while running on the same p-level. a motor current spike was reported in the middle of the case without impacting major issues in pump performance. the motor current spike was indicative of biomaterial interaction during the case; therefore, the thrombus failure mode was added as an investigated failure mode. according to the clinical notes, the patient was on pressure support, and blood products were given due to blood loss during support. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. limb ischemia - reversible: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. access site bleeding-major: the cause of the access site breeding issue was not determined since product was not returned and sufficient clinical information was not provided. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7704 | 1220648-2025-30062 | 22410636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410636&pc=OZD | 06/17/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  | 11/30/2025 |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426647 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced recurrent ventricular tachycardia and ventricular fibrillation during support. the patient was treated with lidocaine. after nine days of impella 5.5 support, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7705 | 1220648-2025-30063 | 22410724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410724&pc=OZD | 05/01/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629145 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured an event with a "possible" relationship to the impella device used: ¿pt with acute kidney injury stage 3. baseline pre impella creatinine of 1.00 mg/dl on (b)(6) 2025 at 03:43. on (b)(6) 2025 creatine rose to 1.73 mg/dl at 03:59. (b)(6) 2025 creatine was 2.72 mg/dl at 03:37. (b)(6) 2025 creatine was 4.03 mg/dl and nephrology was consulted. of note, pt had ¿coca cola colored urine¿ (b)(6) 2025 ¿indicative of hemolysis.¿ felt to have non-oliguric acute kidney injury related to contrast nephropathy, shock/hypotension and pigment nephropathy secondary to hemolysis. creatinine on(b)(6) 2025 was 5.85 mg/dl at 12:02. hemodialysis initiated on (b)(6) 2025. creatinine continued to fluctuate and peaked on (b)(6) 2025 at 05:27 with a value of 7.85 mg/dl. hd continued throughout index admission. on day of discharge from index admission ((b)(6) 2025) creatinine was 5.16 mg/dl at 05:47. pt discharged with plan to continue hd with hopes of renal recovery. pt completed dialysis on unknown date. creatinine on (b)(6) 2025 was 2.03 mg/dl at 06:40.¿. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7706 | 1220648-2025-30064 | 22410822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410822&pc=OZD | 06/17/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672968 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 89 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient went back into atrial fibrillation approximately 18 hours post impella 5.5 pump insertion. the patient had been escalated to and expired on the impella 5.5. however, the impella cp related adverse events were more likely associated with the demise outcome. there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome, which is captured under the impella cp complaint file (b)(4). | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. abiomed medical safety review: the patient was escalated to and expired on the impella 5.5 but is a serious injury for the atrial fibrillation. the patient went back into atrial fibrillation approximately 18 hours post impella 5.5 insertion. however, the impella cp related adverse events (e.g., limb ischemia access bleeding requiring blood transfusion etc.) are more likely associated with the demise outcome. there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome, which is captured the impella cp complaint file (b)(4).
7707 | 1220648-2025-30065 | 22410890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410890&pc=OZD | 06/18/2025 | 08/15/2025 |  | 07/07/2025 |  |  | 06/18/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1643465 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 85 KG | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced a console failure of the supporting automated impella controller (aic). the aic turned black then returned to case start screen. the team was able to restart the impella, and no patient harm was reported. | [5 empty] | Additional information provided in h3, h6 and h10. the console was returned for investigation and testing. the unexpected shutdown was not reproduced while replicating the clinical actions but the subprocess could be forced to crash causing a partial reboot of the system. the impella positioning alarms were similarly not reproduced and optical performance by the console was as expected so the alarms mentioned in the -2 are unlikely to be due to the console. data logs were reviewed and showed that the calculator subprocess crashes triggering a partial reset of the console by the software watchdog. the console boots back up and hemodynamic support is resumed. there are constant impella position unknown and impella position in aorta alarms throughout the case and on 6/19 the user disables placement monitoring as described in the complaint. the unexpected reboot by the console was due to the calculator subprocess crashing, triggering the watchdog. the cause of the calculator to crash is unknown as it was not reproduced. as a precaution, the 4-in-1 and compact flash cards were replaced and a functional check of the console was completed. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7708 | 1220648-2025-30066 | 22410922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22410922&pc=OZD | 06/17/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/18/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148257 |  | 2855 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SERIAL NUMBER (B)(6). | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella automated impella controller (aic) screen was showing a purge connection failure while supporting a patient implanted with an impella cp. the aic was exchanged and no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7709 | 1220648-2025-30067 | 22411011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411011&pc=OZD | 06/18/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/18/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637885 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 57 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had a gastrointestinal bleed. there were concerns about red streaking in recent stools. gastrointestinologist was consulted and heparin was withheld. support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7710 | 1220648-2025-30068 | 22411082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411082&pc=PYX | 06/14/2025 | 08/11/2025 |  | 07/07/2025 |  |  | 06/14/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025587567 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that multiple complications were encountered during impella rp flex support. two days into support, the registered nurse noted a decrease in the purge flow rate. the purge cassette was replaced, but the issue remained. tissue plasminogen activator was then added to the purge and the flow rate improved. the staff reported that it appeared as if a clot had been ingested, leading to the purge flow issues. the patient then developed heparin induced thrombocytopenia. heparin was stopped and agatroban was initiated in response to the condition. in addition, the flow rate of the pump dropped. a clot was seen near the motor at the time of the decrease in flow. per advice from the staff, the pump was explanted the following day. patient survived | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the high purge pressure, low pump flow, thrombus, and thrombocytopenia has been completed. without additional clinical details, the root cause of the thrombus and thrombocytopenia could not be determined. however, data logs were reviewed for the high purge pressure (hpp or pp) and the low pump flow which showed that there was purge system blocked alarms where pp rose for 53 minutes and resolved with a bag change. there were no motor current spikes or changes in flow before this. about 26 hours after, purge pressure began to rise over the span of about 21 hours to about 900mmhg and then slowly began to decrease as there was noted to be tissue plasminogen activator in the purge at this point. about 20 hours later there was a motor current spike but no changes in purge pressure or flows. purge pressure remained stable for the duration of the case. there were no long bag/cassette changes. the data logs further show that pump flows were within specification until about 125 hours in when there was a spike in motor current and ps followed by a reduced pulstaility and lower pump flows. the clinical stated that there was a visible clot seen near the motor. the pump was returned and there were no cannula kinks seen. biomaterial was seen around the impeller. the pump was run and flows were slightly saturated due to the biomaterial. based on the data log and returned product analysis, the root cause of the low pump flow is most likely due to ingested biomaterial. based on the data log and returned product analysis, the root cause of the high purge pressure events throughout support are most likely due to biomaterial.
7711 | 1220648-2025-30069 | 22411162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411162&pc=OZD | 04/29/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171598 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella 5.5 support and after the pump was implanted, there were placement signal not reliable alarms. audio was disabled and staff monitored. additionally, a pump stop occurred with motor spikes. the automated impella controller (aic) was switched out and the pump stopped several times, but was successfully restarted. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7712 | 1220648-2025-30070 | 22411187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411187&pc=PYX | 06/18/2025 | 08/14/2025 |  | 07/07/2025 |  |  | 06/18/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551311 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female | 118 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella rp flex was supporting for a cardiogenic shock patient. during support, the pump came out of position and wave forms look to be in the right ventricle. the pump was malpositioned so it was removed and replaced. the patient's outcome is stable. | [5 empty] | D.2 product code was entered incorrectly on initial manufacturer device report number 1220648-2025-30070. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
7713 | 1220648-2025-30071 | 22411237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411237&pc=OZD | 06/19/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/20/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519865 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump support placed for the cardiogenic shock patient. the cp was placed in the community and was then transferred to tertiary center and escalated to the 5.5 pump. the team observed thrombus at the explant and later limb fasciotomy was performed to the limb the cp had been placed. the patient survived the harm and remains on the 5.5 pump support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7714 | 1220648-2025-30072 | 22411328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411328&pc=OZD | 06/23/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/23/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657737 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (hrpci) was placed on an impella 5.5 for mechanical circulatory support. post explant of the device, the patient exhibited signs of stroke and thrombus was observed. the patient was scheduled for surgery, taken to the operating room, and upon induction, the patient decompensated. the decision was made to place the patient on an impella 5.5 and undergo hrpci within the next couple of days. on day five of support, the patient was noted to be doing "extremely" well, and the plan was to explant the impella 5.5 this morning was confirmed. however, as soon as impella 5.5 was explanted, the patient exhibited signs of slurred, garbled speech and stroke symptoms. stroke alert was activated. additionally, it was reported that there was thrombus/biomaterial observed. it is unclear if the thrombus/biomaterial was observed in the patient or on the device. the patient was reported to have survived the explant. however, the patient was listed in critical condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7715 | 1220648-2025-30073 | 22411359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411359&pc=OZD | 06/15/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/15/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576853 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the perfusionist confirmed persistently elevated purge pressures. tissue plasminogen activator was added to the purge solution to address the high purge pressure. despite intervention, purge pressures remained elevated. the pump was explanted due to the unresolved high purge pressure. | [5 empty] | Section a5 is unknown. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7716 | 1220648-2025-30074 | 22411832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22411832&pc=OZD | 06/23/2025 | 09/24/2025 |  | 07/07/2025 |  |  | 06/23/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026673220 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Dislodged or Dislocated | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support and approximately four days into support, the patient inadvertently dislocated the pump back into the aorta. upon readvancing the pump into the left ventricle, the pump unexpectedly stopped and was unable to be restarted. the pump was subsequently explanted. the patient was noted to be stable and awake on extracorporeal membrane oxygenation. additionally noted was given the patient's stability, no new impella device implant was planned at such time. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of positioning issues and pump stop has been completed. the cut pump(motor side) was returned for investigation. no physical damage was noted on the cannula and impellor. no issue was found on catheter anchor on returned product. no further investigation was performed due to the cut catheter. the remaining cut pump was not returned for investigation. data logs show several pump stops without any motor current spikes, accompanied only by alarm 119 (pump stop). several unsuccessful pump start attempts were recorded with motor current spikes up to 3000 and alarm 119 (pump stop) were noted. the cause of the pump stop issue was not determined since sufficient returned product was not returned and data logs was inconclusive. the cause of the positioning issue was most likely patient movement, based on clinical details and available returned product investigation. g1 reporting contact email revised on manufacturer device report 1220648-2025-30074 in accordance with updated procedures.
7717 | 1220648-2025-30075 | 22412043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412043&pc=OZD | 06/15/2025 | 09/10/2025 |  | 07/07/2025 |  |  | 06/15/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673854 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in the eu with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed ventricular tachycardia (vt), requiring repeated defibrillation. clinical staff reported a red alarm with the message "pump stopped, mc high." multiple unsuccessful restart attempts were made. the device position was verified, and impella support was adjusted to p9. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer as the device remains implanted. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the instructions for use for the impella cp with smart assist system states the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. | The investigation was completed. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details. there were no data logs returned. clinical details state that there was a motor current high pump stop and several attempts to restart. the product was returned and there was a large purge gap. no other abnormalities were observed. the product was run for 10 days. per the cp c9 design trace matrix the design life for the cp is 8 days. based on the clinical details and returned product analysis, the root cause of the pump stop is due to bearing wear that occurred beyond intended design life.
7718 | 1220648-2025-30076 | 22412150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412150&pc=OZD | 06/19/2025 | 08/15/2025 |  | 07/07/2025 |  |  | 06/19/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572051 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and had persistent low flow/suction events. repositioning the pump failed to resolve the issue. the pump was explanted and a kink was observed. a new impella cp was implanted and the patient survived. | [5 empty] | The investigation of low pump flow, product damage (cannula kink), and positioning issues has been completed. low pump flows: since data logs are inconclusive, clinical details provided are limited, and low flows did not reproduce on returned product, the root cause of the low pump flows could not be determined. product damage cannula kink: returned product was investigated, and there was evidence of a kink in the cannula near the motor housing, as reported. since the timing of the cannula kink is unknown and there are several potential caused based on the clinical details provided, the root cause of the cannula kink could not be determined. positioning issues: returned product was investigated and there were no abnormalities, particularly with the repositioning unit. the cannula kink was confirmed. since insufficient clinical details were provided, and there are several reported potential causes, the root cause of the positioning issue could not be determined. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7719 | 1220648-2025-30077 | 22412343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412343&pc=OZD | 06/20/2025 | 09/05/2025 |  | 07/07/2025 |  |  | 06/20/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098584 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an aic (console) displayed a controller error during patient preparation. the console was removed from service as a result of the noted failure. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the case started on (b)(6) 2025. the console logs showed an controller error alarm (opm comm crc error ¿ event set # 1045) was triggered during the clinical procedure on (b)(6) 2025. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as - (b)(6) values due to the crc error. device analysis: the console (b)(6) was sent back for further investigation and was tested. the device issues could not be reproduced. no communication issue were reproduced through out normal operation. no misaligned cables associated with the optical bench's rs232 communication. the voltage supplied to the optical bench was monitored and confirmed to be functioning correctly. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue (optical bench fw anomaly). | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
7720 | 1220648-2025-30078 | 22412384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412384&pc=OZD | 06/19/2025 | 08/12/2025 |  | 07/07/2025 |  |  | 06/19/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811664 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) had a burning smell and the power cord was not connecting to the console properly. the console was due fr preventative maintenance. the controller was removed from service. the reported event did not involve patient use. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the impella device was returned for evaluation. there was no issue found during the case. the device functioned/operated to specification.
7721 | 1220648-2025-30079 | 22412796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412796&pc=OZD | 06/21/2025 | 08/19/2025 |  | 07/07/2025 |  |  | 06/21/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618467 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 57 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had oozing at the impella 5.5 insertion site. a suture was placed by the intensivist. the pump was successfully weaned and explanted three days later. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by placing a suture at the bleeding site.
7722 | 1220648-2025-30080 | 22412859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412859&pc=OZD | 06/15/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/15/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667473 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 65 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced arrhythmia. upon explant of the pump a clot was observed. the patient survived. | [5 empty] | A5, ethnicity, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7723 | 1220648-2025-30081 | 22412948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22412948&pc=OZD | 06/20/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/21/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1362195 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that a placement signal unreliable alarm appeared on an automated impella controller during implant. support remained unaffected and the placement was monitored via motor current and imaging. after two days of support, the patient developed multi organ failure and care was withdrawn. the patient subsequently passed away. there is not enough evidence to exclude the failure as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7724 | 1220648-2025-30082 | 22413130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22413130&pc=PYX | 06/19/2025 | 10/07/2025 |  | 07/07/2025 |  |  | 06/19/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026668081 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 95 KG | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had an impella rp flex placed to support the patient awaiting transplant. the pump was supporting for 17 days when the purge pressure rose. the team troubleshot by adding tissue plasminogen activator to the purge and swapping consoles. hemolysis was observed after approximatley a month of support which prompted explant of the pump on day 29. the patient's outcome is stable. | The complainant also reported that a placement signal loss occurred. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of hemolysis and high purge pressure has been completed. the impella device was returned for evaluation. high purge pressure: the returned product had biomaterial beneath impeller blocking the purge gap. the pump was run in the lab after soaking and high purge pressure did not reproduce. the data logs available do not cover complaint occurrence date. logs returned are from (b)(6) while purge system blocked alarms were reported on (b)(6). clinical mentioned console switch out was performed and tissue plasminogen activator (tpa) through purge line performed. issue resolved. the root cause of the high purge pressure was most likely biomaterial as evidenced by the biomaterial on the returned product and the lack of issue reproduction on returned product after soaking. hemolysis: the returned product had biomaterial beneath impeller and around impeller blades. hemolysis was reported on (b)(6). the data logs available do not cover complaint occurrence date. logs returned are from (b)(6) while purge system blocked alarms were reported on (b)(6). the root cause of the hemolysis was most likely biomaterial as evidenced by returned product inspection confirming biomaterial beneath impeller and by impeller blades thrombus: the failure mode was added based on the investigation findings. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d6b explant date is unknown. d9 device returned to manufacturer date added. instructions for use: impella rp with smartassist system. section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ b2 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30082. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30082. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. insufficient data logs provided. the returned product passed the laser continuity test with 5s parameters within range. there was a kink in the catheter before kink protector of blue pump plug. the data logs available do not cover complaint occurrence date. the root cause of the optical signal issue was unable to be determined as insufficient data logs were provided for review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. corrections that were made from the initial manufacturer device report number. d10 concomitant medical products and therapy dates updated.
7725 | 1220648-2025-30083 | 22413302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22413302&pc=OZD | 06/18/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/21/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1362195 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an impella cp was being setup and during setup, there were placement signal unreliable alarms. the pump was disconnected and primed using a backup automated impella controller (aic). the pump was primed and implanted without issues. three hours post implant, they had the same alarm "placement signal unreliable". audio was disabled for this alarm and support was continued. this complaint was created to document the aic exchange due to the placement signal issue. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7726 | 1220648-2025-30084 | 22413499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22413499&pc=OZD | 06/16/2025 | 08/12/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691229 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the cover on the automated impella controller (aic) no longer closes. there was no patient use reported. | [5 empty] | The investigation for the aic - damage before therapy was completed. the console was received for evaluation console was inspected finding that the optical pump connector dust cover was confirmed to be mechanically damaged. no malfunctions were observed, and console operated within specification. | The reported event did not involve patient use. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7727 | 1220648-2025-30085 | 22413780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22413780&pc=OZD | 06/18/2025 | 09/08/2025 |  | 07/07/2025 |  |  | 06/21/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547765 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was being setup and during setup, there were placement signal unreliable alarms. the pump was disconnected and primed using a backup console. the pump was primed and implanted without issues. three hours post implant, they had the same alarm "placement signal unreliable." audio was disabled for this alarm and support was continued. the patient survived. | [5 empty] | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the cause of the optical signal issue could not be definitively determined as no product or logs were returned for analysis and limited clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7728 | 1220648-2025-30086 | 22414012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414012&pc=OZD | 06/20/2025 | 08/16/2025 |  | 07/07/2025 |  |  | 06/20/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620044 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 89 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and suffered an inferior left ventricle wall tear at the infarct site. the tear was repaired surgically and the patient was transfused 12 units of packed red blood cells. the patient remains on impella support for 10 more days. the pump was successfully weaned and explanted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. the data logs show during the time of the complaint the pump was turned to p-0 and there was "impella position unknown" alarms seen. due to lack of details surrounding the depth of the pump when the physician advanced it and the fact that the patient had a small left ventricle (lv) leading to questions regarding the depth of the pump and the physician did not see the impella exit the lv wall, the root cause of the ventricle perforation cannot definitely be determined.
7729 | 1220648-2025-30087 | 22414239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414239&pc=OZD | 06/20/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/20/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659444 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 50 KG |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Asystole | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and had the pump come out of position. the patient went into asystole. cardiopulmonary resuscitation was rendered but the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7730 | 1220648-2025-30088 | 22414483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414483&pc=OZD | 06/15/2025 | 08/12/2025 |  | 07/07/2025 |  |  | 06/15/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664318 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. shortly after the initiation of support, the device was unable to deliver flow rates exceeding 1.4 liters per minute. the pump subsequently stopped functioning. as a result, the device was explanted. during removal, thrombus was observed on the impella cannula, suggesting a potential cause for the impaired flow and pump stoppage. a new impella cp device was implanted. the second device performed without incident, and full circulatory support was successfully re-established. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of low pump flow, pump stop, and thrombus has been completed. the pump was not returned for investigation. the data log shows that the pump was started in auto mode but was unable to achieve expected pump flow. after a few minutes, a pump stop was reported with a motor current spike. the pump was able to restart and ran for another seven minutes before explant. the pump stop failure mode was replaced to temporary pump stop failure mode after investigating the field data log. the cause of the temporary pump stop issue was most likely biomaterial, based on data log review. the cause of the low pump flow issue was most likely biomaterial, based on data log review. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7731 | 1220648-2025-30089 | 22414562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414562&pc=OZD | 06/15/2025 | 10/31/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1662880 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 115 KG |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while supporting a patient implanted with an impella 5.5 the automated impella controller (aic) experienced low purge flow. no patient harm was reported and the patient remains on impella support. | The patient presented with end-stage heart failure requiring inotropic support and admitted with volume overload, and severe lower extremity edema. patient continued to decompensate and the impella 5.5 device was successfully placed for mechanical circulatory support during work up for a left ventricular assist device (lvad) or orthotopic heart transplantation (oht). approximately five days after start of the support the patient experienced a nosebleed. heparin was suspended. mechanical circulatory support (mcs) continued, and heparin was restarted at one point. heparin protocol changed for a higher therapeutic goal. the patient was denied for oht. work up continued and the patient was optimized for a durable lvad. following, nine days later, rhino rocket was replaced to help nostril oozing. impella mcs continued. however, the patient became septic and was placed on antibiotics and more pressor requirements noted on day 20 of mcs. support continued for the patient. approximately 26 days later, placement signal not reliable (psnr) on console was received in addition to the automated impella controller (aic) alarmed low purge flow. the purge flow dropped below three when the patient moved. however, the flow returned to normal. there was no report or indication of aic-to-aic exchange. the issue self resolved and mcs continued. approximately 14 days later the patient expired as care was withdrawn. of note, there were no further updates noted during this 14-day period. |  |  |  |  | The medical safety officer reviewed the case and agreed with the assessment provided in 1220648-2025-30089-1. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue - other and optical signal issue has been completed. the impella device was not returned for evaluation. aic - purge issue - other: based on the data log no issues were observed with the console regarding the high purge pressure. no problem found with the console. optical signal issue: based on the data log no problems were identified with the console in relation to the reported optical signal issue. no problem found with the console. the demise outcome is more related to the patient's condition given the history which includes but not limited to an ejection fraction of 20-25% and exacerbated heart failure. the patient was denied for orthotopic heart transplant showing signs of right ventricular dysfunction (rvd) and plans were noted for possibly cannulation for veno-arterial extracorporeal membrane oxygenation (va ecmo) for additional support while working up for left ventricular assist device (lvad). however, there is not enough evidence to rule out, dissociate the impella 5.5 from the patient's outcome. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. the patient became septic treated with antibiotic and it is unknown how, if any, the impella 5.5 pump played a role. regarding the automated impella controller (aic), patient expired for reasons unrelated to the aic. there was no interruption in support as the aic does not appear to have been exchanged or issue impacted the therapy. aic device remained in service. b5 updated based on the additional information received regarding the patient outcome. the patient death has been captured in manufacturer device report 1220648-2025-29818 b5 updated based on the additional information received regarding the patient outcome.
7732 | 1220648-2025-30090 | 22414586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414586&pc=OZD | 06/23/2025 | 09/09/2025 |  | 07/07/2025 |  |  | 06/23/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1733097 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support and approximately four days into support, the automated impella controller (aic) displayed a controller error. the patient had inadvertently dislocated the pump back into the aorta. upon readvancing the pump into the left ventricle, the pump unexpectedly stopped in therapy. the patient was noted as doing well. the aic gave out a variety of errors including the "controller error." the pump was unable to be restarted. additionally reported, the display glass of the aic was slightly protruding from the case in the upper left corner. the aic was removed from service. the patient was reported to be in stable condition following the event. it was further noted, given the patient's stability, no new impella device implant was planned at such time. | [5 empty] | The investigation of automated impella controller (aic) - damage during therapy and aic - controller error (yellow alarm) has been completed. the device and data logs were received for investigation. data analysis: pump (b)(6) was connected on 6/19 and run without issue until 6/23 when the imc logs show pe_motasym for motor current differences between phases. this is immediately followed by impella stopped alarms as the pump is stopped as a result. the pump attempts to be restarted 13 more times but fails every time. 8 times a controller failure #138 is triggered for a pump motor driver (pmd) speed deviation as well which is a system of the persistent pump stops. the real time logs confirmed the continuous pump stops that occur. device analysis: the failure mode was not reproduced and the console was able to run a pump successfully. aic - damage during therapy: there was no damage found with the aic. the cause of the pump stops could not be determined because the failure mode was not reproduced and the pump was cut. aic - controller error (yellow alarm): the failure mode was not reproduced and the console was able to run a pump successfully. d9 device returned to manufacturer date added. | The automated impella controller (aic) was not received from the customer. therefore, the investigation of the aic was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. refer to manufacturing report number 1220648-2025-30074 for the impella 5.5 medical device report.
7733 | 1220648-2025-30091 | 22414680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414680&pc=OZD | 06/16/2025 | 08/13/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1253203 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was having trouble slaving to the large screen display. there was no patient involvement reported. | [5 empty] | There was no patient involvement in the reported issue. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. the impella device was returned for evaluation. data log analysis ended up not being informative for this complaint. device analysis: the failure mode was reproduced during testing. rl05051 was connected to ic_test and was outputting video as expected to impellaconnect.com but when the console was hooked up to a monitor via vga port on the rlm, the monitor was getting no signal. the monitor was tried with another console's vga on the rlm and was successful so the failure mode was confirmed that the console was not connecting via vga port. root cause: the cause of the user not getting video output was due to video signal not being outputted by the vga port on the rlm.
7734 | 1220648-2025-30092 | 22414771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22414771&pc=OZD | 06/19/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/19/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662668 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 80 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had atrial fibrillation. the patient was cardioverted and support was continued. additionally, the patient had a gastrointestinal bleed. support was continued. the pump was successfully weaned and explanted after 10 days of support. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7735 | 1220648-2025-30093 | 22415371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22415371&pc=OZD | 06/16/2025 | 09/17/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while an automated impella controller (aic) was supporting a implanted impella pump, the purge cassette priming stopped and a purge fluid notification appeared on the screen. a second aic was used to successfully support the patient. no patient harm was reported. | [5 empty] | Patient data was not provided. section a was left blank. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on the data log and device analysis the cause of the purge pressure not changing triggering a purge fluid not detected alarm could not be determined as the failure was not reproduced but may have been influenced by the user.
7736 | 1220648-2025-30094 | 22416111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22416111&pc=OZD | 06/17/2025 | 10/27/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161464 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SERIAL NUMBER (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced failure to boot up of the supporting automated impella controller (aic). a second aic was used to support the patient. no patient harm was reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. the impella device was returned for evaluation. data analysis: console logs from the reported event date aligned with the complaint. an unexpected controller shutdown event #603 alarm were observed just after the console was started, with no controller failure related alarms were found in the logs. the aic was in use during a previous case from june 13th to 15th, 2025, when a shutdown was requested at 10:50:29. although the shutdown was acknowledged and initiated, it was not completed successfully. two days later on the compliant date the aic was turned on and the unexpected controller shutdown event #603 alarm was triggered. the console was rebooted again without any issues and no alarms were triggered. device analysis: the console was tested for multiple power-cycles, but the issue could not be reproduced, and no shutdown related problems or irregularities within the aic were observed during testing. no power supply interruptions were detected between the pbm and the 4-in-1 assembly. aic was run with known good pump and purge cassette without any issue. additionally, the power supply and batteries of the aic were also inspected and no issues were observed. root cause: no issues were found regarding reported failure mode of the aic - power on issue/blank screen with the console. the root cause of the aic's inability to shutdown was determined to be a software anomaly. d6a/d6b should have been left blank on the initial manufacturer device report number 1220648-2025-30094. h1 was erroneously selected on the initial manufacturer device report number 1220648-2025-30094. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-30094. h8 was erroneously selected on the initial manufacturer device report number 1220648-2025-30094.
7737 | 1220648-2025-30095 | 22417111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22417111&pc=OZD | 06/17/2025 | 08/14/2025 |  | 07/07/2025 |  |  | 06/17/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026682607 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | During implantation of an impella 5.5, the guidewire could not cross the aortic valve despite attempting multiple wires. the impella 5.5 implant was aborted and the patient survived. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to resistance and despite using multiple catheters and techniques. the insertion was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7738 | 1220648-2025-30097 | 22418201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22418201&pc=OZD | 06/19/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/19/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640289 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 107 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, there was oozing from the access site and the internal jugular central line. heparin was used prior to the internal jugular central line bleeding. the patient received two units of blood. support was continued. the pump was successfully weaned and explanted five days later. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7739 | 1220648-2025-30096 | 22418237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22418237&pc=OZD | 06/16/2025 | 08/13/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674769 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female | 83 KG |  | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp developed hemolysis. the impella was noted to be shallow and was repositioned to resolve the issue. additionally, the patient experienced arrhythmia. the patient was successfully weaned from the impella and survived. | [5 empty] | The investigation of hemolysis, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. hemolysis: clinical details report that the patient had red-tinged urine the afternoon of (b)(6) 2025. an echo was done and the pump was noted to be shallow with the inlet in the lvot. pump position was corrected. additionally, the patient was experiencing some bleeding from a different access site, so 1 unit of blood was administered. data logs were reviewed and there were no definitive abnormalities noted. there was only 1 position in aorta alarm at the end of the case. this suggests it may have been during pump explant. there was only 1 brief suction alarm that resolved quickly on (b)(6) 2025. there were no abnormal motor current spikes. during the first 24 hours of support, the placement signal (ps) was variable and there was one instance of what appeared to be a drop in motor current (mc)/flows. however, the real time (rt) log showed a different picture: there was a clear increase in ps pulsatility and values that aligned with a drop in mc min values. however, values remained stable and the dropped mc min values were still well within the expected range for p-9. no suction alarms triggered. although there was some variability noted in the ps and mc during the first 24 hours of support that resolved, data logs do not make it clear whether correcting pump position or administering volume for the bleeding event may have resolved the hemolysis. product was not returned for investigation. since insufficient clinical details were provided, data logs were inconclusive, and product was not returned for investigation, the root cause of the hemolysis could not be determined. positioning issues: data logs were reviewed and there was only 1 position in aorta alarm at the end of the case. this suggests it may have been during pump explant and was unrelated to the reported issue. product was not returned for investigation. the root cause of the positioning issue could not be determined due to insufficient clinical information and no product return. vf/vt: the root cause of the reperfusion beats was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30096. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-30096. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30096. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
7740 | 1220648-2025-30099 | 22420283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22420283&pc=OZD | 06/16/2025 | 08/05/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680548 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female | 59 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding at the hemostatic valve. the bleeding could not be controlled with repositioning of the needle or the pump. the pump was explanted and the peel away sheath was exchanged to resolve the issue. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of access site bleeding and introducer damage ¿ mechanical has been completed. the clinical states that " after single access needle stick, large amount of bleeding noted at hemostatic valve. needle removed and bleeding continued. impella adjusted and unable to fix bleeding. impella cp removed and sheath exchanged with short 14 fr (x13cm) peel away sheath in kit." the product was returned and there was a needle puncture seen in the long sheath. leak testing was performed and the introducer was not able to hold pressure as fluid escaped from the needle puncture seen near the introducer hub. based on the clinical details and returned product analysis, the root cause of the introducer damage - mechanical is most likely due to use as a puncture hole was seen where bleeding occurred after single needle stick access. the root cause of the access site bleeding - major is due to use as bleeding was noted due to a puncture hole in the introducer. d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2025-30099. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30099. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7741 | 1220648-2025-30098 | 22420304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22420304&pc=OZD | 06/16/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  | 09/30/2025 |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393711 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Male |  |  | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 was implanted too deep. the team tried to reposition the pump but was unsuccessful as the tuohy borst would not allow pump movement. the patient died from multi organ failure after 12 days of support. | [5 empty] | Athe impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7742 | 1220648-2025-30100 | 22420375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22420375&pc=OZD | 06/16/2025 | 07/07/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2026679757 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced high purge pressure alarms. the alarms were resolved by adding tissue plasminogen activator to the purge. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7743 | 1220648-2025-30101 | 22420867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22420867&pc=OZD | 06/16/2025 | 08/01/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657678 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient implanted with an impella 5.5 was agitated and pulled the impella catheter tubing out of place. the patient experienced ventricular fibrillation and cardiopulmonary resuscitation was provided. after 30 minutes the patient expired. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. vt/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. positioning issues: data logs were reviewed, and there was a concentration of position in ventricle and position in aorta alarms late the evening of (b)(6) 2025. at this time there were distinct changes in motor current (mc) and placement signal (ps) pulsatility aligned with the alarms. the pump had been functioning normally up to that point. returned product was investigated and the catheter anchor still functioned as normal. there were no visual damages/abnormalities noted on the pump besides the sterile sleeve being separated from both the proximal and distal hub. based on the clinical details provided that the patient pulled the catheter out of the left ventricle (lv0, the root cause of the positioning issue was most likely patient movement. product damage - sterile sleeve damage: investigation of returned product found the sterile sleeve to be separated from both hubs. therefore, the produce damage failure mode was added and investigated. clinical details reported that the patient was agitated, pulled on the impella catheter, causing it to come out of the lv. based on the clinical details that the patient pulled on the catheter, the root cause of the product damage (sterile sleeve damage) was most likely a use issue. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30101. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7744 | 1220648-2025-30102 | 22420879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22420879&pc=OZD | 06/15/2025 | 08/05/2025 |  | 07/07/2025 |  |  | 06/16/2025 | 07/07/2025 |  |  |  |  | 07/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660228 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AIC SN (B)(6). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a placement signal not reliable alarm on the supporting automated impella controller (aic). the impella 5.5 was exchanged to resolve the issue. | [5 empty] | The investigation of optical signal issue has been completed. optical signal issue: clinical details state that there was a placement signal not reliable (psnr) alarm during priming of the 5.5. the pump was never inserted into the patient. the pump was replaced. data log review showed no abnormalities with the 5s optical parameters but the ps was out of range right away. the provided logs were cut off and did not show g0 calibration. the pump was returned and evaluated. the pump had no issues with the 5s parameters but it reproduced the psnr alarm and the placement signal (ps) was out of range. the logs from the lab were evaluated and the calibrated g0=0, indicating eeprom corruption. the root cause of the optical signal issue was most likely eeprom corruption since the calibrated g0=0 based on lab data log review but the cause of the corruption is unknown. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7745 | 1220648-2025-30029 | 22393758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22393758&pc=OZD | 06/18/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/18/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654263 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient awaiting heart/kidney transplant was taken to the or for placement of impella 5.5. the 5.5 was successfully placed. it was noted that the patient was tolerating support well waiting as status ii. during support the patient had tooth ache and was in need of an extraction. the patient was changed to status 7 due to infection. in addition, it was noted that the anti-contamination sleeve detached from locking mechanism on distal end. team consulted surgery for replacement with positive blood cultures. heparin was on hold due to removal of teeth. the patient also had a cyst rupture in right lower extremity unrelated to impella. the patient continues on impella support. status ii heart/kidney. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7746 | 1220648-2025-30030 | 22393851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22393851&pc=OZD | 06/18/2025 | 08/14/2025 |  | 07/03/2025 |  |  | 06/18/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026649559 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the patient had oozing from the access site. manual pressure was held, integrillin turned off, and a femstop was applied. support continued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient condition related since the patient was oozing from all sites.
7747 | 1220648-2025-30031 | 22393992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22393992&pc=OZD | 06/14/2025 | 08/13/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664261 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed access site bleeding, vascular injury, and positioning complications, all of which required medical intervention.. the patient had experienced a cardiac arrest and was taken to the catheterization lab for st-elevation myocardial infarction (stemi) intervention, where the impella cp was placed. during subsequent repositioning of the patient, the nurse reported that the vascular sheath became dislodged, resulting in a large hematoma. manual pressure was applied in an attempt to control the bleeding. however, due to the patient¿s body habitus and the size of the hematoma, the sheath could no longer control the vessel bleeding. the patient was taken for surgical repair, blood products were administered, and the impella cp was explanted. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the cause of bleeding issue most likely was patient movement related since the dislodge occurred while attempted to move and repositioning the patient, due to patient body habitus and hematoma the sheath was no longer able to control vessel bleeding impella explanted. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-30031. h6 medical device problem code 1158 was erroneously entered on the initial manufacturer device report number 1220648-2025-30031. | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation)."
7748 | 1220648-2025-30032 | 22394153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394153&pc=OZD | 06/14/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678607 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The patient, with a recent st-elevation myocardial infarction (stemi), underwent percutaneous coronary intervention (pci) with placement of two drug-eluting stents (des) to the proximal left anterior descending (lad) artery. following pci, the decision was made to initiate impella cp support. the peel-away sheath was removed and replaced with a repositioning sheath. forward tension sutures were applied; however, significant oozing and hematoma formation were noted. a retrospective review of angiography revealed a kink in the peel-away sheath. left femoral artery access was obtained for angiography, which demonstrated extravasation in the superficial femoral artery (sfa). percutaneous transluminal angioplasty (pta) was performed, the impella pump was explanted, and an angio-seal device was placed. a follow-up angiogram confirmed resolution of the extravasation. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7749 | 1220648-2025-30033 | 22394180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394180&pc=OZD | 06/14/2025 | 08/08/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665763 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female | 69 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding at the right femoral artery access site, which was managed by adjusting the insertion angle and redressing the site. the patient received two units of red blood cells, and heparin therapy was held. on day five of impella support, the patient developed significant bleeding from the mouth and nose, prompting discontinuation of anticoagulation. the patient also presented with thrombocytopenia. due to a recent effient dose, thrombocytopenia, and ongoing oral bleeding, the surgeon decided to postpone the planned upgrade to the impella 5.5. in the operating room, the patient received two units of packed red blood cells and two units of platelets. the patient was then returned to the icu in stable condition and remains on impella support. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of access site bleeding, vascular damage - peripheral vessel, and thrombocytopenia has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since angle changed and was redressed, which resolved the oozing. the root cause of vascular damage peripheral vessel was most likely patient condition since patient had bleeding from multiple sites other than access site. the root cause of thrombocytopenia cannot be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a.4. weight was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-30033. e.1. fax number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-30033. phone number was revised since the initial manufacturer device report number 1220648-2025-30033 was submitted. h.6. code 4644 was reported incorrectly on initial manufacturer device report number 1220648-2025-30033. a new code was added. h.10. instructions for use were updated since the initial manufacturer device report number 1220648-2025-30033 was submitted.
7750 | 1220648-2025-30034 | 22394314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394314&pc=OZD | 06/13/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/13/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677283 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male | 57 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was placed to support a patient admitted in with cardiogenic shock and other underlying cardiac and systemic comorbidities. the pump was supporting when there was limb ischemia. the ischemia was diagnosed on day four of support. the patient had been transferred on support to a tertiary center but the pump was not escalated at the tertiary center. the pump was explanted and the patient placed on extracorporeal membrane oxygenation. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7751 | 1220648-2025-30035 | 22394491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394491&pc=OZD | 06/18/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/18/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677150 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that roughly one day into support, the inlet of the implanted impella cp pump was found under the papillary muscle. in addition, it was noted that the peel away sheath was left in place following pump implant and distal pulses were subsequently absent. repositioning was unsuccessful and the patient was ultimately escalated to impella 5.5 support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7752 | 1220648-2025-30036 | 22394516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394516&pc=OZD | 06/13/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/13/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630524 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 60 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was placed on impella 5.5 awaiting a heart transplant. during support, it was noted that the patient was having abdominal pain. biliary tract imaging revealed gallbladder inflammation. the patient also tested positive for staphylococcus and antibiotic therapy was initiated. it was also noted that computed tomography (ct) of abdomen/pelvis and a ct chest imaging would be done to further evaluate aortic dissection. the patient was explanted as a bridge to transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings - ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7753 | 1220648-2025-30037 | 22394616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394616&pc=OZD | 06/18/2025 | 07/29/2025 |  | 07/03/2025 |  |  | 06/18/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656992 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp pump placed to support the patient admitted in with cardiogenic shock. the pump supported and a hematoma developed and anticoagulation was held. the pump was explanted after 20 hours of support. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely patient movement since hematoma to right groin occurred upon arrival from another facility.
7754 | 1220648-2025-30038 | 22394798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394798&pc=OZD | 06/19/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/19/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585020 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that following explantation of an impella cp the interventionalist had difficulty closing the groin site. surgical intervention was required and a patch was placed to achieve hemostasis. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7755 | 1220648-2025-30039 | 22394845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394845&pc=OZD | 06/13/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/13/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663421 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Asystole | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp set to be delivered to support a st-elevation myocardial infarction patient. after the intra-aortic balloon pump was removed, the patient did suffer asystole briefly which needed transcutaneous paced. the impella cp was explanted after 64 hours with a purge sidearm leak and was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
7756 | 1220648-2025-30040 | 22394999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22394999&pc=OZD | 06/14/2025 | 09/25/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639638 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 101 KG | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that an impella cp had saturated flow, and the physician decided to explant the impella cp. impella suction thrombus is likely to have been pulled into the descending aorta until the explant. the patient's outcome is stable. | [5 empty] | The investigation of saturated flow has been completed. the pump was not returned for investigation. the data log shows a motor current spike followed by a hike in pump flow above the specification range at steady p-level p7. the motor current spike may refer to the interaction with biomaterial during the case run. the cause of the saturated pump flow issue was most likely biomaterial, based on data log trend. thrombus failure mode was added as investigated failure mode based on the data log review. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The device was received for investigation. the investigation findings confirmed what was reported in previous manufacturer device report 1220648-2025-30040-1. the pump and purge cassette was returned for investigation. no visual damage was observed on the returned product. significant biological material was observed on the returned product. the pump was then tested in a bucket of water, and flow was noted to be saturated. the cause of the saturated pump flow issue was biomaterial, based on return product investigation and data log review. d9 device returned to manufacture date added.
7757 | 1220648-2025-30041 | 22395044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395044&pc=OZD | 06/19/2025 | 08/15/2025 |  | 07/03/2025 |  |  | 06/19/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681668 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced many suction events. an echocardiogram showed the impella placement was good. in addition, there was bleeding from the access site. four units of packed red blood cells were transfused and a femstop was placed. ultimately, the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of low pump flow and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: patient noted to have coagulopathy. the cause of the vascular damage - peripheral vessel was likely patient condition related as patient had coagulation disorder per clinical details. low pump flow: the pump was not returned. data log review showed suction alarms occurring with lower than expected flow rates, at instances of p-7, flows fell below the expected 2.9l/min, as low as 1.9l/min. p-levels were reduced and gradually brought back up. the cause of the low pump flow was most likely patient condition related due to the noted bleeding occurring with suction events that resolved with reducing the p-level and giving volume.
7758 | 1220648-2025-30042 | 22395087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395087&pc=OZD | 06/25/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/26/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666745 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was implanted with a impella cp for support of acute myocardial infarction. it was reported that during cardiopulmonary resuscitation, the patient experienced pulmonary bleeding and hemothorax. the bleeding could not be controlled and had progressed to mediastinal shift. in addition, suction occurred on the impella cp due to insufficient intravascular volume. it was reported that systemic circulation could not be maintained due to the bleeding, and both the extracorporeal membrane oxygenation and impella cp were unable to provide sufficient flow, leading to the patient's death from circulatory failure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7759 | 1220648-2025-30043 | 22395249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395249&pc=OZD | 05/21/2025 | 12/03/2025 |  | 07/03/2025 |  |  | 06/20/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1091080 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that upon turning the pump on to p2, flows displayed 0.0/0.0, and the pump reverted to p0. several attempts were made, however the p-level consistently defaulted back to p0 before the aic turned off. the aic was restarted prompting impella defective" alarm and alert to "replace impella." a new aic was brought in but did not resolve the issue. after several attempts the decision was made to explant the defective pump and implant a new 5.5 impella. this manufacturer report is for 2nd aic (console) with serial number (b)(6). | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation has been completed. the cause of the issue was not determined since issue could not be consistently. reproduced. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
7760 | 1220648-2025-30044 | 22395278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395278&pc=OZD | 06/27/2025 | 09/30/2025 |  | 07/03/2025 |  |  | 06/27/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622108 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 69 KG | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that roughly fifty days into impella 5.5 support, multiple complications were encountered. it was initially documented that the purge pressure began to rise coinciding with a drop in purge flow rates. a purge system blocked alarm was triggered and a blood clot was subsequently discovered around the outflow cage. up until that point in time, the purge fluid was mixed with sodium bicarbonate instead of heparin. it was additionally noted that an intermittent suction alarm began to appear. the pump was removed in response to the failures and the amount of catecholamines were increased to support the patient. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist - section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | High purge pressure: upon evaluating the returned pump, it was observed that high purge pressure reproduced in the lab, and after analyzing the data logs, the root cause of the high purge pressure was determined to be biomaterial buildup inside the purge gap. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. low pump flow: upon evaluating the returned pump, it was observed that low flow reproduced in the lab, and after analyzing the data logs, the root cause of the low pump flow was determined to be biomaterial found in the outflow cage and purge gap.
7761 | 1220648-2025-30045 | 22395411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395411&pc=OZD | 06/01/2021 | 07/03/2025 |  | 07/03/2025 |  |  | 06/01/2021 | 07/03/2025 |  |  | 03/31/2023 |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063657 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support as a bridge to heart/kidney transplant. the patient had a past medical history of heart failure with reduced ejection fraction, with a recent ejection fraction of 50%, due to nonischemic cardiomyopathy, an automatic implantable cardioverter defibrillator inserted in 2017, severe mitral regurgitation, status post mechanical mitral valve replacement and left atrial appendage closure clip. the patient presented to the hospital on (b)(6) 2021 with syncope and was transferred to keck for advance heart failure management and mechanical circulatory support. on the 23rd day of support, the pump abruptly stopped operation and was unable to be restarted. this pump was replaced with a new impella cp. the patient continued on support. | [5 empty] | The investigation for the reported pump stop has been completed. the returned data logs showed the pump stopped immediately following the initial occurrence of the "impella stopped - motor current high" alarm. the alarms persisted and the pump was unable to restart. visual inspection of the returned product revealed a large purge gap. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the pump stop was bearing wear after 23 days of use, which is beyond indication. the ifu states: impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7762 | 1220648-2025-30046 | 22395456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395456&pc=OZD | 06/14/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669226 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 45 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after impella cp placement, the patient developed a cold leg. the patient had posterior tibialis pulse and not a dorsalis pedis pulse. vascular surgery was consulted and the pump was successfully weaned and explanted four days later. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
7763 | 1220648-2025-30047 | 22395716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395716&pc=OZD | 10/27/2024 | 07/03/2025 |  | 07/03/2025 |  |  | 10/27/2024 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519936 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) intermittently made a whirring and grinding noise and that it appeared to be coming from the purge cassette. the noise was not continuous. however, it was found that the device (aic 11387) had fallen and hit the ground while in use. a console change was performed (aic 2760) and after performing the console change, the grinding and whirring sound continued. therefore, a cassette change was then made. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7764 | 1220648-2025-30048 | 22395760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395760&pc=OZD | 06/14/2025 | 08/11/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662612 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Detachment of Device or Device Component | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with a history of cardiomyopathy, mitral regurgitation, and an anomalous left circumflex artery, underwent insertion of an impella 5.5. during support, the patient experienced an episode of ventricular tachycardia. the bedside nurse stated that the event did not appear to be impella-related. however, it remains unclear whether the device contributed to the arrhythmia. subsequently, the patient spontaneously got out of bed, stepped on the white cable, and stood up abruptly, resulting in a complete tear of the driveline at the red plug. the pump stopped immediately. the clinical representative was notified and advised the staff to retract the impella 5.5 across the aortic valve. the damaged device was explanted and replaced with a new impella 5.5. the patient survived the explant procedure and was reported to be in stable condition following the event. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the ventricular tachycardia and pump stop has been completed. the root cause of the vt was unable to be determined due to insufficient clinical details. the clinical details stated that "pt attempted to get out of bed sporadically without help from staff. impella was secured properly in 3 point fixation. however, pt stepped on white cable and stood up abruptly. this caused a total tear in driveline at red impella plug. impella stopped immediately." there were no data logs returned. the product was returned and the impella catheter with the kink protector was pulled out of the red handle and all lines were broken. there was glue seen around the kink protector. based on the clinical details and returned product analysis, the root cause of the pump stop was due to open electrical lines.
7765 | 1220648-2025-30049 | 22395804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395804&pc=OZD | 06/14/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/14/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026673710 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an attempted insertion of an impella 5.5 was aborted due to challenging vascular anatomy. the patient was transitioned to an impella cp for circulatory support. the support continued as a bridge and the patient remained stable throughout this period. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7766 | 1220648-2025-30050 | 22395988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22395988&pc=OZD | 04/09/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 04/09/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636281 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Atrial Fibrillation | Ischemia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility an impella 5.5 placement for a patient for support during high-risk cardiac procedure. it was noted of successful placement of impella 5.5 via direct approach using tee guidance. during support it was noted that the patient was resting on impella support at p5. however, had a few runs of ventricular tachycardia (vt), the pump was pulled back 1cm and no further vtach was noted. it was noted that the patient had ectopy at baseline. in addition, the patient went into atrial fibrillation, amio was started. the 5.5 was successfully weaned and explanted. however, it was reported that despite explant being uneventful. subsequently the patient was discharged and has been outpatient. however, presented back with complaints of leg pain, found to have an ischemic leg. workup included a ct scan that showed a thrombus in the ascending aorta at the level of the impella graft remnant. no extravasation of contrast noted in graft. patient required percutaneous thrombectomy with resolution of ischemia. notable, patient was on eliquis at home. plan to anti-coagulate with coumadin, no surgical intervention planned. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7767 | 1220648-2025-30051 | 22396013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22396013&pc=OZD | 04/22/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 04/22/2025 | 07/03/2025 |  |  | 06/30/2025 |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360230 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured vascular complication - embolism of left external iliac with a "definite" relationship to the impella device used: ¿after impella removal, clot burden identified within left external iliac, common femoral artery (cfa), and superficial femoral artery (sfa). after a fogarty balloon embolectomy. left external iliac dissection after stenting. bovine patch repair of left cfa arteriotomy. primary repair of sfa arteriotomy¿. the patient recovered with no sequelae. additionally, it was noted in the patient updates that the pump was repositioned at bedside for red urine. on echo and chest xray. pump was shallow on both. the pump was advanced and returned to p6. in addition, there was placement signal alarm noted periodically. pump was measuring 5.35 and was again pulled back 1.5 cm. patient declined escalation to 5.5. so, cp was weaned and removed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7768 | 1220648-2025-30052 | 22396180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22396180&pc=OZD | 06/15/2025 | 07/31/2025 |  | 07/03/2025 |  |  | 06/15/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655873 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 89 KG | Non Hispanic | White | Death | Yes | Yes | Crack | Restricted Flow rate | Malposition of Device |  |  |  |  | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during the implant, the peel away sheath cracked. the patient was bleeding. the peel away sheath was removed and the repositioning sheath was implanted. the bleeding stopped. additionally, the patient was found to have hemolysis. the pump position was shallow but the doctor did not want to reposition so support was continued. the patient also had limb ischemia. a femoral-femoral bypass was placed. furthermore, the pump was going to be explanted and replaced with an impella 5.5 device. upon arrival to the intensive care unit (icu), the patient had a hematoma to the right groin. shortly after arrival, the patient began to decompensate significantly requiring rapid transfusion and massive transfusion protocol. at this time, it was noted patient had a growing hematoma to the left groin. an emergent surgical cut down was performed in the icu to control the bleeding. hemostasis was achieved on the left groin. the doctor surgically repaired the right groin at this time as well. the patient maintained flows on impella cp with minimal suction alarms during resuscitation. the doctor stated this adverse event was due to hemostasis devices used to close the external femoral-femoral bypass and closure of the impella cp insertion site. the patient expired. | [5 empty] | The investigation of positioning issues, introducer damage ¿ mechanical, access site bleeding, vascular damage - peripheral vessel, hemolysis, and limb ischemia has been completed. the impella device was not returned for evaluation. positioning issues: clinical details report that the patient began to experience hemolysis symptoms the morning of (b)(6) 2025. at this time, position was checked with imaging, and it was noted that the pump was shallow. the physician decided not to reposition. of note, the patient was reported to have a small left ventricle (lv). no clinical details were provided to suggest how the pump had become shallow. data logs were reviewed and there was only 1 position in aorta alarm at the end of the case. this was presumably during pump explant. product was not returned for investigation. since insufficient clinical details were provided regarding how the pump came out of position, the root cause of the positioning issues could not be determined. introducer damage - mechanical: clinical details report that during the insertion procedure, the introducer sheath cracked. there were no vessel conditions reported nor excessive force used. of note, however, it was reported that the insertion angle was 45 degrees. the introducer sheath was pulled out of the body and peeled away, then the repositioning sheath was advanced to resolve the bleeding. product was not returned for investigation. based on the clinical detail provided that the insertion angle was steep, the root cause of the introducer damage was determined to most likely be a use issue. access site bleed: clinical details report that during access, the introducer sheath was found to be cracked. the access site was bleeding. the peel away sheath was removed, the repositioning unit was advanced, and the bleeding subsided. it was reported that the introducer was inserted at a 45 degree angle. product was not returned for investigation. clinical details provided that the introducer sheath was inserted at a 45 degree angle and cracked. the bleeding resolved after advancing the repositioning unit. based on this information, the root cause of the access site bleed was determined to most likely be a use issue. vascular damage - peripheral vessel: clinical details report that shortly after the impella cp was explanted and the patient was escalated to an impella 5.5, the cp access site and the left femoral artery (lfa) (other site related to fem-fem bypass) developed hematomas. additionally, both vessels required surgical repair. the physician believed that the adverse event was due to the hemostasis devices used to close the external fem-fem bypass and cp insertion sites. however, no further details were provided, particularly if excessive force was used or if the patient's condition may have played a role. since insufficient clinical details were provided, the root cause of the vascular damage could not be determined. hemolysis: data log review did not confirm any positioning issues. there was 1 position in aorta alarm at the end of the case, presumably during pump explant. there were no abnormal motor current (mc) spikes during the case. that being said, the placement signal (ps) and mc showed variability throughout the case. there was one instance early in the case where there was a significant increase in the ps/map, and it correlated with a drop in mc minimum values. that being said, mc maximum values remained stable and no suction alarms triggered. this was an acute event that resolved quickly. later in the evening of 15/jun/2025, there was a trend down in ps that aligned with a trend down in mc. the mc recovered, but ps continued to be variable. therefore, it is unclear if the patient's condition could have caused this drop in mean flows. product was not returned for investigation. since clinical details reported several potential factors that could play a role in hemolysis, data log review was inconclusive, and product was not returned for investigation, the root cause of the hemolysis could not be determined. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. g1 reporting contact email revised on manufacturer device report 1220648-2025-30052 in accordance with updated procedures. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30052. | Abiomed's medical safety officer reviewed the event and determined the patient was escalated to and expired on the impella 5.5 but was a serious injury for the atrial fibrillation. the impella cp was related adverse events (e.g., limb ischemia, access bleeding requiring blood transfusion, etc.) are more likely associated with the demise outcome. there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome. this is one of two medical device reports associated with the patient death. refer to initial medical device report for the impella 5.5 serial number (b)(6) with date of event 17-jun-2025 and identifier ending in (b)(6). the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7769 | 1220648-2025-30053 | 22396221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22396221&pc=OZD | 05/05/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 05/05/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599668 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient is waiting for a heart transplant with impella 5.5 support. after a month of support the patient experienced ectopy only when ambulating and leaning forward. echo was recommended. soon after the team reported having suction and could not increase the p-level above 4 when the patient was resting, but could increase p-level to 7 with ambulation. the team was advised to follow ifu guidelines for preload and monitoring position with echo. three months later, there was a placement signal not reliable alarm. there was no change in the patient hemodynamics or motor current. the impella position was good. audio was disable. | [5 empty] | The investigation for the arrhythmia and low pump flow has been completed. no product has been returned. the data logs do not show any motor current spikes outside p-level changes. clinical reported having suction and cannot increase p-level above 4 when resting but can increase p-level to 7 with ambulation. flows were within expected range for matching p-level. data logs confirm continuous suction alarms while at p4. correct position was verified via imaging. no details on any patient anatomy complications or volume issues. suction alarms continue with p-level between p4 and p7 for remainder of support with appropriate flows for matching p-level. the root cause of the low pump flow was most likely due to patient condition as evidenced by correct flows for p-level and continuous suction alarms with patient tolerating p7 during ambulation.
7770 | 1220648-2025-30054 | 22396541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22396541&pc=OZD | 06/06/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/06/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670781 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient developed a bleed from their femoral artery during impella cp support. heparin was held and the pump was subsequently removed. surgical intervention was required to manage the bleed and blood products were given to counteract the blood loss. the bleed contributed to a complicated pump removal and it was believed it resulted in ischemia of the limb. the patient remained supported via bipella support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7771 | 1220648-2025-30055 | 22396697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22396697&pc=OZD | 06/12/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/12/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620825 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 on a patient as bridge to advanced therapies, left ventricular assist device or orthopedic heart transplant (oht). during 5.5 support the patient had increased ectopy with movement. echo position was slightly deep 6.5cm, however, there was no decision to reposition. the patient was on p7, flows 4. left radial arterial ultrasound was performed to assess for pseudoaneurysm after bleeding from 7-day old arterial line site. echo again showed impella measuring around 6.5cm from av annulus, pulled back to 5.7cm. in addition there was bleeding from axillary impella site with extravasation around the graft. the patient continues on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7772 | 1220648-2025-30056 | 22396851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22396851&pc=OZD | 08/21/2022 | 07/03/2025 |  | 07/03/2025 |  |  | 08/21/2022 | 07/03/2025 |  |  | 02/29/2024 |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023191125 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was being implanted with an impella 5.5 for mechanical circulatory support. the patient's vessel diameter was evaluated and an attempt was made to insert the pump through the anastomosis graft on the right subclavian region. it was reported that the impella would not advance over the 0.18 guidewire. several attempts were made, all were unsuccessful. the impella 5.5 insertion was aborted and an impella cp was delivered instead. there was no patient injury. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the delivery issue was difficult patient anatomy. patient vessel size was noted to be <7mm. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
7773 | 1220648-2025-30057 | 22398081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22398081&pc=OZD | 06/11/2025 | 08/16/2025 |  | 07/03/2025 |  |  | 06/20/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663002 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed in place of the intra-aortic balloon pump. during use of the aic (console) it was reported that the purge kept priming and could not get to stop, so the aic was exchanged and that resolved the issue. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on the data and device analysis the cause of the priming failure was most likely the purge cassette not properly seated. this then triggered an erroneous piston block in the logs which triggered the pud to reset and interrupt priming. d6a and d6b should have been left blank on the initial manufacturer device report number 1220648-2025-30057.
7774 | 1220648-2025-30058 | 22398301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22398301&pc=PYX | 06/19/2025 | 08/12/2025 |  | 07/03/2025 |  |  | 06/20/2025 | 07/03/2025 |  |  |  |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1411424 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a rp flex placed to support the patient awaiting transplant. the pump was supporting for 17 days when purge pressure rose. the team troubleshot by adding tissue plasminogen activator (tpa) to the purge and swapping consoles for the purge disc failure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ | The investigation of automated impella controller (aic) - a/c power issues has been completed. the impella device was returned for evaluation. data review: console logs were partially consistent with what was reported in the complaint. real time (rt) logs show that at the time of the purge pressure high alarm, the purge pressure was steady at around 500mmhg and then gradually spiked up to about 1000mmhg. device analysis: console was tested after arriving in field service. the reported purge system blocked issue was unable to be reproduced while testing on an engineering lab bench. further inspection of the console revealed that there was a broken purge disc flag. conclusion: the root cause of the purge system blocked alarm was likely use related. it is likely that there was a kink in the purge line after the purge disc causing any pressure to be trapped in the purge disc while the purge piston was advancing. the observed broken purge disc flag reinforces this, as the trapped pressure in the purge disc likely caused the flag to break. d9 device returned to manufacturer date added.
7775 | 1220648-2025-30059 | 22398451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22398451&pc=OZD | 06/19/2025 | 07/03/2025 |  | 07/03/2025 |  |  | 06/20/2025 | 07/03/2025 |  |  | 06/30/2025 |  | 07/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-CA |  |  | OZD |  |  | IMPELLA CP -CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2024378550 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient lost pedal pulse in their impella leg during support. in addition, the patient experienced a gastrointestinal bleed for which two units of blood were given. the decision was ultimately made to explant the pump in response to the conditions. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7776 | 1220648-2025-29987 | 22381634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22381634&pc=PYX | 06/13/2025 | 08/28/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540911 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 89 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella rp flex, the patient experienced blood leaking from where the impella catheter enters the repositioning sheath. when the device was pulled back slightly, the leaking would cease. the patient was successfully weaned from the device and survived. | [5 empty] | The investigation of access site bleeding has been completed. the pump was returned for investigation. data log review was not required for this case. according to the clinical report, blood was leaking from the repositioning sheath entry point, which ceased after a slight adjustment in catheter position. the returned product showed no visual damage to the repositioning sheath. leak testing of the repositioning sheath revealed a leak from the hemostasis valve at 60 mmhg. further close-up imaging confirmed the valve was damaged. the cause of the access site bleeding issue was most likely damaged valve. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿
7777 | 1220648-2025-29988 | 22381734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22381734&pc=OZD | 06/13/2025 | 08/27/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665167 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 to support a patient admitted in for coroanry artery bypass graft and found to have postcardiotomy cardiogenic shock. the pump was supporting despite known heparin-induced thrombocytopenia (hit) positive status when after three days the highpurge pressure was treated with the addition of lysis, despite the hit status. later, care was withdrawn and the patient expired. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | High purge pressure: clinical details state that there was high purge pressure (pp) and purge flow (pf) dropped from 11.4 to 3.1 ml/hr on (b)(6) 2025. pp/pf was in range again by the next day after using lysis. data logs are not available. the product was returned and evaluated. there was biomaterial in the outlet cage partially covering the purge gap and biomaterial around the impeller. no other abnormalities were found. high purge pressure did not reproduce while the pump was run in the lab. the pp was stable around 500 mmhg during the run. without field data logs, the hpp cannot be evaluated for motor current spikes/trends or other patterns. the cause of the high purge pressure was not determined since data logs were not available. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29988.
7778 | 1220648-2025-29989 | 22381854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22381854&pc=OZD | 06/18/2025 | 08/14/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666181 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 79 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Sepsis | Tachycardia | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and after the implant, there was oozing from the access site. forward tension sutures were placed, angle matched, and dressing was changed. support was continued. additionally, the patient had hemolysis during support and continuous renal replacement therapy was started. the p level was decreased and heparin concentration was increased. the staff did not think the hemolysis was caused by the impella. however, it is unknown if the impella contributed to the hemolysis. furthermore, the patient had ventricular tachycardia during support. amiodarone was started. moreover, the patient was possibly septic and had a cold leg. the staff planned on weaning the pump and explanting however explanation was not confirmed. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding, renal failure, vf/vt/af/at, limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. access site bleeding - major: data logs are not applicable and product was not returned. the root cause of the access site bleeding was most likely use-related since the bleeding resolved after placing forward tension sutures and angle matching. hemolysis: clinical details note on the second day of support that the ldh increased but team did not think cp was causing hemolysis at this time. it is unknown when the hemolysis began and if/when it resolved. data log review did not show positioning issues, motor current spikes, suction, or placement signal trends. no product was returned. the root cause of the hemolysis was not determined since insufficient clinical details were provided and no product was returned. renal failure: the cause of the renal failure was not determined due to insufficient clinical details. limb ischemia - reversible: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details.
7779 | 1220648-2025-29990 | 22381910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22381910&pc=OZD | 06/18/2025 | 07/29/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680634 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma in the right groin. four units of packed red blood cells were transfused to the patient. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely patient movement since the patient sat up in their icu bed and bent their right leg, which was not immobilized, which led to a growing hematoma of right groin.
7780 | 1220648-2025-29991 | 22381932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22381932&pc=PYX | 06/18/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  | 02/28/2026 |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529155 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 86 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient went into atrial fibrillation and then into ventricular fibrillation during impella rp flex support. advanced cardiac life support was initiated in response to the condition and the patient was cannulated for veno-arterial extracorporeal membrane oxygenation. support continued with the implanted pump following the intervention. the patient survived and was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7781 | 1220648-2025-29992 | 22381984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22381984&pc=OZD | 06/13/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668961 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 66 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced intermittent ventricular tachycardia upon movement by the patient. proper pump positioning was confirmed by echocardiogram. the patient continued on support and was eventually successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7782 | 1220648-2025-29993 | 22382030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382030&pc=PYX | 06/18/2025 | 08/05/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026652205 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp flex experienced biomaterial in the pump and low flow output. the pump was explanted. no patient harm was reported. | [5 empty] | The investigation of low pump flow and thrombus has been completed. low pump flow: the returned pump was inspected and there were no physical abnormalities. the pump had biomaterial wrapped beneath the impeller, blocking the purge gap and some around impeller blades. the logs show a sudden loss of flow calculation and spike in placement signals with no change in motor current or cvp readings. clinical mentioned this occurred after thoracentesis procedure. a swan was also placed after rp flex implant. issue may be as a result of possible device interaction with the differential pressure sensor. it is unknown if any lines or devices were present at the the time of issue or how close they were to pump sensors. the root cause of the low pump flow was not determined as limited clinical information was provided to confirm possible root cause thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29993. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7783 | 1220648-2025-29994 | 22382103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382103&pc=OZD | 06/18/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613727 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 experienced deep vein thrombosis on the right arm and possibly on the left. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7784 | 1220648-2025-29995 | 22382144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382144&pc=OZD | 06/13/2025 | 08/12/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675659 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 113 KG | Non Hispanic |  | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted as it would not advance across the aortic valve due to the patient's anatomy. due to the forceful attempts, the cannula kinked and the patient experienced access site bleeding. the patients outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of vascular damage - great vessel, delivery issues, and product damage (cannula kink) has been completed. product damage (catheter kink) is added based on return product investigation. vascular damage - great vessel: clinical reports the impella was unable to advance across the aortic valve due to tortuous anatomy and/or anastomosis angle. the pump was returned for investigation and kinks were observed on the cannula and catheter near the motor housing. the cause of the vascular damage was most likely use related, as clinical reports excessive force was used during insertion attempt. delivery issues: clinical reports the impella was unable to advance across the aortic valve due to tortuous anatomy and/or anastomosis angle. the pump was returned for investigation and kinks were observed on the cannula and catheter near the motor housing. the cause of the delivery issues was most likely due to the patients condition, as clinical details report the pump was unable to pass through the aortic valve due to tortuous vascular anatomy. product damage (cannula kink): clinical reports the impella was unable to advance across the aortic valve due to vascular anatomy and/or anastomosis angle and that multiple forceful attempts were made resulting in a cannula kink. the pump was returned for investigation and kinks were observed on the cannula near the motor housing. the cause of the cannula kink was most likely due to the patients condition, as clinical details report the pump was unable to pass through the aortic valve due to vascular anatomy. product damage (catheter kink): clinical reports the impella was unable to advance across the aortic valve due to vascular anatomy and/or anastomosis angle and that multiple forceful attempts were made resulting in a cannula kink. the pump was returned for investigation and kinks were observed on the catheter near the motor housing. the cause of the catheter kink was most likely due to the patients condition, as clinical details report the pump was unable to pass through the aortic valve due to vascular anatomy. d9 device returned to manufacturer date added.
7785 | 1220648-2025-29996 | 22382165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382165&pc=OZD | 06/18/2025 | 08/15/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610812 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 107 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced an "erratic" placement signal. the patient remains on impella support and no patient harm was reported. | [5 empty] | The pump was not returned for investigation. no known pump optical failure patterns were confirmed from the data log review. the cause of the optical signal issue was not determined since the product was not returned for investigation. corrected information provided in h10. | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7786 | 1220648-2025-29997 | 22382260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382260&pc=OZD | 06/18/2025 | 08/05/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610815 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed hemolysis. the patient remains on impella support. no patient harm was reported. | It was further reported that the patient had slight hematoma growing at the cutdown site. it was reported that a washout would need to be done if it grew significantly. systemic heparin was being used. patient update reported that the hematoma at the access site had gone down some, so surgical washout was not needed. systemic heparin no longer being used. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis and hematoma has been completed. the impella device was not returned for evaluation. since clinical details do not report any troubleshooting methods that resolved the hemolysis and data logs were not conclusive, the root cause of the hemolysis could not be determined. since it is unclear when/if the hematoma resolved and what management methods were used, the root cause of the access site bleed could not be determined. b5 updated with additional information. d6b explant date was added. h6 health effect - clinical code 1884 added b2 intervention was erroneously selected on the initial manufacturer device report number 1220648-2025-29997. h6 health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2025-29997.
7787 | 1220648-2025-29998 | 22382291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382291&pc=OZD | 06/13/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594015 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male | 88 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 pump was supporting a patient awaiting transplant. while awaiting the heart transplant, computed tomography of the head showed new right striatocapsular lacunar infarct and perfusion deficit in the right posterior cerebral artery. the patient complained of small left sided weakness but no noted significant neurology deficits per the nurse. the patient still awaits transplant while on the impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7788 | 1220648-2025-29999 | 22382423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382423&pc=OZD | 06/18/2025 | 08/14/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026684327 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 65 KG | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted. due to the patient's anatomy the catheter kinked when advancing the pump. a new impella 5.5 was implanted successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue : the root cause of the delivery issue was most likely patient condition related, the team attempted to deliver the 5.5 but the delivery failed with kinking when trying to advance due to patients anatomy. product damage ( cannula kink) : the root cause of the product damage issue was most likely patient condition related, the team attempted to deliver the 5.5 but the delivery failed with kinking when trying to advance due to patients anatomy.
7789 | 1220648-2025-30000 | 22382442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382442&pc=OZD | 06/18/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611061 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female | 88 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and after the implant, a left ventricle perforation was noted. the pump was explanted. the patient was stable after the explant. however, the patient ended up expiring. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7790 | 1220648-2025-30001 | 22382489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382489&pc=OZD | 06/16/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/16/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673759 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient received tissue plasminogen activator (tpa) for a suspected right coronary artery thrombus. distal pulses past the 14 french peel-away sheath were unable to dopplered in due to the patient receiving tpa. the physician injected contrast and placed a distal perfusion catheter for sluggish flow and to prevent limb ischemia. the patient was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7791 | 1220648-2025-30002 | 22382630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382630&pc=OZD | 06/18/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674567 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 96 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support had a computed tomography (ct) of the brain that showed multiple cerebrovascular accidents. ct was done again showing evolving strokes but not worsening. neurologist concluded that the strokes were not the culprit for the patient absence of neurology status. additionally, the patient experienced run of ventricular tachycardia requiring one round of cardioversion. the patient went into ventricular fibrillation arrest and the decision was made to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7792 | 1220648-2025-30003 | 22382711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382711&pc=OZD | 06/16/2025 | 07/29/2025 |  | 07/02/2025 |  |  | 06/16/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658240 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following insertion, a kink was noted in the cannula after turning around the aortic arch and into the descending aorta. the device was unable to flow more than 0.9 liters per minute. the device could not be repositioned without exiting the left ventricle. therefore, the impella was removed and replaced. a new impella cp was successfully implanted. the patient was noted to be stable post procedure. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. product damage (cannula kink): clinical reported patient had small aortic arch and small lv and that kink in the cannula occurred after making the turn around the arch into the descending. the root cause of the product damage (cannula kink) was most likely patient condition related as evidenced by clinical details of small patient anatomy. low pump flow: clinical reported patient had small aortic arch and small lv and that kink in the cannula occurred after making the turn around the arch into the descending. upon activation, there was low pump flow. the logs do not show any motor current spikes. the root cause of the low pump flow was most likely due to a cannula kink as evidenced by clinical details of small patient anatomy causing cannula kink and low flow upon activation.
7793 | 1220648-2025-30004 | 22382802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382802&pc=OZD | 06/27/2025 | 09/24/2025 |  | 07/02/2025 |  |  | 06/27/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026685950 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 pump when the medical doctor wanted to escalate care to another impella 5.5. however, was unable to advance through the ltsca graft anastomosis during attempted delivery. the insertion was discontinued as a result of the noted resistance. support was maintained with the original 5.5 device. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because failure to pass impella through due to patient vascular anatomy.
7794 | 1220648-2025-30005 | 22382875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382875&pc=OZD | 06/18/2025 | 07/29/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678526 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 91 KG |  | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Adhesion(s) | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was placed for a patient admitted in with coronary artery disease, cardiomyopathy, and plan for coronary artery bypass graft. the patient was supported for seven days when explanted. the surgeon planned to attempt to remove the pump through the graft but had the team aware of potentially opening the chest. the team was ready and opened the chest with no complications. the two cardiac surgeons, who were on the case, said they sutured the anastamosis very tight which required resutures and surgicel at the initial implant. they were aware of the potential challenge when explanting the patient. the patient's outcome is unknown. | [5 empty] | The investigation of removal issues has been completed. no product was returned for analysis. data log review not required. clinical reported unable to remove pump past the anastomosis. chest opened and removed through the graft with ease. the root cause of the removal issue was most likely patient condition related as evidenced by clinical detail of being unable to remove pump past the anastomosis. b2 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30005. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7795 | 1220648-2025-30006 | 22382984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22382984&pc=OZD | 06/16/2025 | 08/01/2025 |  | 07/02/2025 |  |  | 06/16/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645815 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp started once the guidewire was out and the impella cp was seen to be in good position. once the impella cp started, the placement signal was not reading appropriately. although the pump was flowing at 3.6 liters per minute and there was a pulsatile motor current, the placement signal was not reliable. the placement signal was not registering appropriately and the staff decided to switch the impella cp to new automated impella controller. once that was done, the placement signal normalized. it was noted that this event did not result in any harm to patient. the patient remained hemodynamically stable throughout the event. additionally, the patient converted from normal sinus rhythm to rapid atrial fibrillation. the patient was put on amiodarone. the patient had atrial fibrillation with a ventricular response rate between 80 to 120 beats per minute. the patient was occasionally paced. the impella cp was repositioned under echocardiogram from 93 centimeters to 91 centimeters. it was noted that the right ventricle's function was depressed due to the arrhythmia but it believed to be inaccurate due to the rapid atrial fibrillation (as the echocardiographic measurement decreased from 14 to 8). the pulmonary artery pulsatility index was around 0.8 at that time. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation of optical signal issue and vf/vt/af/at has been completed. the impella device was not returned for evaluation. optical signal issue: data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed on console within 3 minutes of starting the pump. as soon as the pump was connected, the snr was roughly 300. around the time the pump started, it trended down to ~78, and the psnr alarm triggered 2 minutes later. contrast also trended down along with snr. when the pump was connected to a second console, all optical signal metrics were generally stable for the duration of support. this suggests that the pump was not the cause of the optical signal issue. vt/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. | The device was not received from the customer, therefore no evaluation is possible. if the device or additional information is received, a supplemental mdr will be filed.
7796 | 1220648-2025-30007 | 22383091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22383091&pc=OZD | 06/18/2025 | 08/05/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674538 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that they encountered removal issues. during removal of the impella, closure attempt with angioseal was unsuccessfully completed. perclose was attempted but groin was still unstable. manual pressure for 30 mins was held and then femstop applied. pulses in foot were present. patient was weaned and explanted. | [5 empty] | The investigation of access site bleeding and removal issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b1 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30007. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30007. h1 was erroneously selected on the initial manufacturer device report number 1220648-2025-30007. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-30007. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30007. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
7797 | 1220648-2025-30008 | 22383147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22383147&pc=OZD | 06/18/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674683 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed to support an cardiogenic shock patient. the pump was placed but there was bleeding at the access site. the pump was explanted and escalated to the 5.5 pump after just over 24 hours. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7798 | 1220648-2025-30009 | 22383187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22383187&pc=OZD | 06/19/2025 | 08/13/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672996 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support and developed pocket hematoma. it is unclear when the hematoma occurred but was noticed by the nurse two days after implant. the hematoma did not appear to be worsening. anticoagulation was held. the patient was successfully weaned from the pump and survived. | A notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry reported upon 2 ae's for the patient. ae #3 for right axillary hematoma impella site. the relationship is listed as "possibly" related. "loqi states: cta was negative for extravasation but showed hematoma. 1-unit prbc (b)(6) 2025 and (b)(6) 2025. 1 unit of ffp 6/18. h/h was 7.8/23.1 ((b)(6) 2025 1032). baseline unknown. impella was removed. pt had some oozing post op. 2 additional units of prbc given and 1 unit of plts also given. jp left in place until (b)(6) 2025. monitored cbc w/ platelets." loqi also reported upon ae #4 of"hemolysis with relationship noted to be "possibly" related. loqi states: peak k+ is 6.4 on (b)(6) 2025 at 1533); peak ldh is 290 ((b)(6) 2025 0203). peak for plasma free hgb was 7.4 ((b)(6) 2025 at 0452); total bil 0.5 peak on (b)(6) 2025 at 0452). pt was hyperkalemic last night. pd restarted. hyperkalemia may be d/t impella related hemolysis and off pd for a period of time yesterday. lokelma given. paitromer restarted. labs monitored." |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b5 and h6 updated with additional information received from the clinical site.
7799 | 1220648-2025-30010 | 22383363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22383363&pc=OZD | 04/24/2025 | 09/05/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249361 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient that was a bridge to left ventricular assist device (lvad) has a pump exchanged due to persistent positioning alarms. it was reported that during support on the second impella 5.5 that there was an impella stopped alarm. the automated impella controller (aic) was exchanged without success. the cannula remained across aortic valve, however repeated attempts to restart were unsuccessful. the patient remained stable hemodynamically and marinating at baseline throughout the pump stop period. patient was taken to operating room for impella explant and lvad. at the customer request to also to send aic (b)(6) back to get it looked at in addition to the pump investigation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - pump stop has been completed. the impella device was not returned for evaluation. analysis: 5.5 pump was connected to ic4457 on 4/24 and ran until 5/18 with no pump stops. on 5/18 the pump was moved to ic10758 and there were multiple pump stops as the motor current spiked to try and fail to restart the pump. root cause: the console most likely performed as expected and the pump was most likely damaged by the user causing the pump stops see manufacturer device report 1220648-2025-29041.
7800 | 1220648-2025-30011 | 22383401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22383401&pc=OZD | 06/18/2025 | 08/14/2025 |  | 07/02/2025 |  |  | 06/18/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665411 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 73 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was placed for support of a patient admitted in with cardiac and systemic comorbidities. the pump was supporting when purge pressures rose and purge blocked alarms were present. the team troubleshooted but no pump explant was planned. the pump remained on for use despite the malfunction. the patient's outcome is stable. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was returned for evaluation. the pump was returned for investigation. a kink in the catheter was observed at the kink protector. some biomaterial was noted at the outlet cage and was easily removed with a dental pick. the purge gap was found to be free of biomaterial, and further priming testing revealed no blockage in the purge line, with the pump operating as expected in a bucket of water. the data log shows a 1-hour gradual rise in purge pressure followed by a high purge pressure (hpp) alarm for about a day, after which the high purge pressure issue was resolved with a gradual decrease trend. the pump was utilized for another 3 days without issue. the cause of the high purge pressure issue was biomaterial buildup, based on data log review and returned product investigation. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7801 | 1220648-2025-30012 | 22384045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22384045&pc=OZD | 06/12/2025 | 07/26/2025 |  | 07/02/2025 |  |  | 06/12/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615239 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 101 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Defective Device | [5 empty] | Pericardial Effusion | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient presented with continued chest pain. the patient was taken urgently to the catheterization lab. the medical doctor believed the left anterior descending artery (lad) to be the culprit lesion and the patient received two drug-eluting stents to lad without complications. the patient was moved over to the stretcher where the patient went into ventricular fibrillation (vf) and arrested. the patient was shocked numerous times and placed back on the catheter table where the medical doctor placed an intra-aortic balloon pump (iabp). the patient was intubated and continued to have vf requiring defibrillation and cardiopulmonary resuscitation. the decision was made to remove the iabp and place an impella cp via the right femoral artery. the patient was successfully placed on impella support and vf ceased. however, it was noted that during impella cp access, the medical did not unlock introducer from sheath and instead pulled introducer out of peel away sheath, breaking introducer connection, and leaving orange screw on portion on peel away sheath. once the peel away sheath was backed out, the medical doctor was able to break the peel away and connection to introducer. the patient began needing to be externally paced due to heart rate of 20. the medical doctor came back to table and placed a transvenous pacemaker. the patient was taken to operating room for upgrade to an impella 5.5. after multiple attempts to advance the impella 5.5, the medical doctor was unsuccessful. the decision was made to explant the impella 5.5 and use graft for extracorporeal membrane oxygenation (ecmo) cannulation. the patient was successfully cannulated for ecmo. after ecmo was initiated, the impella and ecmo were unable to flow. it was determined on transesophageal echocardiogram that the patient had a large pericardial effusion. the medical doctor opened the chest and drained the effusion and multiple blood products were given. the impella was repositioned, but ecmo and the impella remained unable to flow. after long continued effort, the team was unable to flow. the decision was made to stop efforts and time of death was called. | Abiomed's medical safety officer reviewed and determined there was an issue with the peel away sheath. however, there was no impact to support or injury noted. the patient was supported for approximately seven hours and needed more support and was escalated to an impella 5.5 (delivery issues with failed attempt and adverse event) and subsequently expired. there is not enough evidence to exclude the impella 5.5 device used as an associated factor in the patient's outcome, which is captured under manufacture device report 1220648-2025-30015. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B5 and h10 medical safety officer review was added. abiomed's medical safety officer reviewed and determined there was an issue with the peel away sheath. however, there was no impact to support or injury noted. the patient was supported for approximately seven hours and needed more support and was escalated to an impella 5.5 (delivery issues with failed attempt and adverse event) and subsequently expired. there is not enough evidence to exclude the impella 5.5 device used as an associated factor in the patient's outcome, which is captured under manufacture device report 1220648-2025-30015. the follow fields were corrected from the initial manufacturer device report 1220648-2025-30012 based on the medical safety officer review b1: revised to product problem only. b2: death and date of death were removed. h1: type of report was revised to malfunction. h6: health effect: impact code 1802, 3271, and 4643 should not been reported on the initial manufacturer device report per medical safety officer review. a new code was added. | Introducer damage - mechanical: no product was returned for analysis. clinical reported md did not unlock introducer from sheath and instead pulled introducer out of peel away sheath, breaking introducer connection, and leaving orange screw on portion on peel away sheath. issue was fixed and the case progressed. the root cause of the introducer damage - mechanical was most likely due to a use-related issue as evidenced by clinical details of steps taken before issue occurrence low pump flow: clinical reported patient was successfully cannulated for ecmo. after ecmo was initiated, impella and ecmo were unable to flow. the data logs show a motor current spike with speed deviation and a drop in flow on 6/12 at 16:13 earlier during support. there were also suction alarms. time of ecmo initiation was not noted but noted to be after patient was taken to the or. the logs show placement signal trending down. low pump flow could have been as a result of biomaterial ingestion exacerbated by patient condition. the root cause of the low pump flow was most likely patient condition related as evidenced by log pattern and clinical summary of low pump flow occurring after ecmo initiation thrombus: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7802 | 1220648-2025-30013 | 22384234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22384234&pc=OZD | 06/19/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680852 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 73 KG | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Great Vessel Perforation | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and had low platelets and thrombocytopenia was suspected. upon removal of the impella, the patient suffered a type b dissection using a j wire while attempting to close groin site with perclose device. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk. percutaneous coronary intervention. section: warnings & cautions: cautions, ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7803 | 1220648-2025-30014 | 22384423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22384423&pc=OZD | 06/11/2025 | 08/08/2025 |  | 07/02/2025 |  |  | 06/11/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679070 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 81 KG | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Cardiac Arrest | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [8 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and the patient coded. there were concerns for low hemoglobin/hematocrit levels. the patient required packed red blood cells. no source of bleeding was identified. the patient went into arrythmia and cardiopulmonary resuscitation was performed. ultimately, the patient passed away. the medical doctor did not feel this to be impella related. | It was reported multiple blood products given for bleeding. the source of the bleeding was never determined. on the last day of support the patient was noted to have hemolysis and was transfused with 2 units of red blood cells (rbc) blood product. patient expired. cause of death noted as unknown; bleeding suspected as labs showed a hematocrit was 5. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist for use during: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b2 intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30014. b5 additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30014. h6 health effect - clinical codes 1888 and 1886 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30014. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30014.
7804 | 1220648-2025-30015 | 22384732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22384732&pc=OZD | 06/12/2025 | 08/11/2025 |  | 07/02/2025 |  |  | 06/12/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680986 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female | 101 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiac Perforation | Insufficient Information | [11 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after multiple failed attempts to advance an impella 5.5 for support escalation, the procedure was aborted. the physician was unable to advance the impella due to the patient¿s anatomy. immediately following the delivery failure, the patient was placed on ecpella support with an existing impella implant. however, the patient decompensated and passed away within the hour. it is unknown if the failure to escalate support contributed to the patient's passing shortly thereafter. | A transesophageal echocardiogram showed that the patient had a large pericardial effusion. the physician opened the chest and drained effusion and multiple blood products were given. however, the patient decompensated and passed away within the hour. |  |  |  |  | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and ventricle perforation has been completed. the impella device was returned for evaluation. failure mode product damage (catheter kink) added after product evaluation since a catheter kink was found. guidewire was delivered successfully but unable to advance impella 5.5 after multiple attempts due to patient anatomy. delivery was aborted and patient was cannulated for extracorporeal membrane oxygenation (ecmo) instead. impella 5.5 was returned and a catheter kink was found below the motor housing. delivery issue - the root cause of the delivery issue was most likely patient condition related since the pump was unable to pass through difficult anatomy. catheter kink - the root cause of the product damage (catheter kink) was most likely patient condition related since the pump was unable to pass through difficult anatomy. ventricle perforation - the root cause of the ventricle perforation was not determined due to insufficient clinical information. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30015. h6 health effect - clinical code 4580 was erroneously entered on the initial manufacturer device report number 1220648-2025-30015. h6 health effect - clinical code 213 and health effect - impact codes 4643 and 4624 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30015. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-30015.
7805 | 1220648-2025-30016 | 22384892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22384892&pc=OZD | 06/13/2025 | 08/11/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669118 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was implanted for support of a patient with low ejection fraction and acute myocardial infarction/cardiogenic shock. the access site was a bleeder and was treated by femostop and an extra suture. within just over 24 hours, the pump was explanted and escalated to an impella 5.5. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient movement because the bleeding occurred after transfer.
7806 | 1220648-2025-30017 | 22384941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22384941&pc=OZD | 06/13/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662730 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 96 KG | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was placed to support an acute myocardial infarction patient intubated with flash pulmonary edema. the impella cp was placed but there were signs of hemolysis and poor positioning in the left ventricle so the pump was explanted and escalated to an impella 5.5 after 14 hours. the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7807 | 1220648-2025-30018 | 22385394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22385394&pc=OZD | 05/14/2025 | 07/08/2025 |  | 07/02/2025 |  |  | 05/14/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657719 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | Device Sensing Problem | [5 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and during support, the placement signal became intermittent later noted to be from a kink in the pump catheter. the patient would subsequently expire. the patient presented to the hospital noted with "more" heart failure symptoms and was implanted with the impella 5.5 device for mcs and bridge to transplant. approximately five days after start of the impella mcs, the placement signal became intermittent, erratic. the staff was advised that the device could still function without the placement signal and walked through how to disable the placement signal if the alarm could not be resolved. additionally, the staff was informed of how to use hemodynamics, echo monitoring and trends to monitor the support status. some time thereafter, the attending noticed a kink in the impella catheter at the dressing site. the kink was fixed, and the placement signal normalized. the status of the patient following the event was indicated as unknown. however, impella mcs continued. approximately nine days later it was noted the patient was taken to the electrophysiology lab for arterial lead placement. the patient was noted to have remained with respiratory issues requiring daily bronchoscopy once lungs were stable. later that afternoon, it was reported the patient arrested for approximately two hours. the family was called to bedside, the decision was made for comfort measures only, and the patient subsequently expired. | [5 empty] | It is unknown if there was any impact to the flow as a result of the kinked catheter. thus, there is not enough evidence to exclude the impella device used as an associated factor to the overall outcome. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation of optical signal issue and product damage (cannula kink) has been completed. no product or data logs were returned for evaluation. optical signal issue: clinical detail noted that intermittent "placement signal not reliable" alarms were noted during support with erratic ps waveforms. all other waveforms within normal limits. issue resolved in the morning. additionally, ao placement signal was reading high aortic pressures which was out of proportion from the clinical a-line. a kink in the impella catheter was observed at the site dressing. kink was fixed and the ao placement signal normalized. the root cause of the optical signal issue could not be definitively determined as no product or logs were returned for evaluation. product damage (cannula kink): clinical details noted that a kink in the catheter at the site dressing was observed during support. kink was able to be fixed. the root cause of the product damage was most likely a catheter kink per clinical details, however, the cause of the catheter kink was unable to be determined. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30018. h6 type of investigation 4112 and investigation findings 3211 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30018. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-30018.
7808 | 1220648-2025-30019 | 22385428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22385428&pc=OZD | 06/13/2025 | 08/12/2025 |  | 07/02/2025 |  |  | 06/13/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667485 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 76 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pericardial Effusion | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a who patient presented to the emergency department after a syncopal episode and an impella cp was placed. it was noted that the impella cp was started in auto mode. however, an echocardiogram was done after the patient continued to decompensate. a large pericardial effusion with coagulated blood was found which was presumed ruptured and was taken to the operating room and the chest was opened. a cross clamp was placed above the motor housing unit. cardioplegia was given with the impella on p2. the patient was placed in surgical mode and the contained rupture was repaired. the patient came off cariopulmonary bypass onto impella support without issue. during impella cp support, it was reported that the peel away sheath remained in place. however, no pulses and right lower extremity with dpc function. the impella cp was successfully weaned and explanted. | The free wall rupture was identified before the cp went it, so no concern that it was related to the impella in any way. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the free wall rupture (pericardial effusion) was identified before the implant of the impella cp, so no concern that it was related to the impella in any way. the following fields were corrected. b2 should have been other serious event as no intervention was reported for limb ischemia. b5 explanation of exclusion of pericardial effusion. h6 health effect - clinical code 3271.
7809 | 1220648-2025-30020 | 22385453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22385453&pc=OZD | 05/26/2021 | 07/02/2025 |  | 07/02/2025 |  |  | 05/26/2021 | 07/02/2025 |  |  | 03/31/2023 |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022065389 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support following a coronary artery bypass graft procedure after which the patient continued to decompensate and in acute myocardial infarction/cardiogenic shock. it was noted that the motor current had increased for the last few days and the medical team was advised to watch for pump failure. the pump level was decreased to p-8, but saturated flows at 4.2 remained. the pump abruptly stopped operating and was unable to be restarted. the device was pulled back into the aorta and the patient was observed. the patient remained hemodynamically stable following the pump stop, and the physician decided to explant the impella cp without replacing the device. no patient harm was reported. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. the returned data logs revealed the pump stopped immediately following the occurrence of the "impella stopped - motor current high" alarm. about 24 hours prior to the pump stop, the motor current and purge pressure increased suddenly. visual inspection of the returned product revealed a large purge gap. no other abnormalities were found. in conclusion, it was determined that the cause of the pump stop was bearing wear.
7810 | 1220648-2025-30021 | 22385679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22385679&pc=OZD | 05/28/2021 | 07/02/2025 |  | 07/02/2025 |  |  | 05/31/2021 | 07/02/2025 |  |  | 03/31/2023 |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059843 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support as a bridge to a durable left ventricular assist device left ventricular assist device (lvad). after 19 days of pump support, the pump abruptly stopped operating and could not be restarted. the impella cp was explanted and replaced with an lvad. the patient was noted to be stable on the lvad. | [5 empty] | The investigation has been completed. the pump was returned with the catheter cut proximal to the 50 mark. biomaterial was observed around the impeller and the shaft. the purge gap was noticeably large. no blood was noted in the purge lumen. only rt data logs were provided, with data ranging from (b)(6) 2021 at 16:28 through (b)(6) 2021 at 22:45. no pump stops were noted in the provided logs; however, the logs did not contain data from the reported date of the pump stop ((b)(6) 2021). the root cause of the pump stop was most likely bearing wear due to use beyond the intended design life of the impella cp c8. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7811 | 1220648-2025-30022 | 22385698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22385698&pc=OZD | 06/19/2025 | 08/15/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677369 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was loaded over the wire and entered the body but was not able to be advanced beyond the bifurcation due to iliac stenosis. after 40 minutes, the iliac was open with a balloon angioplasty but the impella cp was not able to be used as it had entered the body once. a new impella cp was successfully implanted. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to the tip of sheath not able to pass the heavy calcified lesion at iliac bi-furcation.
7812 | 1220648-2025-30023 | 22385949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22385949&pc=OZD | 05/14/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 05/14/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122630 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Pump | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and was transferred to a different hospital. upon aic (console) to aic transfer, an "impella defective¿ alarm appeared. a new aic used and ¿placement signal not reliable¿ alarmed. during this time, the impella was shut off three times and restarted. it was found that the impella cord not completely plugged in as a small gap was visible between console plug in and the impella cord. impella cord unplugged to assess prongs, which were all straight then replugged in, still not fitting flush into plug in. the doctor was ok with disabling placement signal (ps) unreliable alarm. the alarms were muted. support continued. however, it was further reported that care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
7813 | 1220648-2025-30024 | 22386157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22386157&pc=OZD | 05/23/2025 | 07/02/2025 |  | 07/02/2025 |  |  | 05/23/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663956 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient underwent ventricular tachycardia during support on a second impella cp. norepi weaned off and started lidocaine drops. the patient later expired however that will be captured on manufacture report number 1220648-2025-29558 follow up report number 1. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. abiomed medical safety officer reviewed this complaint and determined that ventricular tachycardia can be associated with the 2nd impella device on this report. however, there is not enough evidence to exclude the impella cp pump 1 used as an associated factor in the patient's outcome. there was an interruption in the procedure as a result of the device malfunction (no purge noted exiting the pump). however, it must be noted care was withdrawal because the patient never quite recovered. the demise outcome will be captured under the first impella cp manufacture report number 1220648-2025-29558 follow up report number 1.
7814 | 1220648-2025-30025 | 22386259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22386259&pc=OZD | 06/19/2025 | 08/18/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026682494 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult or Delayed Positioning | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was being prepped for implantation. the first .018 wire was removed from packaging and found to be bent at the tip. the team took another .018 wire from another cp box and it too was bent. the bent wire was used anyway and may have contributed to the pump coming out of position from the left ventricle and back to the aorta. due to the malpositioned pump, the wire was again replaced with a steelcore .018. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and product damage (guidewire kink) has been completed. the impella device was returned for evaluation. positioning issues: clinical details state that when impella cp went in over 2nd bent 0.018 wire, and wire was removed, cp jumped back into aorta and md said it was related to the wire. ultimately, the physician decided to remove cp and re-insert over 0.018¿ steelcore. data logs are not applicable. the 2nd guidewire was not returned. the cause of the positioning issues was not determined since insufficient clinical details were provided. product damage (guidewire kink): clinical details state that the 0.018" wire in the second cp kit was also bent, but they proceeded anyways. the 2nd guidewire was returned and evaluated. the guidewire appeared bent at the tip and the coil appeared to be damaged. the cause of the product damage (guidewire kink) was not determined since clinical details noted an out-of-box bend in the 0.018" wire but the returned guidewire could not be assessed in an out-of-box state since it had been inserted into the patient. d9 device returned to manufacturer date added.
7815 | 1220648-2025-30026 | 22386410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22386410&pc=OZD | 06/19/2025 | 08/12/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678669 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 92 KG | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a pump stop when the physician used a cross clamp to complete a coronary artery bypass graft. the procedure was completed successfully. | [5 empty] | The investigation of pump stop has been completed. the impella device was returned for evaluation. pump stop: the returned product was inspected and a kink in the catheter evident of clamping close to the kink protector of the red handle was observed. there was also a cannula kink observed. there was no biomaterial observed by the impeller. clinical mention that there was cross clamping of the pump after which there were motor current and flow issues. the data logs show a sudden drop in motor current, speed, and flows after which there were motor current spikes likely as a result of pump attempting to generate flow. a few moments later there were 'impella stopped - motor current spikes' alarms. the pump was run in the lab (unclamped) and there was no pump stop. the root cause of the pump stop was determined to be due to a use issue as evidenced by the motor current and flow drops and motor current spikes observed in the logs aligning with clinical narrative of issue occurring when pump was cross clamped. product damage (catheter kink): the returned product was inspected and a kink in the catheter close to the kink protector of the red handle was observed. clinical mention that there was cross clamping of the pump after which there were motor current and flow issues. no clinical mention of difficult patient anatomy. the root cause of the product damage (catheter kink) was most likely use related as evidenced by clamped/kinked catheter on returned product aligning with clinical details of clamping. product damage (cannula kink): the returned product was inspected and a kink in the cannula was observed. no clinical mention of difficult patient anatomy. the root cause of product damage (cannula kink) was not determined as insufficient information was provided to confirm root cause of cannula kink. d9 device returned to manufacturer date added. | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
7816 | 1220648-2025-30027 | 22386461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22386461&pc=OZD | 06/19/2025 | 08/18/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026680681 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male | 89 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a hematoma and oozing at the access site and hemolysis. a mattress suture was applied to obtain homeostasis, and heparin was withheld. purge system alarms were present. the pump was found to be positioned deep so it was repositioned and the patient was given additional volume. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding, hemolysis, optical signal issue, and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding - major: the root cause of the access site bleeding was most likely a use related issue per clinical details noting that mattress sutures were applied after oozing was observed. hemolysis: data logs were reviewed and no persistent suction or positioning alarms noted during reported hemolysis. repeated "placement signal low" alarms noted at time of hemolysis. pump was not repositioned unit the following morning. no clinical details noted if hemolysis resolved after repositioning. the root cause of the hemolysis could not be definitively determined as limited clinical details were provided and no product was returned for evaluation. optical signal issue: data logs show that a slight decrease in signal not reliable (snr) and contrast at time of "placement signal low" alarms occurring. no hard failures of the optical sensor are observed in returned data logs. clinical details noted that persistent "placement signal low" alarms triggered overnight. note that pump position was deep on imaging and not repositioned until am. the root cause of the optical signal issue could not be definitively determined as no product was returned for evaluation. positioning issues: clinical details noted that pump was deep and contacting mitral apparatus. no additional details provided on when pump came out of position. the root cause of the positioning issue could not be definitively determined as limited clinical details were provided as to when the pump came out of position. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30027. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-30027.
7817 | 1220648-2025-30028 | 22386863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22386863&pc=OZD | 06/19/2025 | 08/15/2025 |  | 07/02/2025 |  |  | 06/19/2025 | 07/02/2025 |  |  |  |  | 07/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues while implanting an impella cp. the team was unable to gain vascular access due to difficult vascular anatomy. implantation was aborted. | [5 empty] | A2 and a4 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since unable to gain vascular access due to difficult vascular anatomy.
7818 | 1220648-2025-29919 | 22369856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22369856&pc=OZD | 06/24/2025 | 08/26/2025 |  | 07/01/2025 |  |  | 06/24/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633293 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp smartassist was attempted to be inserted but the impella cp was not able to pass through the blood vessel. therefore, insertion was aborted. it was noted that the aortic arch had been replaced in the past and that the impella cp got caught on the graft and could not be passed through. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella cp device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7819 | 1220648-2025-29920 | 22370030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370030&pc=OZD | 06/11/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/11/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677276 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 84 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. right axillary was exposed and an 8 millimeter graft was sewn. left ventricle (lv) access was obtained with al1 and 0.018 wire diag catheter was removed and the impella 5.5 was backloaded onto the wire and placed into lv. support was slowly initiated up to p9 flows/5.3 liters. the first measurement on transesophageal echocardiogram (tee) showed the impella 5.5 at 5.4 centimeters with good waves and flow. the patient began to have ventricular tachycardia (vt). amiodarone/lidocaine drip was started and the patient was defibrillated twice to convert to normal sinus rhythm after repositioning with the tee which is now measuring at 4.5 centimeters and no ectopy was noted. additionally, two days later, oozing at the jp drain site and hematoma were noted. redressing was appropriate. the next day, a low dose of heparin was restarted. bleeding at the jp drain site and hematoma improved. the patient remains on impella 5.5 support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7820 | 1220648-2025-29921 | 22370168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370168&pc=OZD | 06/11/2025 | 09/03/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670359 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed bleeding soon after arrival to the intenstive care unit. blood products were administered and the impella cp was explanted by the vascular surgeon with surgical wound closure. the pump was discarded in the operating room. the patients outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement since bleeding occurred soon after arrival on icu.
7821 | 1220648-2025-29922 | 22370261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370261&pc=OZD | 06/12/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1410103 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomedical engineer reported that for an automated impella controller (aic), the pump stopped in the middle of a case. no other information was provided. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) ¿ pump stop has been completed. the impella device was returned for evaluation. analysis: this console performed as expected during testing further confirming the pump stops were not caused by the console. unrelated to the pump stops, during testing the 5s was found to fail due to optical pump connector and the usd card was corrupt. root cause: there was no problem found that would cause the pump stops.
7822 | 1220648-2025-29923 | 22370333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370333&pc=OZD | 08/20/2024 | 07/01/2025 |  | 07/01/2025 |  |  | 06/10/2025 | 07/01/2025 |  |  | 02/28/2026 |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025502874 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The protect iv study reported ae1: the patient on impella cp support had pericardial effusion not requiring drainage. this event was resolved 10 days later. cec relatedness assessment: possible related to the impella device and the impella procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7823 | 1220648-2025-29924 | 22370369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370369&pc=OZD | 06/13/2025 | 09/08/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1257686 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) could not be set to maintenance mode. the aic kept restarting with every attempt. this issue occurred during prep. there was no reported patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. data analysis: there were no rlm logs recorded on the date of the complaint. device analysis: this console was tested upon arrival and the reported failure mode was reproduced, and the rlm was unable to enter maintenance mode. after replacing the microsd card with a known good microsd card, the issue was resolved. root cause: the root cause of the inability to enter maintenance mode was microsd card corruption. the cause of the microsd card corruption could not be determined. no log evidence was found for the aic shutdown or restart issue. most likely that the inability to enter maintenance mode and the rlm restart issue were misreported as an aic shutdown or restart problem. d.9. device returned to manufacturer date added h5 was erroneously entered on the initial manufacturer device report number 1220648-2025-29924. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. e1-initial reporter name is unknown.
7824 | 1220648-2025-29925 | 22370370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370370&pc=PYX | 06/12/2025 | 07/25/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025587566 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella rp flex for mechanical circulatory support. during the removal of the peel away sheath and repo in, the impella rp flex appeared to dive deep into the right ventricle. during this flows dropped to 2.0. the impella rp flex was attempted to be repositioned but was unable to get flows above 2.0. the decision was made to replace the device. the pump was pulled back out of body and large tissue inside of inflow cage around impeller was noted. 0.018 wire was inserted into the outflow and the pump was explanted. a new flex sheath was put in followed by impella rp flex insertion without issue. flows were back to 3.8 on p-8. no patient harm was reported. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Clinical details report that on the first day of impella cp support, a purge leak was noted at the luer-lock connection in front of the microfilter on the purge sidearm. the purge filter bypass protocol was done, and reported to work well through explant on (b)(6) 2025. data logs were not available for review. returned product was investigated and the purge filter bypass was confirmed. when the tape was removed, it was confirmed that the entire purge sidearm had been cut off from the pump and not returned. therefore, the location and characterization of the damage to the sidearm could not be confirmed and investigated. based on the clinical details provided that the leak was noted to be from the purge sidearm and doing a purge filter bypass resolved the complication, the root cause of the purge leak was determined to most likely be a damaged sidearm. the exact location and characterization of the damage could not be determined since the sidearm was not returned for investigation.
7825 | 1220648-2025-29926 | 22370480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370480&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679089 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump to support the patient admitted in with a st-elevated myocardial infarction and planned workup for an left ventricular assist-device. the pump was placed when the patient suffered a stroke (cva) on day 2 of the support. the cva was noted to be treated by thrombectomy due to mca occlusion. the team gave tenecteplase and held off on anticoagulation. later there was thought the clot was due to older clot. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7826 | 1220648-2025-29927 | 22370566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370566&pc=OZD | 06/13/2025 | 07/28/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645967 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and after the implant, suction alarms began. suction was unable to be remedied or broken by decreasing p level. there was no issues with positioning. the pump was explanted and replaced. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Low pump flow: returned complaint cp pump visually inspected. cannula kink near outlet observed. no biomaterial observed. no broken blade found. impella cp connected to aic to reproduce reported low pump flow issue at it run for more than 10 minutes. flow was with specified limit at p-9. no low flow can be reproduced. data logs reviewed, suction occurred shortly after impella support started, followed by drop in flow at p4. clinical stated, shortly after impella insertion upon removing peel away sheath and inserting repo sheath suction events persisted. suction unable to be broken even at p-2. echo and fluoro used to confirm correct position. no other obvious issues with position or rh failure. patient had small lv. the cause of low pump flow issue most likely was cannula kink due to patient small lv. product damaged (cannula kink): the cause of cannula kink was most likely patient condition related due to patient small lv.
7827 | 1220648-2025-29928 | 22370606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370606&pc=OZD | 06/12/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662678 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support for just over three hours, the pump was then explanted and due to groin closure issues with the perclose. the team performed imaging and found that the superior femoral artery (sfa) had dissected. the vessel was surgically repaired and the patient's outcome was stable. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7828 | 1220648-2025-29929 | 22370787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370787&pc=OZD | 06/12/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622525 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, there was significant bleeding around the access site. the doctor then tightened the perclose sutures to stop the bleeding. the doctor proceeded with the percutaneous coronary intervention and performed an intravascular ultrasound. two drug-eluting stents were placed in the proximal/ostial right coronary artery followed by non-complaint balloons to optimize the stents. it was at that time bleeding was noted to still be around the sheath. five liters of volume (normal saline and albumin) were given. due to the patient not tolerating the procedure well and loss of volume, the doctor made the decision to leave the impella cp in and send to the intensive care unit. the position of the impella was confirmed under fluoroscopy and echocardiogram. the t-lock was locked and the tuohy was tightened. the blue suture hub was secured with angle maintained. the site was dressed and no bleeding was around the repositioning sheath and no hematoma was present. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
7829 | 1220648-2025-29930 | 22370826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370826&pc=OZD | 06/12/2025 | 09/26/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026627768 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 for the support of a mitraclip patient. the 5.5 failed to deliver as it kinked. the pump was removed and replaced by a 2nd 5.5 pump. | [5 empty] | The investigation of product damage (cannula kink) and delivery issues has been completed. per returned product investigation, product damage (catheter kink) was added as a failure mode. delivery issues: clinical reported pump became kinked during insertion into the vessel as a result of the advancement and rotation during the process. the pump was removed and replaced. the product was returned for investigation and kinks were observed in the cannula and catheter. the cause of the delivery issues is not determined due to insufficient clinical details. product damage (cannula kink): clinical reported pump became kinked during insertion into the vessel as a result of the advancement and rotation during the process. the pump was removed and replaced. the product was returned for investigation and a kink was observed in the cannula near the inlet cage. the cause of the cannula kink is not determined due to insufficient clinical details. product damage (catheter kink): clinical reported pump became kinked during insertion into the vessel as a result of the advancement and rotation during the process. the pump was removed and replaced. the product was returned for investigation and a kink was observed in the catheter near the pump. the cause of the catheter kink is not determined due to insufficient clinical details. d9. device returned to manufacture date added b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29930. d1. brand name was erroneously entered on the initial manufacturer device report number 1220648-2025-29930. g1. manufacturing site facility name was erroneously entered on the initial manufacturer device report number 1220648-2025-29930. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7830 | 1220648-2025-29931 | 22370992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22370992&pc=OZD | 06/13/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607683 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient implanted with an impella cp due to chest pain. first cp removed due to suspected clot ingestion and persistent suction alarms and a new cp inserted. unfortunately, the new cp that was inserted had constant suction at all p levels that could not be resolved. the device was repositioned multiple times. no right heart failure. appropriate volume. it was concluded that anatomy or possibly clot ingestion was causing the issue. decision was made to remove the second cp and put in an intra-aortic balloon pump. patient was stable during all troubleshooting. | [5 empty] | The investigation for the low pump flow has been completed. impella cp was returned for evaluation. upon visual inspection, no biomaterial or pump damage was observed on returned pump. pump was connected to lab console and run for approximately 9 minutes at p-9 and low pump flow did not reproduce. data logs were reviewed and logs showed low pump flows with suction alarms and some positioning alarms. no motor current spikes or ingestion signatures observed. clinical details noted that immediately after pump was inserted and peel-away was removed, suction alarms were triggered even down to p-2. a new aic was used in attempt to resolve the issue without success. no obvious issues noted on fluoro or echo. pump was repositioned multiple times without success in resolving low flows. it was noted that patient had a small lv but no kinks or supporting patient volume or condition info available. the cause of the low pump flow issue was not determined since no abnormalities found with return product and due to insufficient clinical info supporting small lv. e3 corrected to physician. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
7831 | 1220648-2025-29932 | 22371010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371010&pc=PYX | 06/12/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625249 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. the initial plan was to take the patient to cvor urgently for escalation of impella cp smartassist to an impella 5.5 and leave on bipella support. the impella rp flex had a clot ingestion issue and little flows. the decision was made to convert to veno-arterial extracorporeal membrane oxygenation (va ecmo)/impella cp smartassist. a drainage cannula was placed using the existing impella rp flex site utilizing the trojan horse technique. the patient was placed on va ecmo. the patient will remain on va ecmo/impella cp smartassist support and the team will reevaluate the need to escalate to an impella 5.5. the patients outcome is stable. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow and thrombus has been completed. the impella device was returned for evaluation. the data logs showed that about 21 hours into support, there was a large spike in motor current, placement signal and purge pressure, where the motor current remained elevated after. there was a clot returned in a cup, with the returned pump. there were no cannula kinks found but the pump had biomaterial wrapped around the impeller. based on the data logs and returned product, the root cause of the low pump flow is most likely due to biomaterial ingestion. the root cause of the thrombus was unable to be determined due to insufficient clinical details. b.2 required intervention was added as it was omitted on the initial manufacturer device report number 1220648-2025-29932. h.1 type of report was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-29932.
7832 | 1220648-2025-29933 | 22371031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371031&pc=OZD | 06/05/2025 | 08/12/2025 |  | 07/01/2025 |  |  | 06/06/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657569 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that a placement signal dashed out with placement signal unreliable alarms. likely optical sensor failure. | The user facility reported a 5.5 pump support placed for the patient admitted in with known cardiac history and worsening heart failure. during the support period, the patient experienced tachyarrhythmia, which was treated with amiodarone and bisoprolol. the arrhythmia resolved without further issue. the patient remained on impella support. |  |  |  |  | The investigation of vt/vf/at/af has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. additional information received 09aug2025 reporting placement signal issue that occurred (b)(6) 2025. the following fields were updated to reflect the additional information. the investigation is not complete should investigation results become available a supplemental report will be submitted. b1 added product problem. b5 updated . h6 medical device problem code added. | The impella device was not received from the customer as the device remains implanted supporting the patient pending an orthotopic heart transplant. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7833 | 1220648-2025-29934 | 22371047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371047&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544873 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 89 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after implanting an impella 5.5 the p level could not be increased. the pump was explanted and replaced with a new impella 5.5. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7834 | 1220648-2025-29935 | 22371187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371187&pc=OZD | 06/12/2025 | 08/08/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655018 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 62 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella 5.5 for mechanical circulatory support. after multiple failed attempts to cross with the pump and due to patient's anatomy, the physician made the decision to abort the impella 5.5 placement and decided to place an impella cp instead. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
7835 | 1220648-2025-29936 | 22371321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371321&pc=OZD | 06/13/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026679371 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that despite multiple efforts and techniques an impella 5.5 pump could not be implanted successfully due to tortuous anatomy. the pump was explanted and a second impella 5.5 was attempted unsuccessfully. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the delivery issue has been completed. clinical details states that they unable to delivery the impella 5.5 through the innominate artery due to tortuous anatomy. multiple guidewires were attempted without success and delivery of the pump was abandoned. the impella 5.5 was returned and no cannula kinks or catheter kinks were observed. the root cause of the delivery issue was most likely patient condition related since the patient had tortuous anatomy.
7836 | 1220648-2025-29937 | 22371385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371385&pc=OZD | 06/13/2025 | 08/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623880 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of two impella 5.5 devices was unsuccessful due to patient anatomy. the patient remained on va ecmo. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. initial impella 5.5 delivery was aborted and the complaint device delivery was aborted due to tortuous patient anatomy. catheter kink was noted in the field. patient was left on va ecmo. impella 5.5 was returned and a catheter kink was found. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient had tortuous anatomy. catheter kink - the root cause of the catheter kink was patient condition related since the patient had tortuous anatomy. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29937. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7837 | 1220648-2025-29938 | 22371405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371405&pc=OZD | 06/12/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 62 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, micro bubbles were found at the aorta. after having placed the catheter, positioning was verified. while assessing the position, the anesthesia provider observed micro bubbles at the aorta. micro bubbles were observed again after a couple of minutes. no harm was caused to the patient. it is unsure where the micro bubbles could have originated from. the impella purge was examined and no issues were noticed on tubing, catheter, or alarms on the automated impella controller (aic). the procedure was successful and completed with this impella device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | The investigation of embolism has been completed. the pump was not returned for investigation. the data log shows that the pump was inserted into the patient's body prior to completing the case start. the cause of the embolism issue was most likely use related since pump was inserted into patient before completion of case start priming steps.
7838 | 1220648-2025-29939 | 22371530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371530&pc=OZD | 06/18/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/18/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674734 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female | 79 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed ischemia during impella cp support. it was reported that the patients left foot was colder than right foot and vascular ultrasound was completed showing there was minimal blood flow down the leg but not completely occluded. the patient was brought to back to catheterization lab for antegrade sheath for distal limb reperfusion. the procedure was successful and was completed with this device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
7839 | 1220648-2025-29940 | 22371805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371805&pc=OZD | 06/13/2025 | 10/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657973 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Ventricular Fibrillation | [8 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient implanted with an impella 5.5 for support due to coronary artery bypass graft. the 5.5 was successfully placed. during support it was noted that the 5.5 site was oozing significantly as well as all other access sites. the physician expressed frustration with impella and said ¿it¿s not helping the patient and not flowing enough". the team attempted to increase impella to p5 however had suction event so it was left at p4. it was noted that the patient remained on amiodarone infusion for intermittent non-sustained ventricular tachycardia (nsvt). however, the patient had an episode of nsvt with sedation. additionally there were intermittent alarms with loss of pulsatility with runs of ventricular fibrillation (vf/vt). decision to switch to procainamide, however was on procainamide for 24 hours, and was noted to not have helped with intermittent vt. the patient was taken in for vt ablation and still had runs of vt. it was also noted that the patient had maroon stool and had been bleeding out of his mouth, so holding heparin. when heparin was restarted at a low dose, there was still bleeding present at the mouth and groin. also noted was that the patient was taken to or emergently for cold leg. placement signal significantly elevated compared to art line. motor current remained stable. no hemolysis concerns or other issues were noted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of access site bleeding, vascular damage - peripheral vessel, ventricular fibrillation/ventricular tachycardia, limb ischemia, and low pump flow has been completed. the impella device was not returned for evaluation. no product was returned for analysis. clinical noted bleeding from impella access site as well as other access sites. the root cause of the access site bleeding major was most likely patient condition related as evidenced by clinical mention of bleeding from multiple access sites as well as impella access site in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the ventricular fibrillation/ventricular tachycardia and limb ischemia was unable to be determined due to insufficient clinical details. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided. clinical reported suction alarms when p5 was attempted from p4. there were no motor current spikes outside p-level changes. suction alarms in the logs were confirmed. the logs confirm slightly lower flow for p5 but appropriate flows for p4. the patient was also on extracorporeal membrane oxygenation (ecmo) during support. the root cause of the low pump flow was mist likely patient condition related as evidenced by suction alarms when higher p-level attempted and concurrent ecmo use with no motor current spikes and appropriate pump flows at lower p-levels. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ a.4 age was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-29940. d.4 revised unique identifier (udi) # as it was entered incorrectly on the initial manufacturer device report number 1220648-2025-29940. e.1 revised facility name as it was entered incorrectly on the initial manufacturer device report number 1220648-2025-29940. added phone number and fax number as they were inadvertently omitted from initial manufacturer device report number 1220648-2025-29940. h.6 component code 1721 was reported incorrectly on the initial manufacturer device report number 1220648-2025-29940. codes 2095 and 2130 were added. code 4644 and 3038 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29940. h.10 revised instructions for use since the initial manufacturer device report number 1220648-2025-29940 was submitted. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned for analysis. the clinical report mentioned placement signal became significantly higher than the arterial. motor current remained stable. there were no placement signal not reliable alarms, but placement signal monitoring was disabled. the pattern observed in the data logs is characteristic of a possible fiber break. the root cause of the optical signal issue was most likely due to a damaged fiber line as evidenced by data log pattern observed but details related to possible damage were not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-29940. d10 concomitant medical products and therapy dates updated.
7840 | 1220648-2025-29942 | 22371880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371880&pc=OZD | 06/12/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667544 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp, the peel away sheath could not be advanced due to the patient's anatomy. the patient experienced cardiac arrest and required epinephrine, atropine, and bicarbonate. the physician resumed implantation using a 11 french terumo sheath and an intra-aortic balloon pump was implanted instead. cardiopulmonary resuscitation continued but return of spontaneous circulation was not obtained and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation difficulty inserting/removing introducer and cardiac arrest has been completed. the impella device was not returned for evaluation. the cause of the difficulty inserting introducer issue was most likely patient condition related, based on given clinical details. the root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. b.2 required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2025-29942.
7841 | 1220648-2025-29943 | 22371886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371886&pc=OZD | 06/21/2025 | 09/10/2025 |  | 07/01/2025 |  |  | 06/23/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561819 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following startup during weekly routine inspection, the display brightness of the automated impella controller would drop significantly. the device was rebooted multiple times, but the issue persisted. there was no patient involvement at the time of the failure. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. data analysis: upon reviewing the logs, no evidence was found to confirm the display flashing or brightness issue. device analysis: this console was tested. during testing with the cassette and pump, neither the lashing/flickering screen issue nor the brightness issue could be reproduced. the lvds cables and backlight led cables were confirmed to be appropriately connected. the voltage at the 4-in-1 connector (x25) was measured at 23.78v under ac power, as intended. root cause: the root cause of the flashing display and brightness issue could not be determined due to the inability to reproduce it. d9 device returned to manufacturer date added b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29943. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. e1-initial reporter name is unknown.
7842 | 1220648-2025-29944 | 22371938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22371938&pc=OZD | 06/24/2025 | 08/21/2025 |  | 07/01/2025 |  |  | 06/24/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666662 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support placed for the acute myocardial infarction patient. after just 15 minutes of support the pump was removed due to low pump flows. the team replaced the pump and support proceeded. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for low pump flow has been completed. the pump was returned for investigation. data log shows low pump flow from the start of the case run, with several motor current spikes observed 30 minutes later, after which the pump was explanted. the returned pump had biomaterial in the inflow cage. no biomaterial was found on the impeller and inflow cage. the pump was started in a bucket of water, and biomaterial was seen blocking the impeller, causing the pump to stop. after removing the biomaterial with a dental pick, the pump was able to run according to specifications. the cause of the low pump flow issue was most likely biomaterial, based on data log review and returned product investigation. thrombus failure mode was added as an investigated failure mode based on low pump flow and returned product investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7843 | 1220648-2025-29945 | 22372341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372341&pc=OZD | 06/12/2025 | 08/26/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637509 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was unable to be completed due to the patient's vascular anatomy. an impella cp pump was placed successfully instead and the patient's outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to be delivered due to patient anatomy.
7844 | 1220648-2025-29946 | 22372345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372345&pc=OZD | 06/13/2025 | 07/30/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504204 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, the patient had arrhythmias. amio bolus was given to the patient. additionally, there were placement signal issue. placement was confirmed with tee. there was also a motor spike, motor current went flat, flows increased and impella position in aorta alarm started. placement signal was disabled and flows returned to normal ranges at p-2. support continued. | [5 empty] | Vf/vt/af/at: the root cause of the arrythmia and vt was unable to be determined due to insufficient clinical details. optical signal issue: data log review showed a motor current spike at 6:34 pm on (b)(6) 2025 as mentioned in the clinical. after the spike, motor current lost pulsatility and pump flow increased. impella in aorta alarms triggered as a result. the 5s parameters were stable. product was returned and biomaterial was found near the sensor, in the outlet cage, and wrapped around the impeller. prior to soaking the pump in water, the 5s parameters were out of spec with snr near zero. after soaking the pump in water, the 5s parameters greatly improved. the root cause of the optical signal issue was most likely biomaterial that was triggering false impella in aorta alarms since biomaterial was observed on the returned product and data log review showed a motor current spike prior to the alarms. saturated flow: data log review showed a motor current spike at 6:34 pm on (b)(6) 2025 as mentioned in the clinical. after the spike, motor current lost pulsatility and pump flow increased. product was returned and biomaterial was found near the sensor, in the outlet cage, and wrapped around the impeller. there was no blood ingress into the purge line past the motor housing. the root cause of saturated flow was most likely biomaterial ingestion since data log review showed a motor current spike followed by increased flows and loss of motor current pulsatility and biomaterial was observed on the returned product. thrombus: the failure mode thrombus was added since data log review showed a motor current spike and biomaterial was observed on the returned product likely related to the saturated flow and false impella in aorta alarm. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
7845 | 1220648-2025-29947 | 22372454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372454&pc=OZD | 06/13/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608533 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump failed initial wireless delivery into the ventricle. it was noted that the pump would not cross the aortic valve. a j wire and pigtail were subsequently placed, followed by a v-18 wire, with which the pump was guided across the aortic valve. there was no patient harm resulting from the initial delivery failure. | [5 empty] | The investigation for the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. impella 5.5 pump failed initial wireless delivery into the ventricle but later a j wire and pigtail were subsequently placed which the pump was guided across the aortic valve. the root cause of the delivery issue was most likely use related. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
7846 | 1220648-2025-29948 | 22372556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372556&pc=OZD | 06/03/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/03/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female | 81 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was in cardiogenic shock from an acute myocardial infarction and was implanted with an impella cp for mechanical circulatory support (mcs) and experienced an access site hematoma requiring intervention. the patient presented to the emergency room with palpitations and anxiety. the patient went into ventricular fibrillation arrest requiring epinephrine and one shock with return of spontaneous circulation achieved after eight minutes. the patient was taken to the catheterization lab and found to have multi-vessel disease. the impella cp was selected and placed without incident. however, approximately 1.5 hours later it was noted that a hematoma had developed at the right groin (impella cp access site) requiring one unit of packed red blood cells with femostop placement. the patient was noted to be in stable condition following the event and the impella cp mcs continued. approximately three hours later, the impella cp was explanted and an impella 5.5 was placed for continuation in therapy. however, the patient would experience an impella 5.5 related event and expire approximately eight days later. | [5 empty] | Medical safety officer reviewed the event and determined that the impella cp related event is a serious injury for the access hematoma requiring blood product. however, there is not enough evidence to exclude the impella 5.5 device as an associated factor to the patient's outcome. the last day of support, the patient was noted to have hemolysis and was transfused with two units of blood product. the patient expired. and the cause of death was noted as unknown, bleeding suspected. the impella cp is one of two devices associated with the patient's implant experience. refer to medical device report for the impella 5.5 serial number (b)(6) with date of event (b)(6) 2025 and patient identifier ending in (b)(6). the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7847 | 1220648-2025-29949 | 22372893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372893&pc=PYX | 06/12/2025 | 09/17/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026668109 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella rp flex and generated pulmonary artery pressure readings of 100/100. the impella rp flex was explanted and replaced with a new pump to resolve the issue. the patients outcome is stable. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Corrections to the initial mfr report: h6 type of investigation code 10 added.
7848 | 1220648-2025-29950 | 22372956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372956&pc=OZD | 06/02/2025 | 08/13/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1922223 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that there was a controller error on the console. the console was exchanged without issue. there was no harm to the patient. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the logs show the console had motor communication issues at 20:05 followed shortly after by the console rebooting, correctly triggering an unexpected controller shutdown alarm. 8 hours later pump motor driver (pmd) and epm were reset followed by the ossi subprocess crashing and triggering a sw wdg as well as a controller error #1021 for loss of communication to the epm and #1047 for pmd comm loss and triggering a pump stop as a result. there were numerous communication issues and reboots until the pump was removed 5 minutes later with various other subprocesses failing. device analysis: the failure mode was reproduced as the console was unable to boot up due to communication issue. replacing the 4-in-1 resolved the issue. root cause: the cause of the communication issues to the epm and pmd as well as unexpected reboots and pump stops were all cascading failures of a malfunctional 4-in-1.
7849 | 1220648-2025-29951 | 22372974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22372974&pc=OZD | 06/12/2025 | 08/08/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481981 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male | 92 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced rising purge pressure. there were concerns of biomaterial ingestion. the team explanted and replaced the pump. the patients outcome is stable. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was returned for evaluation. the pump was returned for investigation. biomaterial was observed on the impeller and in the purge gap. the pump was started for priming, and high purge pressure was seen initially, which gradually resolved after pump start. at the case start, data log shows a gradual rise in purge pressure around 800 mmhg, followed by a small motor current spike, which caused the purge pressure to rise further until the high purge pressure alarm occurred. the cause of the high purge pressure issue was biomaterial ingestion, based on data log review and returned product investigation. thrombus failure mode was added as an investigated failure mode based on the returned product investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7850 | 1220648-2025-29952 | 22373105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373105&pc=OZD | 06/15/2025 | 08/13/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643967 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 90 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient was being transferred to a different hospital. the seatbelt was placed over the impella insertion and this caused bleeding and a hematoma. a femstop was placed and this was effective. the pump was explanted and replaced with an impella 5.5. it was stated that the bleeding was not the sole reason for the replacement. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation of access site bleeding has been completed. site bleeding: clinical reported that during transportation, the transport team placed the seatbelt over the insertion site, which led to bleeding and a hematoma. the pump was returned for investigation and no kinks or abnormalities were found. the cause of the access site bleeding was most likely a use issue, as the clinical reports a seatbelt was placed over the insertion site during transport, causing the hematoma.
7851 | 1220648-2025-29953 | 22373187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373187&pc=OZD | 06/15/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026686614 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 66 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed after the patient deteriorated after the prior 5.5 pump was explanted and prior cardiac surgery. the cp was placed along with extracorporeal membrane oxygenation (ecmo). the ecmo limb got perfusion line placement, no allegation of the impella causing the need for perfusion. the cp support continued while condition monitored and again 5.5 escalation discussed. the team observed left atrial clot on day 9 of the cp support. the cause of the clot could be the prior mitral repair. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.' "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7852 | 1220648-2025-29956 | 22373340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373340&pc=OZD | 06/15/2025 | 08/13/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675450 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient(pt) presented to emergency room and coded a total of 5 times. once return of spontaneous circulation (rosc)was achieved dr. (b)(6) decided for impella cp placement if adequate vasculature and percutaneous coronary intervention (pci). it was noted that once the patient was moved over to the table and the patient coded again. cardiopulmonary resuscitation performed while patient was prepped for impella placement. rosc achieved and impella inserted via the right femoral artery utilizing best practices without difficulty. flows of 3.4 l/min obtained. pci of lad performed with 3 stents being deployed. pci of circ was then done with 1 stent being deployed. during support impella alarmed suction and unable to flow above 1.0 l/min and waveforms became dissociated. an echocardiogram (echo) performed again showing impella in the proper position free in the mid-ventricular space. discussion had with physicians about possible bio-material ingestion. decision made to explant the pump. while preparing for transfer an echo was performed to check for thrombus in the left ventricle and none was visualized. the impella was explanted without difficulty. impella cp was examined and bio-material was found in the inlet. decision made by dr. to not implant another cp due to possibility of more biomaterial ingestion. additional information received: the pump is not available for return; it was disposed of by the staff. the patient passed away the next day. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. low pump flows: data logs were reviewed and the drop in flows with suction was confirmed. while running at p-9, pump flows suddenly dropped from 3.65 l/min to 1.02 l/min, and 8 seconds later, the suction alarm triggered. the suction alarm triggered because motor current (mc) max values were very low, indicating a high differential pressure even during systole, when the valve is open. during the subsequent low flows, the placement signal (ps) remained pulsatile and aortic with values around 119/63 mmhg, suggesting that the patient's left ventricle (lv) was contracting. this all aligns with the clinical report of biomaterial being visualized in the inflow cage upon explant and no cannula kink. since biomaterial was in the inflow cage, it had not migrated down to the impeller, and therefore, no motor current spike is to be expected. the biomaterial likely occluded the cannula from typical flow rates at p-9. product was not returned for investigation. based on the clinical details provided that biomaterial was visualized in the inflow cage and data logs showed a sudden drop in pump flows while ps remained stable, the root cause of the low pump flows was determined to most likely be biomaterial. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7853 | 1220648-2025-29955 | 22373348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373348&pc=OZD | 06/23/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/24/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661757 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complainant reported the patient had an impella 5.5 implant and after the pump was implanted, the purge pressure gradually increased and a purge pressure high alarm appeared. various measures were attempted such as replacing the purge set and increasing the amount of purge heparin, but there was no improvement. additionally, it was stated there was a bend in left subclavian artery of about 90 degrees, and a kink was suspected. the pump was explanted and replaced. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7854 | 1220648-2025-29941 | 22373411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373411&pc=OZD | 06/15/2025 | 09/05/2025 |  | 07/01/2025 |  |  | 06/17/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026646680 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was attempted to be placed but it was difficult to access the left ventricle due to dextrocardia. the shaft was confirmed to be kinked and bent. the impella insertion was aborted. | [5 empty] | The investigation of delivery issues and product damage (cannula kink) has been completed. pump was returned for investigation. data log review was not required for this case run. cannula kink damage was observed on the returned product. no other physical damage was observed on the returned product. the case start menu appeared on the pump connected to the console, indicating that the case start was not initiated in the field. as per the clinical report, the cannula was kinked and bent during insertion, and the doctor had difficulty with placement due to dextrocardia. the cause of the product damage (cannula kink) was most likely patient condition related to abnormal anatomy. the cause of the delivery issue was most likely patient condition related to abnormal anatomy. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29941. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system. to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
7855 | 1220648-2025-29957 | 22373495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373495&pc=OZD | 06/18/2025 | 08/14/2025 |  | 07/01/2025 |  |  | 06/18/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642063 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 74 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient presented with st-segment elevation myocardial infarction and shortness of breath. during impella cp support, it was noted that pulses were not palpable or appreciated on doppler. bilateral lower extremities were cool but not cold. cap refill was less than three seconds. additionally, there were reports of hemolysis. transthoracic echocardiogram was done and the pump was repositioned. later, the urine started to clear up. the impella cp was successfully weaned and explanted. the patient survived and the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. limb ischemia - reversible: clinical details noted that clinical team was unable to find distal pulses of extremities. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. hemolysis: data logs were reviewed and logs showed repeated suction alarms with placement signal spiking to 3000 at time of reported hemolysis. no repeated positioning alarms occurred during reported hemolysis. clinical details noted that patient was experiencing red urine overnight on first day of support. transesophageal echocardiogram (tee) was performed, pump was repositioned, and turned down to p-5. noted that patient had high map and had an afterload issue that was treated the next morning with hemolysis resolving. the root cause of the hemolysis was most likely patient condition related as data logs showed placement signal trend supported with suction alarms, and clinical details noting hemolysis was able to resolve after patient's afterload was addressed.
7856 | 1220648-2025-29960 | 22373969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373969&pc=OZD | 06/15/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673714 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella cp support who was admitted in and awaiting transplant. it was reported that the team was concerned for heparin-induced thrombocytopenia as platelets have gone from 89 to 6. this did not prompt explant or replacement. the pump remained on for support until successful explant and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7857 | 1220648-2025-29954 | 22373995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22373995&pc=OZD | 06/13/2025 | 10/06/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578563 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | During evaluation and investigation of the impella 5.5, there was low pump flow noted. the data logs confirmed slightly low flow for p6. | The complainant also reported that, during support, there were placement signal issues. the team was aware and silenced the alarm. support continued. |  |  |  |  | The impella device was not returned, and the investigation has been completed. the data logs confirmed slightly low flow for p6. the logs do not show any motor current spikes before low flow. there were impella position unknown alarms at the start of the case. clinical did not mention any position related issues or patient conditions or extracorporeal membrane oxygenation support. the root cause of the low pump flow was unable to be determined as no product was returned for analysis and limited clinical information was provided. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the data logs show steady 5s parameters with three placement signal spikes in isolation before later during support on 12/25 placement signal goes out of range suddenly. at the same instant, contrast increases with an increase in amplitude, lamp steps up, and gain decreases. no clinical related information was given to align with timeline of placement signal not reliable alarm. the log pattern shows a potential fiber break issue but that cannot be definitively determined as logs do not match pattern and no product was returned. the root cause of the optical signal issue was not determined as no product was returned to confirm potential issues identified via logs and limited clinical information related to failure was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated.
7858 | 1220648-2025-29961 | 22374036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22374036&pc=OZD | 06/18/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/18/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662070 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 75 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient experienced runs of ventricular tachycardia during impella cp support. the patient required three shocks in response to the condition and lidocaine and amiodarone drips were initiated. support was continued with the implanted pump following the intervention. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7859 | 1220648-2025-29962 | 22374110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22374110&pc=OZD | 06/16/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/16/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662793 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had a ct scan that revealed a brain bleed. heparin was withheld and one unit of red blood cells were given to the patient. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7860 | 1220648-2025-29958 | 22374243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22374243&pc=OZD | 06/15/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662345 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that a patient with acute myocardial infarction and cardiogenic shock was placed on impella cp for hemodynamic support. during support it was reported that the patient had experienced leg ischemia that began during/after introducer insertion. a senior doctor checked position and leg perfusion(fem.-fem bypass). it was noted that the pump ran without problems on p-6, however higher p-level was not possible due to suction alarms. discussions of escalation to bipella were had however, not an option due to patient overall bad conditions and bad prognosis. additionally the patient had a early onset of sepsis, as a result of that extravasal volume shift and suction alarms. ultimately care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7861 | 1220648-2025-29959 | 22374363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22374363&pc=OZD | 06/18/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/18/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674768 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 72 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a thrombus was discovered in the apex of the left ventricle during impella cp support. in addition, it was noted that the patient's platelet level had dropped into the 30's prompting an heparin-induced thrombocytopenia (hit) panel to be taken. the patient was switched to argatroban in response to the potential hit. support was continued uninterrupted following the intervention. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
7862 | 1220648-2025-29963 | 22375020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375020&pc=OZD | 06/16/2025 | 08/01/2025 |  | 07/01/2025 |  |  | 06/16/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645714 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 106 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported after impella cp was placed, in intensive care unit intensive care unit an echocardiogram at bedside showed potential thrombus near the end of pigtail. it was further reported that on day 2 of impella support the cp was inadvertently pulled out of the left ventricle. the cp was removed and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of positioning issues and thrombus has been completed. the impella device was not returned for evaluation. positioning issues: cp removed in icu after pump by dr. lawson was inadvertently pulled out of left ventricle (lv). the cause of the positioning issue was most likely use related as it was inadvertently pulled out of the lv. thrombus: echo bedside shows potential thrombus near end of pigtail, md aware. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7863 | 1220648-2025-29964 | 22375101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375101&pc=OZD | 04/16/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401317 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a console exhibited erratic pressure support behavior during use. upon investigation, additional malfunctions were discovered. the investigation confirmed defects in the power board module and the blue receptacle, which were determined to be the root cause of the pressure support irregularities. no harm was reported. | [5 empty] | The investigation has been completed. the console logs, although incomplete, were mostly consistent with the complaint and confirmed optical signal issue, purge system failure alarm and battery-related alarm. partial logs show battery level low alarm, shortly before pump with serial number (b)(6) was transferred to ic7938. battery and power data was embedded in teh long term log file for pump (b)(6) and showed that the alarm was triggered due to a single erroneous sample of remaining capacity parameter read as 0. real time log data showed raw optical signal often barcoding between two values separated by approx. 110mmhg, which coincides with moments of lower signal-to-nosie ratio (snr) (which was low from the start, as evidenced by long term log data). although most spikes were removed by the software algorithm therefore not showing on thw automated impella controller (aic) screen, those values affect placement signal calculation. the console was returned for evaluation finding a deficiency of its optical system was revealed, as the device would fail optical ¿5s¿ parameter specification due to low snr. analog output of ccd detector in theoptical bench did not show any significant distortion, but overall signal level was lower than typical. inspection of the optical fiber in pump connector revealed some contamination that could not be fully removed. battery-related issue, resulting in erroneous values, were not reproduced during testing, but most likely was related to malfunction of communication between power battery manager (pbm) and batteries. the cause of the aic issue was a defective blue receptacle. the cause of the aic issue was most likely a defective pbm board.
7864 | 1220648-2025-29965 | 22375128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375128&pc=OZD | 06/14/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592238 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump supporting a patient admitted in with conversation of high risk cardiac procedure. the cp was supporting for 16.5 days when the pump stopped. the pump was removed and replaced by a new cp. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
7865 | 1220648-2025-29966 | 22375292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375292&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669147 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a coronary artery bypass graft patient was placed on impella 5.5 as a bridge to decision. it was noted that he pump was running without any problems . however after the nurse placed the feeding tube, the patient went into ventricular tachycardia and was cardioverted and noted to be stable. the impella position was verified. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7866 | 1220648-2025-29967 | 22375375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375375&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1725570 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller power cord frayed. the power cord was not in use. there was no patient use reported with this event. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7867 | 1220648-2025-29968 | 22375385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375385&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654944 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Hematoma | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient implanted with an impella 5.5 support experienced sinus tachycardia requiring defibrillation. additionally, a stroke code was called though a ct was performed and showed no signs of bleeding. it was further reported that the patient had hematoma near the impella access site and therefore, heparin was withheld. the patient's family decided to withdraw care, and the patient expired. | [5 empty] | A4 and a5, race, are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7868 | 1220648-2025-29969 | 22375590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375590&pc=OZD | 06/13/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026649354 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the impella cp was being set up and primed, the impella stopped. the nurse changed the p-level to p0 to resolve the issue, and the pump was placed without issue and operated in auto mode. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. the data log shows that the pump started before insertion into the patient's body. additionally, the nurse confirms that the red cheater was still through the pump during the pump stop, and pump was later inserted successfully and ran without issue. the cause of the temporary pump stop issue was most likely red lumen removal issue. red lumen should be removed before insertion of pump and pump start.
7869 | 1220648-2025-29970 | 22375831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375831&pc=OZD | 06/13/2025 | 09/08/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617983 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced low flow. the pump was repositioned using echo; however, the issue did not resolve. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. data log reviewed: no motor current (mc) spike or placement signal issue seen. no purge spike observed. impella in ventricle alarms observed after case initiated at p9 followed by continued suctions alarms witch correlate with clinical details position remained problem, flows dropped during the impella support. positioning issues: the cause of the positioning issue most likely was patient condition due to dry heart. low pump flow: the cause of low pump flow issue most likely was improper positioning related due to improper technique. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7870 | 1220648-2025-29971 | 22375901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375901&pc=OZD | 06/12/2025 | 08/08/2025 |  | 07/01/2025 |  |  | 06/12/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced atrial flutter with hemodynamic collapse during impella cp support, which was treated with cardioversion. the patient also developed access site bleeding, managed with manual pressure and blood transfusions, as well as hemolysis, evidenced by red urine. the care team administered albumin and additional blood products to manage the hemolysis. mechanical support continued until the patient stabilized and escalated to impella 5.5. the impella cp was subsequently explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the instructions for use for the impella cp with smart assist system states the following: impella cp with smartassist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of the access site bleeding was most likely patient condition related due to the bleeding occurring at multiple access sites. vf/vt/af/at: the root cause of the af / sinus tachycardia was unable to be determined due to insufficient clinical details. hemolysis: data log review showed suction alarms on days of hemolysis with no motor current (mc) spikes and adequate positioning. the pump was not returned. the cause of the hemolysis was most likely patient condition related as the pump positioning and flows were deemed good and the hemolysis resolved after giving the patient blood volume. instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-29971.
7871 | 1220648-2025-29972 | 22375915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375915&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support developed limb ischemia requiring device explanation. the patient with a history of hypertension suffered a cardiac arrest. the decision was made to implant an impella cp device for mechanical circulatory support. the device was placed successfully with local vascular support, and immediate hemodynamic improvement was noted post-implantation. after pressure was applied following explantation, mottling of the right thigh was noted, and distal pulses were absent on doppler examination, raising concern for limb ischemia. manual pressure was then released; hemostasis was maintained, but poor perfusion persisted initially. however, the mottling improved, and an arterial duplex study confirmed the presence of distal flow. a heparin drip was initiated for anticoagulation. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
7872 | 1220648-2025-29973 | 22375925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375925&pc=OZD | 06/13/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/13/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026687392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (hrpci) with mechanical circulatory support using an impella cp device developed access site bleeding complications requiring bedside device explantation. the patient, with a history of hypertension, was found to have multivessel coronary artery disease. the patient declined coronary artery bypass grafting (cabg) and instead elected to proceed with impella-supported hrpci. an impella cp device was successfully implanted. upon arrival in the icu, significant groin bleeding and hematoma formation were noted at the impella access site. initial attempts to control the bleeding with redressing and manual pressure were unsuccessful. the decision was made to explant the impella device. the device was removed, perclose sutures were deployed, and manual pressure was applied to achieve hemostasis. the procedure was successful, and the bleeding was controlled. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7873 | 1220648-2025-29974 | 22375928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22375928&pc=OZD | 06/14/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599822 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and, during support, developed a purge leak alarm requiring device explantation. the patient had experienced a cardiac arrest and was found to have multi vessel coronary artery disease. an impella cp device was initially placed for mechanical circulatory support. due to ongoing cardiogenic shock and the need for continued hemodynamic support, the decision was made to escalate from impella cp to impella 5.5. an air-in-purge alarm was triggered on the impella console. the care team attempted to purge the line; however, semi-continuous purge alarms persisted. the automated impella controller was replaced, but the alarms continued. the device was then connected to a second console, yet the alarm persisted despite the change. due to the continued purge system malfunction, the care team proceeded with sending the patient to the operating room for emergent surgical left ventricular assist device (lvad) placement. the impella 5.5 pump was explanted. the patient¿s procedural outcome was unknown at the time of reporting. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
7874 | 1220648-2025-29975 | 22376030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376030&pc=PYX | 06/14/2025 | 10/24/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025528732 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 82 KG |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient (pt) was admitted for unexplained anemia, weight loss, and abdominal pain. cardiac eval/tte revealed vegetation on the aortic valve leaflets. ef 45%. additionally, pt was diagnosed with acute cholecystitis. abx treatment started and pt optimized for surgery. the patient underwent bentall. surgery was complicated by complex removal of valve and friable tissue. decision was made for urgent placement of impella 5.5 as they were unable to wean from bypass. on arrival in the or, the impella 5.5 was inserted and initiated at p2. upon assessment of hemodynamics and heart function, physician made the decision to proceed with insertion of rp flex. during support, a nurse called about placement signal alarm with the rp. alarm explained and audio was silenced. additionally, there was a notification with the rp of purge pressure low. the purge flow 30, pp in 300's. however, was intermittent after bag change. no leaks noted. it was noted that the patient on bipella support was critically ill with poor prognosis, absent of any neurological function, on multiple high-dose pressors/inotropes with acute-onset multiple organ dysfunction syndrome. family made decision to withdraw care and the patient expired. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. after 3 days on support, the placement signal not reliable (psnr) alarm occurred. data logs were not recovered. the cause of the placement signal issue was not determined since product was not returned, data logs were not recovered, and insufficient clinical details were provided. also, after 3 days on support, the customer service center was called for a purge pressure low alarm, the pump was not returned. data logs were not recovered. the cause of the purge leak - pump was not determined since product was not returned, data logs were not recovered, and insufficient clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. a4 updated. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-29975. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | D6b "if explanted, give date" corrected. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella rp system with automated impella controller and impella rp flex with smart assist system with automated impella controller instructions for use and clinical reference manual. section: potential adverse events (united states): "arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia."
7875 | 1220648-2025-29976 | 22376042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376042&pc=OZD | 06/14/2025 | 12/02/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 182293 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Obstruction of Flow | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and, during support, the automated impella controller developed a purge issue requiring device replacement. the patient had experienced a cardiac arrest and was found to have multivessel coronary artery disease. an impella cp device was initially placed for mechanical circulatory support. due to ongoing cardiogenic shock and the need for continued hemodynamic support, the decision was made to escalate from impella cp to impella 5.5. while on impella 5.5 support, an air-in-purge alarm was triggered on the automated impella controller. the care team attempted to purge the line; however, semi-continuous purge alarms persisted. the impella 5.5 device was then connected to a second console, yet the purge alarms persisted despite the equipment change. the impella 5.5 pump was explanted. the patient¿s procedural outcome was unknown at the time of reporting. | [5 empty] | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported issue has been completed. data analysis: the case associated with the reported complaint on the (b)(6) was resumed on april 18, 2025, with the pump sn: (b)(6). note: this pump was moved from (b)(6), which was used from april 9, 2025, to april 18, 2025. during this period, there was no ¿air in purge system ¿ alarm.¿ on (b)(6), during support on the complaint date, june 14, 2025, the console displayed event #404: ¿air in purge system ¿ alarm.¿ the rt log shows that the purge pressure does not increase by more than deltaairpress (5 mmhg) during the first forward movement after the return stroke. the user performed two de-air procedures, and one cassette change procedures, but event #404 was not resolved. this aligns with the reported failure mode. note: there was no evidence in the cloud to confirm the claim that ¿staff switched impella to another console and the alarm continued to persist.¿ device analysis: this console was tested after arriving at fs. although it was unable to reproduce the ¿air in purge system¿ alarm by performing the cassette change and de-air procedure, the purge pressure accuracy was not within specifications. when applying an external pressure of 100 mmhg, the aic displayed only 86 mmhg, whereas the expected range was between 90 = pr = 125. root cause: the root cause of the ¿air in purge system ¿ alarm¿ was most likely an intermittent malfunction of the purge pressure sensor. h6 investigation type, findings and conclusions. | Corrections to the initial mfr report: updated h3 to device evaluation completed h6 component code should be 765 only.
7876 | 1220648-2025-29977 | 22376479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376479&pc=OZD | 06/14/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665940 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a large groin hematoma and access site bleeding that required medical intervention. the patient initially presented with an acute inferior st-elevation myocardial infarction (stemi) due to right coronary artery (rca) occlusion. upon arrival, the patient experienced a ventricular fibrillation arrest and was later found to have a ventricular septal defect (vsd). an impella cp device was implanted for hemodynamic support. during support, a large hematoma developed at the access site, along with active bleeding. a femstop was placed, the access angle was maintained, and perclose sutures were tightened to achieve hemostasis. due to the ongoing need for mechanical circulatory support, the impella cp was later explanted and replaced with an impella 5.5. the patient survived the event. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7877 | 1220648-2025-29978 | 22376484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376484&pc=OZD | 06/14/2025 | 12/03/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681850 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (hrpci) with mechanical circulatory support using an impella cp device developed access site bleeding and hemolysis complications requiring bedside device explantation. the patient, with a history of coronary artery disease, presented with an acute st-elevation myocardial infarction (stemi) and underwent hrpci supported by an impella cp device. during support, the patient developed bleeding at the impella access site, resulting in saturation of multiple dressings and requiring transfusion of two units of packed red blood cells. at the time of impella explantation, pink-tinged urine was observed, raising concern for possible hemolysis. the patient survived the explant procedure and was reported to be in stable condition following the event. | [5 empty] | The investigation has been completed. access site bleeding - major: clinical mentioned bleeding occurred overnight and hemostasis achieved by removing pump. not enough detail was provided to confirm root cause of bleeding onset. the root cause of the access site bleeding - major was not determined as not enough information related to failure was provided. hemolysis: the returned product was visually inspected and there was no biomaterial observed. a few spots of blood were found in the delo joint between cannula and cage. the pump was hemolysis tested and failed with an mih value greater than 20. the pump was torn down and no impeller abnormalities were observed. the data logs show no motor current spikes or position related alarms. there were suction alarms which occurred when p-level was reduced probably weaning but not during complaint period. the root cause of the hemolysis issue was unable to be definitively determined based on returned product and log analysis. limited clinical information was provided h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
7878 | 1220648-2025-29980 | 22376489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376489&pc=OZD | 06/14/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026673744 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A patient with a ventricular septal defect (vsd) underwent placement of an impella 5.5. bleeding was noted at the access site, and a diagnostic computed tomography scan demonstrated hemorrhage originating from the graft anastomosis. surgical suturing was performed, but bleeding persisted into the arm, necessitating further surgical intervention. the patient received multiple units of blood and antithrombin. the patient remained on impella support. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section a5 is unknown.
7879 | 1220648-2025-29981 | 22376492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376492&pc=OZD | 06/14/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629111 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed hemolysis and renal failure requiring medical intervention. the patient was admitted following a syncopal episode and chest pain, and upon arrival to the cardiac catheterization laboratory, angiography showed multivessel coronary artery disease including acute occlusion of the left anterior descending (lad) artery. an impella cp was placed via the right femoral artery. upon transfer to the icu, hematuria prompted evaluation; echocardiography confirmed the device positioned 4.5cm across the aortic valve, the physician addressed through repositioning. despite repositioning, the patient became anuric and required initiation of continuous renal replacement therapy (crrt). a decision was made to explant the impella cp and escalate support to an impella 5.5. the device was successfully removed, and the patient remained hemodynamically stable post-explant. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." | The investigation for the hemolysis and the renal failure has been completed. the pump was not returned for investigation. data logs were available which showed no abnormalities were reported regarding motor current spikes, placement signal trends, or positioning issues during the time of hemolysis. the cause of the hemolysis issue was not determined since the product was not returned and the data log was inconclusive. the cause of the renal failure was not determined due to insufficient clinical details.
7880 | 1220648-2025-29982 | 22376494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376494&pc=OZD | 06/14/2025 | 08/11/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665341 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed bleeding at the access site, which required medical intervention. the patient had a history of st-elevation myocardial infarction (stemi) and had presented with chest pain, nausea, and vomiting. the decision was made to implant an impella cp device for mechanical circulatory support. a mattress suture was placed by the physician to address minor groin site oozing. following clinical deterioration, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the instructions for use for the impella cp with smart assist system states the following: impella cp with smartassist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation)." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be use issue because hemostasis was achieved by placing mattress sutures at the site.
7881 | 1220648-2025-29983 | 22376532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376532&pc=OZD | 06/14/2025 | 08/01/2025 |  | 07/01/2025 |  |  | 06/14/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026701559 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with an acute myocardial infarction and cardiogenic shock had been implanted with an impella cp device for mechanical circulatory support. the patient had been involved in a motor vehicle accident, during which she developed hypoxemia and hypotension. due to ongoing cardiopulmonary decompensation, she was cannulated onto venoarterial extracorporeal membrane oxygenation (va-ecmo), and an impella cp device was placed for mechanical circulatory support. during placement, the impella cannula became kinked while attempting repositioning. despite multiple attempts, the issue could not be resolved. the device was explanted, and a new impella cp was inserted through the existing sheath. the patient survived the explant and was reported to be in stable condition following the event. | [5 empty] | The investigation of positioning issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issues was most likely patient condition related to the patient's short aortic arch. product damage cannula kink: the cause of the product damage cannula kink was most likely patient condition related to the patient's short aortic arch contributing to positioning issues where cannula kinking was observed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7882 | 1220648-2025-29984 | 22376534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376534&pc=OZD | 06/15/2025 | 07/01/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545608 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with a low ejection fraction (5%) and undergoing right heart catheterization was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed limb ischemia. the impella was removed as the impella leg was mottled without pulses. the doctor made the call to explant the impella. the patient's status post explant was noted as critical. | [5 empty] | The patient's race was listed as unknown. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
7883 | 1220648-2025-29985 | 22376541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376541&pc=OZD | 06/16/2025 | 07/08/2025 |  | 07/01/2025 |  |  | 06/16/2025 | 07/01/2025 |  |  |  |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1675422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with decompensated cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the bedside nurse did a purge cassette change using two different purge cassettes and neither were read by the automated impella controller (aic). the aic was exchanged for another aic without any issues. the patient was stable post procedure. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29985. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-29985. h6 type of investigation code 10, 4112,4121, 4110, investigation findings 4203, and investigation conclusions 4307 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29985. h6 type of investigation, investigation findings, and investigation conclusions codes were erroneously entered on the initial manufacturer device report number 1220648-2025-29985. h10 narrative erroneously entered stated that the device was not received on the initial manufacturer device report number 1220648-2025-29985. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29985. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7884 | 1220648-2025-29986 | 22376708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22376708&pc=OZD | 06/14/2025 | 12/11/2025 |  | 07/01/2025 |  |  | 06/15/2025 | 07/01/2025 |  |  | 09/30/2025 |  | 07/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393802 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a male patient of an unknown age and unknown ethnicity was undergoing an operation was implanted with an impella 5.5 with smartassist device for mechanical circulatory support. while on support the impella 5.5 started leaking from the reservoir. there the clinical staff had to change the dextrose 5% solution to a dextrose 10% solution to reduce purge flow and proceeded with basic fixing to the reservoir. the patient was moved to the cardiac care unit (ccu) later after the procedure. in the ccu, the impella went into a complete stop. a device restart was attempted without success. the device was removed due to the failure mode. the impella 5.5 pump was exchanged for an impella cp. | [5 empty] | Although the impella 5.5 was exchanged for a cp, the date of explant is blank as this information is unknown at this time. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of purge leak ¿ pump and pump stop has been completed. the pump was returned and evaluated. dried dextrose covered the outside of the sidearm retainer. the outer sidearm retainer was removed and purge testing was completed. the low purge pressure and high purge flow reproduced. there was clear leaking from the filter holes on the sidearm. the reservoir had dried dextrose on the exterior but there was no leaking seen inside/around the reservoir. there was also blood ingress found in the purge line past the motor housing. data logs were not available. purge leak - pump: the cause of the pump purge leak was most likely a damaged sidearm filter since the leak was reproduced with the returned product but the cause of the damage is unknown. pump stop: the cause of the pump stop was most likely blood ingress since there was blood ingress observed past the motor housing of the returned product. d9 device returned to manufacturer date added. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29986. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29986. | Additional information was received in the form of medical safety review and noted the following: adult patient requiring mechanical circulatory support (mcs) due to severe hemodynamic compromise. device involved: impella 5.5 (primary support), impella cp (lv unloading), va-ecmo (backup support). event description: impella 5.5 implanted for mcs support. post-initiation, a leak was observed at the purge sidearm; clinical team advised to continue support as no immediate risk was identified. device subsequently experienced sudden pump stop resulting in hemodynamic instability. patient was emergently taken to the or for device removal; clot was noted on the impella 5.5. va-ecmo initiated for primary support with impella cp for lv unloading. patient expired while on support, exact date of death unknown. device performance issues: purge sidearm leak (initial observation), unexpected pump stop during use, and clot formation on device upon removal. clinical outcome: patient death. the death complaint is attributed to impella 5.5 due to pump failure contributing to hemodynamic compromise. device malfunction (pump stop) is considered reportable; underlying clinical condition also a significant factor. after review of the case by medical safety, it was determined the event is a death type of reportable event. sections b1 (adverse event/product problem), b2 (outcomes attributed), h1 (type of reportable event) and h6 (health effect-clinical and impact codes and medical device problem codes) have been updated/corrected, repopulated accordingly. please note date of death is unknown and has been left blank. note: h6: component coding and investigation with coding remains unchanged.
7885 | 1220648-2025-29865 | 22353055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22353055&pc=OZD | 06/05/2025 | 06/28/2025 |  | 06/30/2025 |  |  | 06/06/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026676251 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 64 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 was found to have a small subarachnoid hemorrhage on computed tomography (ct) of the head without contrast. a ct angiogram of the head/neck were negative. argatroban was on hold. neurosurgery was following and the attending medical doctor does not believe findings are to be impella related. the following day, there were concerns for thrombocytopenia. argatroban might be restarted with conservative partial thromboplastin time goal. approximately a week later, advanced therapy was denied. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7886 | 1220648-2025-29866 | 22356461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22356461&pc=OZD | 06/02/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/02/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639396 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female | 71 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. after the impella cp was inserted, the catheter for the percutaneous coronary intervention (pci) was brought forward. the patient was having problems with breathing, lost consciousness and her blood pressure dropped. the doctors intubated the patient and initiated catecholamine therapy. the impella cp was left in place after the pci. there were no signs of a dysfunction of the impella device and there was not a sign that this was impella related. the patient survived the procedure. the patient's outcome was noted to be critical. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
7887 | 1220648-2025-29867 | 22356584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22356584&pc=OZD | 06/02/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/02/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667346 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 79 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced oozing from the access site. blood loss was noted and blood products were transfused. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was discarded, and the investigation has been completed. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
7888 | 1220648-2025-29869 | 22356669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22356669&pc=OZD | 06/10/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/10/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues. following impella cp placement into the left ventricle, the 0.18 guidewire was attempted to be removed by the physician. however, resistance was encountered and the catheter and wire began to buckle. the decision was then made by the physician to remove completely the impella cp and the wire was removed from the ventricle without complication. the decision was made to place an additional impella cp. there was no patient harm. | [5 empty] | The root cause of delivery issue is not determined because the the pump was inserted through guide wire into the left ventricle but the guide wire buckled while pulling out from the ventricle. there were no sufficient details that explains the cause of the buckling and the product was not returned for investigation. hence, the root cause is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
7889 | 1220648-2025-29868 | 22356691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22356691&pc=PYX | 06/02/2025 | 07/16/2025 |  | 06/30/2025 |  |  | 06/02/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661827 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. the complainant reported that the impella rp flex had high purge pressure in the operating room. purge blocked post manipulation of the heart. no kinks were found. therefore, the impella rp flex was removed and a new impella rp flex was implanted. the patients outcome was noted to be stable. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | High purge pressure: the returned pump was severed at the catheter and only the cannula was returned. biomaterial was observed by the impeller and purge gap. reproduction testing was unable to be performed. no data logs were returned for analysis. clinical mention onset of high purge pressure and purge blocked coincided with manipulation of the heart. the root cause of the high purge pressure was not determined as no logs were returned for analysis and insufficient product was returned for analysis. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
7890 | 1220648-2025-29870 | 22356837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22356837&pc=OZD | 06/10/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/10/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677257 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues while attempting to implant the impella 5.5. due to the impella 5.5 not crossing through axillary access into the aortic arch, the pump was explanted and replaced with an impella cp for hemodynamic support. the impella cp was implanted via the axillary access with success. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through patient anatomy despite multiple attempts.
7891 | 1220648-2025-29871 | 22356856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22356856&pc=OZD | 06/10/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/10/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675453 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 102 KG | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while attempting to implant an impella cp, extreme tortuosity was encountered causing the pump to kink and resulting in wire damage and damage to the impella optical sensor. the case was aborted. | [5 empty] | The impella pump and guidewire were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of product damage (cannula kink) and delivery issues has been completed. the impella device was returned for investigation and additional findings were added as failure modes. delivery issue : the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through vasculature due to extreme tortuosity. product damage (cannula kink) : returned product showed cannula had signs of excessive force with kink observed. the root cause of the product damage (cannula kink) issue was most likely patient condition related as evidenced by clinical note mentioning during insertion encountered extreme tortuosity causing pump to kink. optical signal issue: the returned product was inspected and a kink was observed on the catheter between 35cm and 40cm mark. at the kink location, light exited confirming a damaged fiber line and failed laser continuity test. the log pattern observed matches that of a damaged fiber line which aligns with clinical details of extreme tortuosity of patient anatomy causing pump to kink. the root cause of the optical signal issue was determined to be due to a damaged fiber line as evidenced by the light exiting the location of catheter kink where fiber was damaged and as evidenced by log pattern and returned product 5s parameters product damage (catheter kink): kink was observed on the catheter between 35cm and 40cm mark. the root cause of the product damage (catheter kink) issue was most likely patient condition related, during insertion encountered extreme tortuosity causing pump to kink. product damage (guidewire kink): kink was observed on the guidewire. the root cause of the product damage (guidewire damage - peripheral vessel) issue was most likely use related as a result of excessive force as evidenced by kink on the returned product and clinical information revealing extreme tortuosity causing pump to kink and wire to damage. product damage (sheath kink): no sheath was returned. clinical mentioned unable to advance impella past tortuosity kinking peel away sheath. the root cause of the product damage (sheath kink) was most likely patient condition related as evidenced by clinical mention of being unable to advance impella past tortuosity kinking peel away sheath.
7892 | 1220648-2025-29872 | 22357096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357096&pc=OZD | 06/02/2025 | 07/29/2025 |  | 06/30/2025 |  |  | 06/02/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631235 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that after a couple days, the flow rate began to fluctuate and spikes in motor current were observed. the pump was removed in response to the issue and a clot was discovered in the inlet cage of the pump. a new pump was placed for ongoing support. | [5 empty] | The investigation low pump flow and thrombus has been completed. the impella device was not returned for evaluation. data log reviewed and motor current (mc) spike observed followed by many suction alarms and drop in flow , more mc spike occurred after. resolved impella position in aorta alarms occurred. no placement signal issue seen. purge spike observed after last mc spike. clinical stated: flows change noticed and spikes on the mc and clot was caught in the inlet cage of the 5.5 and impella 5.5 replaced to a new 5.5. pump successfully. low pump flow: the cause of the low pump flow issue most likely was biomaterial ingestion. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system: section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.".
7893 | 1220648-2025-29875 | 22357250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357250&pc=OZD | 06/01/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/02/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658581 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory during support access site bleed prompted multiple units of blood replacement. the patient continued on support until bridged to transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
7894 | 1220648-2025-29876 | 22357396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357396&pc=OZD | 06/05/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/05/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665939 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | Sepsis | Tachycardia | Ventricular Fibrillation | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. overnight, the patient experienced runs of ventricular tachycardia (vt)/ventricular fibrillation (vf) when aminodarone was turned off so lidocaine/aminodarone were started. the next morning, the patient developed atrial tachycardic rhythm which caused drops in mean arterial pressures (maps)/pulsatility on the patient arterial line and placement signal was less than 60. the patient required high doses of levophed to sustain maps above 65. positioning of the impella was assessed with echocardiogram and showed that the inlet was close to the mitral valve. the impella 5.5 was pulled back from 48.5 centimeters to 47.5 centimeters. concerns of hemolysis occurred due to high plasma free hemoglobin. the p level was decreased from p8 to p6 with hopes of decreasing hemolysis. epipinephrine was preemptively added with decreasing p level. continuous renal replacement therapy (crrt) was started later that day. the next day, the patients hemolysis labs were stable and showed no worsening signs of hemolysis. the patient continued on crrt. days later there were concerns for infection and sepsis. the central line was replaced and the patient was put on a broad spectrum antibiotic. the patient experience vt again that required shocking. echocardiogram showed proper placement. the level was increased back to p-8 without ectopy. two days later, the patient remained on crrt. the patient was still on antibiotics for positive blood cultures and the impella 5.5 was weaned to p-6. the patient remained on support and was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7895 | 1220648-2025-29873 | 22357413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357413&pc=OZD | 06/11/2025 | 08/29/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548998 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported an automated impella controller (aic) (not in patient use) had an alarm, which was noted to be for suction/yellow alarm. the aic was also due to have preventive maintenance. | [5 empty] | Data analysis: after 22 days of support, from (b)(6) 2025, to (b)(6) 2025, suction alarms (event #14) were logged throughout the logs. this confirms the reported failure mode. additionally, on (b)(6) 2025, a placement signal not reliable (psnr) alarm (event #1048) was observed. the logs show that the ps inconsistently spiked to 3000 and then drifted into negative values. device analysis: this console was tested and the suction alarm could not be reproduced when running the pump simulator. the 5s parameter and the bulb power were within specifications according to the preventative maintenance procedure. root cause: the root cause of the suction alarm could not be determined due to the inability to reproduce it. there were no occurrences of a controller error alarm in the logs. the suction alarm was most likely misreported as a controller error alarm. optical signal issue failure mode added based on the investigation results. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7896 | 1220648-2025-29878 | 22357480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357480&pc=OZD | 06/11/2025 | 09/05/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627820 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Pumping Stopped | Battery Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a battery/power issue with an automated impella controller. there was no reported patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. console logs show that ic9401 was not used on the reported day of event (2025-06-11) however prior data shows that support with pump 570550 started around 9pm on (b)(6) 2025, with few battery low alarms (#23, battery level 50%), multiple suction alarms, placement signal low alarms, and high number of psnr events (but no alarms). there was also a pump stop (impella stopped due to motor current high), which self-resolved within 2 seconds and did not repeat. case ended on (b)(6) 2025. data logs from the console are consistent with transport between approximately 03:48 and 05:25 on (b)(6) 2025. values of battery remaining capacity and relative state of charge are consistent with ac disconnections and reconnections as expected during transport, and meet criteria for battery level low alarm, as console was not plugged into ac during transport and batteries were being discharged. after reconnecting to ac for the first time after arrival, battery did not charge enough, therefore when disconnected, the alarm recurred. a spike in pump motor current prior to pump stop is consistent with biomaterial ingestion. inspection of the console revealed broken corner of the rear housing of the aic, but also compromised backstrap cable (usb wire insulation cut by edge of rlm pcb). no defects of the ac cord were observed. console had been tested with regards to any possible ac power interruptions, but no anomalies were found, and device operated within specification. aic - a/c power issues: the cause of the issue was not determined since issue could not be reproduced. aic - damage before therapy: the cause of the issue was most likely use related. failure mode added based on investigation findings. aic - pump stop: the cause of the issue was not determined since product (pump) was not returned, and issue was not reproduced during console testing. failure mode added based on investigation findings.
7897 | 1220648-2025-29880 | 22357594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357594&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639678 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 149 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed a hematoma at the access site. manual pressure and femstop were applied. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7898 | 1220648-2025-29874 | 22357695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357695&pc=OZD | 06/02/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/02/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658987 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was bleeding at the access site. the bleed was contributing to the diagnosis of hypovolemic shock, the patient was treated by blood product delivery, stopping heparin, and evaluating the angle of the access. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7899 | 1220648-2025-29882 | 22357760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357760&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 74 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced suction events and low flow. the pump was found directed toward the mitral valve. the pump could not be repositioned despite multiple attempts. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. low flow - the root cause of the low pump flow was most likely improper positioning of the device since the pump was contacting the mitral valve. positioning issue - the root cause of the positioning issue was most likely patient movement since suction began after moving patient off the table.
7900 | 1220648-2025-29877 | 22357786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357786&pc=OZD | 06/10/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/10/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656794 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 114 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 had oozing at the access site. one unit of blood was transfused and heparin was withheld. additionally, the patient's labs were hemolyzed the patient was transferred to palliative care and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
7901 | 1220648-2025-29885 | 22357938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22357938&pc=OZD | 06/06/2025 | 09/02/2025 |  | 06/30/2025 |  |  | 06/06/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667261 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and cannulated for extracorporeal membrane oxygenation (ecmo). the patient was sedated and intubated. the following day, it was reported that the patient developed limb ischemia on the impella leg. there were no palpable feet pulses and livid colored feet. vascular surgery was consulted. however, a femoro-femoral bypass was not performed. the vascular surgeon checked the left leg and there were no signs of limb ischemia. the impella position and groin management were checked and observed without issues. however, approximately five days later, the impella cp device became caught in the papillary muscle and was unable to be repositioned. the impella cp was running on performance level p-4 to unload the left ventricle, unloading continued to function. additionally, it was noted that hemolysis was present. it was recommended that the impella cp device remain in place until the following week for ventricular septal defect (vsd) procedure as the physician noted a replacement was not possible due to the circumstances. as planned, the patient underwent the procedure. the vsd closure surgery was completed. the impella cp device was successfully weaned and explanted with a survived outcome. the patient was transferred to the unit for further observations. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. positioning issues: clinical details noted that pump pigtail was positioned in papillary muscles. pump was not attempted to be repositioned until pump was explanted. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided on how the pump came out of position. hemolysis: clinical details noted that pump pigtail was positioned in papillary muscle and hemolysis was present after positioning issue. pump was not repositioned and hemolysis continued, blood products were given but no details were provided on if hemolysis resolved before pump was explanted. the root cause of the hemolysis was most likely due to improper positioning based on clinical details that pump pigtail was in the papillary muscle and then hemolysis occurred.
7902 | 1220648-2025-29881 | 22358138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358138&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634556 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, there was complete loss of pulsatility with each lesion. due to lactate and dependency on levophed, the decision was made to leave the impella overnight to cool the patient off with potential explant the next day. the peel away was removed and repo was advanced. the device was secured in place. some bleeding was noted around repo. perclose was tightened and surgicel was applied to the site. good pulse present. during support, it was reported that the patient was stable. however, the patient had an unexpected arrythmia and went asystolic and expired on support. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7903 | 1220648-2025-29887 | 22358253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358253&pc=OZD | 06/11/2025 | 07/18/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606262 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the patient's leg became ischemic. angiogram was performed which showed superficial femoral artery dissection. the cp was removed and intra-aortic balloon pump (iabp) was inserted. patient had drop foot and may require below knee amputation. | It was reported that a patient implanted with an impella cp suffered an intracranial hemorrhage. impella support continued. |  |  |  |  | B5 update with additional information from the clinical site. d6b explant date added. h6 health effect - clinical code 1942, 3160, health effect - impact codes 4627, 4641, and component code 3038 added. limb ischemia: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. stroke/cva: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h6 health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2025-29887. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7904 | 1220648-2025-29879 | 22358276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358276&pc=OZD | 06/10/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/10/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676345 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Death | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced a cerebral vascular accident which was confirmed with a ct. subsequently, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of cva/stroke and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. limb ischemia based on investigation findings/review of the clinical data.
7905 | 1220648-2025-29883 | 22358312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358312&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585476 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male | 79 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that although an impella cp was successfully implanted in a patient there was no distal flow past the impella sheath. a fem-fem bypass with ante grade sheath was used to restore pedal pulses. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
7906 | 1220648-2025-29884 | 22358423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358423&pc=OZD | 04/30/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656710 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient implanted with an impella 5.5 endured hemolysis with a "possible" relationship to the impella device used: ¿baseline unknown for pfhgb, ldh and total bil. pfhgb peak is 245.2 on (b)(6) 2025, ldh peak was 653 ((b)(6) 2025) and peak total bilirubin is 1.8. haptoglobin was less than 30 on (b)(6) 2025. patient was given 1.5l fluid for low cvp. patient's labs were monitored daily. patient's impella settings went from p7 to p6 d/t hemolysis on (b)(6) 2025. on (b)(6) 2025, pfhgb was 105, patient's impella decreased from p6 to p5. patient received a heart transplant on (b)(6) 2025 (procedure completed (b)(6) 2025 morning).¿ the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
7907 | 1220648-2025-29891 | 22358654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358654&pc=OZD | 06/12/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673736 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump site had access bleed and hematoma which prompted the team to wean for explant and give a unit of red blood cells (rbc) and the pump was weaned after 41 hours of support. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
7908 | 1220648-2025-29889 | 22358880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358880&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 53 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the limb showed signs of ischemia and with ultrasound imaging, the flow was deemed "sluggish". the patient was also exhibiting signs of hemolysis which did not prompt pump explant and remains on for support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. limb ischemia: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. hemolysis: the data logs do not show any motor current spikes or changes in purge pressure. on the 11th there was some increases in placement signal (ps) and slight changes in motor current (mc) and flow, that correspond to the patient events. based on the data logs and clinical details, the root cause of the hemolysis was most likely due to the patients condition as volume given was said to have cleared the hemolysis present.
7909 | 1220648-2025-29892 | 22358955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22358955&pc=OZD | 06/07/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674381 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. when the impella cp was removed, no pulses or dopplerable were found for the patient. there was also a clot in common femoral artery (cfa) after the impella was removed. the patient was taken to vascular surgery, and a cut down was done to remove the clot. it was noted that there was also a circumferential cfa dissection. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
7910 | 1220648-2025-29890 | 22359056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359056&pc=PYX | 06/12/2025 | 07/08/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625125 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 74 KG | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella rp flex being attached to an automated impella controller (aic) for use in a patient admitted in with acute myocardial infarction/cardiogenic shock. when the pump was connected to the aic, the optical signal was not functioning appropriately so the pump was replaced. the impella rp flex was never inserted into the body, just connected to the aic. the device was still in plastic packaging when connected and had not been handled. care was withdrawn and the patient expired. | [5 empty] | H.6 health effect - clinical code 4582 was incorrectly entered on the initial manufacturer device report number 1220648-2025-29890. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7911 | 1220648-2025-29895 | 22359166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359166&pc=OZD | 06/14/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/14/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646058 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support an acute myocardial infarction patient. the pump was placed but there were issues with the limb and flow. the patient had no pulsatility and as such the flow was sluggish in the limb accessed by the pump. there were discussions to possibly escalate to the axillary impella 5.5 but this was not done. the pump also had signs of hemolysis. volume was given. nipride was started for afterload reduction. the pump was repositioned and ran at a lower p-level. the pump was weaned and explanted after 94 hours. the patient survived and the patients outcome is stable. | [5 empty] | The investigation for the limb ischemia and hemolysis has been completed. the clinical details state that the "pt was hemolyzing, urine red. running at lower p-level and giving volume. started on nipride last night for afterload reduction. starting to show some signs of recovery." the data logs show that there was a trend in the motor current, placement signal and flows from the 14th -15th. some suction seen on the 15th. there were no motor current spikes. based on the clinical details and data log analysis, the root cause of the hemolysis is most likely due to the patients condition. the root cause of the limb ischemia could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
7912 | 1220648-2025-29897 | 22359318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359318&pc=OZD | 06/12/2025 | 07/07/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145602 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) was showing a system self-check failed error. the aic was removed from service. there was no report of patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs from the reported day of event are consistent with complaint and show a communication issue with the pbm (power battery manager), evidenced by a controller error alarm (1023, loss of communication with pbm1) during the second bootup after console returned from preventative maintenance (pm). console was power-cycled. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. analysis of log data from preventive maintenance procedure revealed pbm1 communication issues occurring already, but without system self-check failure. during one of the tests, controller error alarm (1023, loss of communication with pbm1) presented twice upon console boot-up, but each time self-resolved promptly. after aic was power-cycled, there were no power or battery related alarms for the rest of the pm. after console was received and visual inspection of the pbm revealed contamination, which has impacted a neighboring rs232 communication channel (for pbm1). corrosion was present on the underside of the pbm, but not visible on top of pbm the cause of the aic issue was contamination.
7913 | 1220648-2025-29898 | 22359421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359421&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669159 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, suction alarms were noted. the pump was repositioned in the ventricle and observed on fluroscopy in good position. despite reposition and adequate volume, pump flows continued low. the pump was explanted and found to have biomaterial in the outlet of the pump. the procedure was successfully completed with another device and the patient's outcome is stable. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was returned for evaluation. low pump flow: returned cp pump visually inspected. biomaterial observed wrapped around impeller and purge gap. cannula kink observed at bending area. no broken blade or observed. data log reviewed: suction alarms occurred; motor current (mc) spike was observed prior to the drop in flow. no placement signal, positioning issues seen. clinical states, impella flows was low 0.8-1.2l/min, suction alarms noted. despite reposition and adequate volume pump flows continued low. when pump removed pigtail was flushed the physician observed ¿something in the outlet¿. the cause of the low pump flow issue was most likely biomaterial ingestion, since it observed during evaluation and mc spike seen pre-drop in flow by logs review. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. failure mode cannula kink was added since it was observed on the returned product. product damage (cannula kink): cannula kink at bending area observed during evaluation, unknown how and when it occurred. the cause of cannula kink cannot be determined due to insufficient clinical details. instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29898. b2 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29898. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29898. h6 health effect: clinical code 4440 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29898. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29898. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7914 | 1220648-2025-29888 | 22359629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359629&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670733 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed limb ischemia that requiring the placement of distal perfusion catheters. impella support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿.
7915 | 1220648-2025-29901 | 22359719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359719&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658527 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 92 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, after the doctor applied soft/soft clamp to the graft and impella catheter, the high purge pressure alarm was triggered. the doctor removed the clamp but the alarm did not resolve. the purge system blocked alarm was triggered and the decision was made to replace the impella 5.5. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. data logs were reviewed and real time (rt) log of full case showed the pump was able to prime and achieve stable purge pressure and purge flow values for approximately 30 minutes before the pump was inserted into the patient. purge flow ticks normal at that time. rt log at time of pump insertion and run showed the purge pressure increase of approximately five minutes with a drop in purge flows and elongation of the purge flow ticks. pump flows during run were observed to be partially and fully saturated. clinical details noted that after the pump was inserted, the physician applied a soft clamp to the graft and the impella catheter then "high purge pressure" alarms were triggered. the root cause of the high purge pressure was most likely due to biomaterial per rise in purge pressure with saturation of pump flows in data logs. however, it was not known if the biomaterial was caused by buildup or ingestion. a.4 weight was inadvertently omitted on initial manufacturer device report number 1220648-2025-29901. b.7 added information was added that was inadvertently omitted on initial manufacturer device report number 1220648-2025-29901. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7916 | 1220648-2025-29899 | 22359788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359788&pc=OZD | 06/12/2025 | 08/08/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672945 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 61 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the surgeon encountered removal issues. there was difficulty while attempting to pull back and remove pump at bedside in cardiovascular intensive care unit (cvicu). patient transferred to cardiovascular operating room (cvor) for pump removal with vascular surgeon¿s assistance and the pump was successfully removed. | [5 empty] | The investigation of removal issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the cause of the removal issue was most likely patient condition related due to the observed calcification of the vessels and the ballooning of the vessel. | H.2 correction was inadvertently omitted on manufacturer device report number 1220648-2025-29899-1 for including the patient age in field a.4. correction statement about age was also omitted from h.11 in manufacturer device report number 1220648-2025-29899-1. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7917 | 1220648-2025-29902 | 22359807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359807&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671896 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male | 67 KG | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient underwent a successful procedure with an implanted impella cp. upon explantation, the sheath fractured and vessel damage occurred prompting surgical repair. the patient survived and the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel and introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. during explant, impella cp was removed with companion sheath in the 14fr sheath. sheath fractured and brachial artery degloved with sheath removal. excessive force was stated to be a contributing factor. vascular damage - peripheral vessel: the cause of the vascular damage - peripheral vessel was most likely use related due to the returned image showing an abnormal peeling not on the score line and stated excessive force applied. introducer damage - mechanical: a returned image of the introducer showed abnormal cracking along the sheath that was not along the score lines. the cause of the introducer damage - mechanical was most likely use related due to the returned image showing an abnormal peeling not on the score line and stated excessive force applied. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29902.
7918 | 1220648-2025-29896 | 22359810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359810&pc=PYX | 06/10/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/10/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026640585 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient had severe pulmonic regurgitation during impella rp flex support. the patient did not seem to be getting benefit of flows demonstrated on console. as a result, the patient was removed from support. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7919 | 1220648-2025-29893 | 22359996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22359996&pc=OZD | 06/12/2025 | 10/02/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611523 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced suction alarms due to lack of volume. 3l had already been administered. it was recommended to reduce the p-level as far as necessary, administer volume and increase the p-level again. in addition, the puncture site was bleeding heavily, and the patient received several blood bags. the vascular surgeon circumcised the puncture site and thus stopped the bleeding. additionally, the patient had a pericardial effusion with 100 ml/h aspirate. they had already administered 5 blood bags and 5 ffp`s. ultimately the patient deteriorated cardiopulmonary overnight. the patient experienced a silent aspiration and became increasingly worse, which is why he was intubated. despite the impella cp and catecholamine administration, he continued to deteriorate, and the doctors discontinued the treatment, and the patient passed away. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding - major: the cause of the access site bleeding was most likely patient condition related per clinical details of patient's coagulopathy. low pump flow: the cause of the low pump flow could not be definitively determined as no product or log were returned for evaluation and limited clinical details were provided. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7920 | 1220648-2025-29900 | 22360072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22360072&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026680986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella 5.5 device for mechanical circulatory support. during the implant, the doctor was unable to advance the impella 5.5 due to patient¿s vascular anatomy. impella insertion had to be aborted after numerous attempts to advance catheter. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The root cause of delivery issue is determined to be patient condition because the pump cannot be advanced through patient vasculature because of tortuous vessels despite multiple attempts. it was reported that the catheter was kinked while attempting to insert the pump and insertion was aborted. d6a and d6b were erroneously entered on the initial manufacturer device report number 1220648-2025-29900.
7921 | 1220648-2025-29903 | 22360137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22360137&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674672 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella and exhibited a perforation of the left ventricle and a hemopneumothorax. the patient expired from the injury. | [5 empty] | The impella device was not received from the customer as it was not explanted from the deceased patient's body. the pump was turned off by the medical doctor. therefore, the investigation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system with smartassist for use during cardiogenic shock section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
7922 | 1220648-2025-29894 | 22360168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22360168&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670763 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was advanced with notable kink in the impella sheath providing more push than normal for insertion. the impella cp did eventually advance through kink without issue. single access percutaneous coronary intervention (pci) could not be completed due to the kink. the physician was fine and just gained lymphocyte function-associated antigen (lfa) access and went on to complete the procedure. impella explanted fine through kinked sheath without incident. | [5 empty] | The investigation of product damage (sheath kink) and difficulty inserting/removing pump through introducer has been completed. the impella device was not returned for evaluation. difficulty inserting/removing pump through introducer: the cause of the difficulty inserting/removing pump through introducer issue was most likely a sheath kink. product damage (sheath kink): the cause of the sheath kink could not be determined due to lack of clinical details and complaint device not returned for analysis. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 & cp impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during impella section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller v ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
7923 | 1220648-2025-29905 | 22360885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22360885&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026684784 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, during pci procedure, there was a distal wire perforation in the posterolateral branch (plb) of the heart which occurred after thrombectomy of the distal right coronary artery (rca). balloon tamponade for several minutes did not close perforation. two coils placed to close perforation, patient lost posterolateral artery (pla) and posterior descending artery (pda) branch. angiographically the plb perforation was closed. pericardium was filling with blood. right ventricle and left ventricle accessed for perforations but no obvious source. transesophageal echocardiography (tee), transthoracic echocardiogram (tte) and angiograms performed but were unable to locate and the patient expired in the cath lab. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. the impella cp pump was returned. no physical abnormalities or damages found with pump. the root cause of the ventricle perforation issue was not determined due to insufficient clinical information and unknown relation to impella. a.4 weight was added as it was inadvertently omitted on initial manufacturer device report number 1220648-2025-29905. b.7 information was added that was inadvertently omitted on initial manufacturer device report number 1220648-2025-29905. h.6 health effect - clinical code 2135 was entered incorrectly on initial manufacturer device report number 1220648-2025-29905. a new code was added.
7924 | 1220648-2025-29906 | 22361033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22361033&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674699 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 83 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced ventricular tachycardia (vt). heparin was administered. the impella cp appeared to be placed slightly deep when observed under echocardiogram. the patient also exhibited red urine and the pump was repositioned. two days later, the patient expired a run of vt again. aminodarone was down to 5. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7925 | 1220648-2025-29904 | 22361389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22361389&pc=OZD | 06/12/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663685 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Remove | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was product damage and removal issues. the 0.18-inch guidewire was molded normally, and the insertion of the wire went well. however, in the left ventricle, the wire was kinked in the transition of the floppy part to the stiff part. the implantation was in a normal way, but it was not possible to remove the wire through the pump. they removed the pump with the wire and the transition part of the wire was kinked in a 90-degree angle. by manipulating the pigtail, the team was able to remove the wire. they used a platinum plus wire and the further implantation was possible without any problems and went successful. | [5 empty] | The guidewire was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
7926 | 1220648-2025-29908 | 22361747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22361747&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667345 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during an attempt to implant an impella cp, the patient had a vertical left ventricle and the impella would not sit right. the implant was aborted. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was most likely patient anatomy as the patient had a vertical left ventricle (lv) that prevented the impella from obtaining the correct position.
7927 | 1220648-2025-29907 | 22361777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22361777&pc=OZD | 06/12/2025 | 08/29/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026677281 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Unspecified Infection | Gastrointestinal Hemorrhage | No Clinical Signs, Symptoms or Conditions | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was pump saturation. there was a change in the waveforms and equalization in max and min flows. it was recommended to explant the pump due to potential pump stop. the staff decided to continue support, the pump was turned to p 9 and flows equalized device the waveforms returned to normal and flows began to settle to an average of 5.3 max 6.0 min 4.7. it was still recommended to wean and explant, however support continued. | Additionally, the patient had an infection. cultures were positive and antibiotics were started. it is unknown if the impella contributed to the infection. furthermore, the patient had gastrointestinal bleed. blood products were given. moreover, high purge pressure was noted. tissue plasma activator was added to the purge. care was withdrawn and the patient expired. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. no product was returned. clinical details state there was a change in waveforms and equalization in max and min flows on the second day of support. pump was verified to be in correct position. after ~2 hours, pump waveforms returned to normal. data log review showed a motor current spike with a step up in motor current, speed deviation, and immediate saturated flows. no product was returned. the root cause of the saturated flow was most likely biomaterial ingestion since data log review showed a motor current spike and increase as the saturated flows began. the failure mode thrombus was added since data log review indicated biomaterial ingestion related to the saturated flow. the root cause of the thrombus was unable to be determined due to insufficient clinical details. clinical details stated that purge pressure was in the 1100s mmhg on (b)(6) 2025. they administered tpa and the purge pressure decreased back to the 700s mmhg. data log review showed a pre-disturbance where the purge pressure increased from ~600 to ~800 mmhg for about 11 hours before resolving. after that resolved, the purge pressure gradually rises from ~600 mmhg to >1100 mmhg over about 24 hours. the high purge pressure is sustained for ~8 hours before gradually resolving. there were multiple long bag/cassette changes throughout this time. although none of these occurred around the times that the purge pressure started to increase, they could have contributed to a continued rise. the root cause of the high purge pressure was most likely biomaterial buildup due to the gradual increase of purge pressure over ~24 hours. clinical details mention an infection treated by antibiotics during the case. product was not returned and logs are not applicable. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. clinical details mention a gi bleed during the case. the patient was on systemic heparin for part of the case. product was not returned and logs are not applicable. the cause of the vascular damage was not determined since insufficient clinical details were provided and no product was returned. | B1 revised to reflect both adverse event and product problem. b2 updated to reflect intervention, death, and date of death. b5 updated with additional information received from the clinical site. d6b added. h1 revised to reflect death. h6 updated with additional failure modes. h10 instructions for use added. the impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use. impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist for use during. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ e1 was erroneously selected on the initial manufacturer device report number 1220648-2025-29907.
7928 | 1220648-2025-29909 | 22363309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22363309&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675694 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 62 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported there was difficulty implanting an impella 5.5 was due to patient anatomy. also, a kink was observed in the catheter. a new catheter was used to implant the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (cannula kink) has been completed. the impella device was discarded. the root cause of the product damage (cannula kink) issue was most likely patient condition related with difficulty delivering pump due to difficult anatomy based on clinical information. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number: 1220648-2025-29909. d.4 revised unique identifier (udi)# as it was previously entered incorrectly on initial manufacturer device report number: 1220648-2025-29909. h.6 code 3036 was inadvertently omitted from the initial manufacturer device report number: 1220648-2025-29909 and was added. coded 4114 was previously entered incorrectly on initial manufacturer device report number: 1220648-2025-29909. code 4115 was added as the device was discarded. h.10 additional manufacturer narrative on initial manufacturer device report number 1220648-2025-29909 reported incorrectly that the device was not returned. device discarded should have been reported.
7929 | 1220648-2025-29910 | 22363426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22363426&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655567 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 96 KG | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted slightly deep in the patient. attempts to reposition the pump were unsuccessful. after implantation the patient began to decompensate. the patient was put on ecmo support to improve low impella flows. the patient was successfully weaned from the impella and survived. | It was noted postoperatively that there was bleeding concerns which were resolving. plans were made to initiate heparin. |  |  |  |  | A5, race, is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. | The investigation of positioning issues, access site bleeding, and low pump flow has been completed. the impella device was not returned for evaluation. the data log showed no signs related to motor current spikes or positioning issues during the case run. several suction alarms were noted throughout the case run without any observed motor current spikes or positioning alarms. the placement signal exhibited a trending/fluctuating pattern during the case run, and some intermittent drops in pump flow. according to the clinical notes, the patient had volume issues and was on venovenous extracorporeal membrane oxygenation support during pump use. the cause of the low pump flow issue was most likely patient condition, based on data log review and provided clinical details. the cause of the access site bleeding issue and positioning issue were not determined since the product was not returned and sufficient clinical information was not provided. b.5 information was inadvertently omitted on initial manufacturer device report number 1220648-2025-29910. h.6 health effect - clinical code 4582 was erroneously entered on initial manufacturer device report number 1220648-2025-29910. health effect - clinical code 1888 and health effect - impact code 4641 were inadvertently omitted on initial manufacturer device report number 1220648-2025-29910.
7930 | 1220648-2025-29912 | 22363476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22363476&pc=OZD | 06/12/2025 | 08/08/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 60 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient implanted with an impella cp experienced bleeding at the right femoral insertion site. the site was re-dressed with angle match dressing and suture was applied. the family withdrew care, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, site re-dressed with angle match dressing and oozing stopped.
7931 | 1220648-2025-29913 | 22363847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22363847&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the pump access site was oozing and was redressed, sutures were tightened, and blood was given for dropped hemoglobin. in addition, the pump had suction despite all troubleshooting. the family made the choice to withdraw care as prognosis was in question with possible encephalopathy and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data log reviewed: suctions alarms occurred after impella insertion above p-4, drop in flow observed at p4 & 5. looks flow resolved after blood infusion, no placement signal or positioning issue alarms seen. no mc spike outside of p-level changes occurred. the cause of low pump flow most likely was patient condition related since after blood infusion was able to increase flows. access site bleeding: the cause of bleeding (oozing) issue most likely was use related since retightened perclose sutures within syringe/stopcock and site redressed required.
7932 | 1220648-2025-29911 | 22364241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22364241&pc=OZD | 06/11/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/11/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664265 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male | 92 KG | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, there was bleeding from the esophageal from the endotracheal tube and suspected gastrointestinal bleed. heparin was held and the patient continued on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7933 | 1220648-2025-29914 | 22364494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22364494&pc=OZD | 06/12/2025 | 06/30/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662732 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 73 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 developed hematuria so heparin was withheld to resolve the issue. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿.
7934 | 1220648-2025-29915 | 22364521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22364521&pc=OZD | 06/12/2025 | 07/31/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633488 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 63 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon implantation of an impella cp the patient suddenly lost consciousness, had difficulty breathing, and her heart rate dropped to the 40s. the patient recovered. the patient experienced bleeding at the access site so one unit of red blood cells and cardene was transfused. because the patient experienced sinus tachycardia and atrial tachycardia an echocardiogram was performed, and the physician repositioned the pump from a deep setting. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients.". | The investigation of access site bleeding, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issue was not determined due to lack of clinical details. vt/vf: the root cause was unable to be determined due to insufficient clinical details. access site bleeding: there are multiple potential contributing factors to the bleeding such as bleeding at other sites and patient unable to keep leg straight. ultimately the clinical states that bleeding resolved once blue hub attached more to skin. the cause of the access site bleeding was most likely use related as bleeding resolved with re-adjusting the blue hub to the skin.
7935 | 1220648-2025-29916 | 22365405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22365405&pc=OZD | 06/13/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/13/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566581 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp was on day 15 of support when it stopped. the impella was able to restart. additionally, impella flow saturation was observed. the patient was eventually weaned and explanted. | [5 empty] | The investigation of temporary pump stop and saturated flow has been completed. the impella device was not returned for evaluation. temporary pump stop: data logs were reviewed and logs showed a motor current high pump stop occurred with pump being able to be restarted immediately after stopping. approximately 18 minutes before pump stop occurred, a motor current spike with speed deviation and full saturation of pump flows occurred. purge pressure also increased slightly with decrease in purge flows right after motor current spike. clinical details noted that temporary pump stop occurred on day 15 of support and pump was able to be restarted with saturated flows after stopping. the root cause of the temporary pump stop could not be definitively determined as no product was returned for evaluation to confirm cause of motor current spike pattern observed in data logs. saturated flows: data logs show saturated flows occurring immediately after motor current spike and speed deviation. pump flows remained saturated for remainder of case and pump weaning. clinical details noted that saturated flows were observed after pump was restarted. the root cause of the saturated flows could not be definitively determined as no product was returned for analysis and limited clinical details were provided. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7936 | 1220648-2025-29917 | 22365456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22365456&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/30/2025 |  |  | 06/12/2025 | 06/30/2025 |  |  |  |  | 06/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090625 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female | 59 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a ¿controller failure¿ alarm occurred while prepping an automated impella controller (aic) for patient use. the aic was turned off and turned back on and primed with new purge but the same alarm continued. the aic was exchanged and the associated pump was implanted without issue. | [5 empty] | The investigation has been completed. the console was booted up and began going through the case start but during step two, the ossi ran a 5s due to a bad lamp and found light to be at 0.7v. the 5s was measured and the lamp failed again triggering a controller error for defective optical lamp as the user moved to step three in case wizard. this is what the complaint was referencing as everything else aligns clinically with the description and the controller error was mistakenly called a controller failure. the console was returned and the failure mode was reproduced. it was found that the optical pump connector was heavily contaminated. after cleaning the connector, the console passed 5s. there was no controller failure found during the case. the controller error was caused by the contaminated optical connector. the contaminated connector absorbed the majority of the light so only a small amount of light was received back (~0.5v would be no light from the lamp received back so 0.7v indicates the lamp is sending out light). | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7937 | 1220648-2025-29837 | 22344287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344287&pc=OZD | 06/10/2025 | 10/02/2025 |  | 06/27/2025 |  |  | 06/10/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1413761 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller was returned from the customer. during functional check, the console was boot looping. later on while trouble shooting, the console coudl be power cycled normally 30 times. however, it was noticed that one unexpected shutdown occurred shortly before puting the console back on the shelf. there was no patient use reported. | [5 empty] | The investigation of automated impella controller (aic) - a/c power issues has been completed. the impella device and data logs were returned for evaluation. aic - a/c power issues: there is no log evidence indicating issues related to a/c power. aic - shutdown or restart issue: the root cause of the intermittent unexpected shutdown was not determined due to the inability to reproduce the issue. it could be attributed to either flash card corruption or a 4-in-1 malfunction. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7938 | 1220648-2025-29838 | 22344370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344370&pc=OZD | 06/10/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/10/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590768 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 113 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was on impella cp support and exhibited a bleed with midline shift. the impella was explanted and the patient was placed on comfort care until he expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7939 | 1220648-2025-29839 | 22344450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344450&pc=OZD | 06/11/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/11/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664096 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 63 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cpwas placed in a st-segment elevation myocardial infarction patient. after six days of support, there was an observation made of a gastrointestinai bleed. the bleed was prompting surgical intervention. the pump remained on supporting for nine more days despite the bleed. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7940 | 1220648-2025-29840 | 22344493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344493&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/27/2025 |  |  | 06/11/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp which was supported by an automated impella controller (aic) that exhibited continuous priming. the purge kept priming and could not stop so the aic was exchanged to resolve the issue. the patients outcome is stable. | [5 empty] | The console was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. device analysis: the console was returned for investigation but the failure mode was not reproduced and there was no evidence of any difficulty with de-airing from the console. there was no evidence in any logs of priming issues and the pump was not switched onto consoles. despite the correct pump number on this complaint, manufacturer device report 1220648-2025-30057 aligns more closely with the complaint description. the air in purge alarm was properly captured and eventually mitigated with de-airing procedure. the complaint was about manufacturer device report 1220648-2025-30057.
7941 | 1220648-2025-29841 | 22344576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344576&pc=OZD | 06/12/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/12/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627601 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 180 KG | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support and during support, the patient had red urine and p level was dropped. continuous renal replacement therapy was started and the concern for hemolysis was resolved with labs not hemolyzing. additionally, the patient went into ventricular fibrillation (vf) arrest during support. the patient was shocked four times. support was continuedand the patient was successfully weaned. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7942 | 1220648-2025-29842 | 22344705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344705&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/27/2025 |  |  | 06/12/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666187 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 78 KG | Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was being placed rather than an impella 5.5 that could not be used as the arm/axillary was to be used for cannulation of extracorporeal membrane oxygenation. the impella cp was placed femorally but the delivery was complicated and the .018 wire kinked. the team replaced the abiomed wire. further details provided stated that care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely use related since non abiomed product was used.
7943 | 1220648-2025-29843 | 22344841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22344841&pc=OZD | 06/12/2025 | 08/11/2025 |  | 06/27/2025 |  |  | 06/12/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666889 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male | 84 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that roughly two days into impella 5.5 support, the patient developed some bleeding at their access site. the bleed appeared after the patient was rolled for cleaning. the patient was given three units of blood over two days and the heparin dose was down titrated. roughly four days later, the patient required an evacuation and washout of an intra-abdominal hematoma. the patient was stable following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding : bleeding at access site since 4am following being rolled. the root cause of the access site bleeding issue was most likely patient movement, bleeding at access site since 4am when patient being rolled for cleaning . vascular damage : intra-abdominal hematoma. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided.
7944 | 1220648-2025-29844 | 22345203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22345203&pc=OZD | 06/01/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/01/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657712 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 111 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient's access site was a little oozy. one unit of red blood cells was given to the patient and support was continued. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7945 | 1220648-2025-29845 | 22345299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22345299&pc=OZD | 06/06/2025 | 09/17/2025 |  | 06/27/2025 |  |  | 06/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653832 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and had a chest wall hematoma on the pump side. a washout was performed and three units of packed red blood cells were transfused. additionally, a left ventricle thrombus and an embolic cerebral vascular accident were noted via computed tomography scans. impella support continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of access site bleeding, thrombus, and stroke/cva has been completed. the pump was returned and no damages or abnormalities were found. access site bleeding: clinical states right chest wall hematoma with a small amount of active contrast extravasation on impella 5.5 side. no active bleed found, chest wall hematoma evacuated. patient has coagulopathy. the root cause of the access site bleeding was most likely patient condition related due to the patient's coagulopathy. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details.
7946 | 1220648-2025-29846 | 22345461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22345461&pc=OZD | 06/06/2025 | 09/02/2025 |  | 06/27/2025 |  |  | 06/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601544 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female | 96 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was difficult as the guidewire would not advance into the graft. placement was achieved but was suboptimal. attempts to reposition the pump were unsuccessful due to catheter kinking. the pump was explanted and a new pump was implanted. no patient harm was reported. | [5 empty] | The investigation of positioning issues and product damage (cannula kink) has been completed. pump returned for investigation. no catheter or cannula kinks observed. multiple attempts made to reposition. pump ended up deep, in the papillary, and towards mitral. attempted repositioning multiple times with resistance. catheter kinked and decision made by implanter to explant and re-implant a new pump. patient had tortuosity. positioning issue: the cause of the positioning issue was most likely patient condition related due to patient's tortuosity. product damage catheter kink: the cause of the catheter kink was most likely patient condition related as it occurred during multiple repositioning of the device through patient's tortuosity. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7947 | 1220648-2025-29847 | 22345587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22345587&pc=OZD | 06/01/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/01/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620846 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 86 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, there was a hematoma at the access site. the heparin was stopped in the pump purge and sodium bicarbonate was used instead. the pump remains on for support and the hematoma is in control now. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7948 | 1220648-2025-29848 | 22345806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22345806&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/27/2025 |  |  | 06/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661706 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male | 112 KG | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced low pump flow. the issue was resolved by increasing the p level of the pump. the patient's labs were hemolyzed but the pump was not repositioned. care was withdrawn and the patient expired. | [5 empty] | The investigation of low pump flow, pump stop, positioning issues, and hemolysis has been completed. the impella device was not returned for evaluation. data log review: the data log shows a trending placement signal without any reported motor current spikes during the low pump flow issue while running on a steady p-level of 3. "pump position in aorta" alarms were present during the time of low pump flow. no major suction alarms were reported during the entire case run. according to the clinical notes, the pump was left shallow due to a small lv, and blood was given during the noted low pump flow; the pump was also providing ecmo support. low pump flow: the cause of the low pump flow issue was most likely related to the patient's condition, based on the data log review and the provided clinical information about the patient's condition related to small left ventricle (lv), low blood volume and extracorporeal membrane oxygenation (ecmo) support during pump use. positioning issue: the cause of the positioning issue was most likely patient condition related small lv. hemolysis: the cause of the hemolysis issue was most likely related to the patient's condition, as the patient experienced low blood volume and positioning issues according to the clinical details. the placement signal trend also confirms aspects of the patient¿s condition. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7949 | 1220648-2025-29849 | 22345992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22345992&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/27/2025 |  |  | 06/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 78 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 failed and could not be advanced into the axillary artery. instead, an impella cp was implanted. care was withdrawn and the patient expired. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the doctor tried multiple times to advance the 5.5 into the axillary artery. we tried multiple insertion tricks with no success. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
7950 | 1220648-2025-29850 | 22346352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22346352&pc=OZD | 06/16/2025 | 10/03/2025 |  | 06/27/2025 |  |  | 06/16/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581670 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 64 KG | Non Hispanic | Asian | Malfunction | No | Yes | Detachment of Device or Device Component | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was on impella cp support. it was noted that during transfer to the intensive care unit on a bed, tension was applied to the impella cp smartassist tip which caused it to slip out to the aorta. when tension was applied, the connection of the sterile sleeve was damaged. mechanical circulatory support was still needed so a new impella cp will be replaced. the patient survived. | A patient with ami/cgs was admitted. lesion for left main and underwent percutaneous coronary intervention (pci) with intra-aortic balloon pump (iabp) assistance. there was hemodynamic disruption during pci and extracorporeal membrane oxygenation (ecmo) was added. additionally, there was difficulty withdrawing from ecmo plus iabp, and the patient was transported from iabp on monday morning for the purpose of impella upgrade and an impella cp was placed. impella cp pump 1: during the transfer to the intensive care unit on a bed, tension was applied to the tip, causing it to slip out to the aorta. when tension was applied, the connection part of the sterile sleeve was damaged. the device was replaced (same day as implanted). impella cp pump 2 was implanted and 20 days after, the pump stopped. the pump was restarted and ran at low flow rate (p-2) for a while. per the physician it did not affect hemodynamics significantly. however, there was a possibility that the patient would still become unstable without mechanical circulatory support. the patient was escalated to an impella 5.5. the patient expired 28 days later on the impella 5.5 noted from multiple organ failure. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage) and positioning issues has been completed. the impella device was not returned for evaluation. product damage (sleeve damage): based on the clinical details, the root cause of the of the product damage (sleeve damage) is most likely due to use as there was tension placed on the device during transfer causing the anticontamination sleeve connection to become damaged. positioning issues: based on the clinical details provided, the root cause of the positioning issues is most likely due to patient movement as an additional load was placed on the cp during transfer causing it to become out of position. | Medical safety review of the event was completed. impella cp pump 1, the event describes a malpositioning of the device and anti-contamination sleeve damage. pump 1 was therefore replaced. this malfunction is not attributable to the demise outcome. however, it is a product malfunction and should be reported. impella cp pump 2 involves a malfunction with pump stop that was restarted with inadequate flow. hemodynamics were not affected significantly per the reporting from the physician. the pump was replaced. the patient expired approximately 28 days later. after being upgraded to the 5.5, an attempt to stabilize the patient's hemodynamics were made but there was little improvement; the patient died of multiple organ failure. in this case, the patient's underlying condition likely contributed most to an outcome of death (cardiogenic shock stage e). however, there is not enough evidence to exclude the impella cp pump 2 as an associated factor given the pump stop malfunction in this complex case. additionally, this device malfunction if occurring in another patient was repeated (as in the case above) could result in significant hemodynamic decompensation and therefore, should be reported as a death type of reportable event. impella 5.5: no complaint ¿, the upgraded to 5.5 and tried to stabilize his hemodynamics, but there was little improvement, and he died of multiple organ failure.¿ this event story involves two product complaints impella cp (b)(6) pump 1 (25-45645-1): malfunction (mdr (b)(4)). impella cp (b)(6) pump 2 (25-45618-1): death (mdr (b)(4)) b5 including patient outcome h10 medical safety officer review.
7951 | 1220648-2025-29851 | 22346666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22346666&pc=OZD | 06/18/2025 | 08/26/2025 |  | 06/27/2025 |  |  | 06/18/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623532 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed to support an acute myocardial infarction patient with ventricular fibrillation. the pump and extracorporeal membrane oxygenation were supporting but with signs of hemolysis, the impella cp was explanted after 28 hours. the thought was it was poor pump positioning that may have contributed to the hemolysis. when the impella cp was explanted, an intra-aortic balloon pump was placed. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: no product was returned for analysis. the data logs show 2 motor current spikes with speed deviations within an hour of implant and pump start. the motor current spikes had speed deviations and drops in flow characteristic of biomaterial ingestion. there were suction alarms. the root cause of the hemolysis was most likely biomaterial as evidenced by the motor current spikes in the data logs. thrombus: failure mode was added based on the investigational findings. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
7952 | 1220648-2025-29852 | 22346915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22346915&pc=OZD | 06/01/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/01/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482771 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 137 KG | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Brain Injury | Ectopic Heartbeat | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced sustained ventricular tachycardia (vt). the impella was repositioned, and the patient was given 500 cubic centimeters of albumin, which resolved the ventricular tachycardia. in addition, the p level was decreased due to the vt. they also noted that heparin was stopped due to bleeding. the patient was not following commands. however, the patient moved all four extremities. two units of red blood cells were administered and heparin was restarted. the impella 5.5 was pulled back one centimeter after ectopy. it was noted that the patient was off sedation. however, there was no neurological response. there was concern for coagulopathy and heparin was stopped again. a computed tomography scan showed anoxic brain injury. the impella 5.5 was repositioned under transesophageal echocardiography due to contact with the mitral apparatus. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7953 | 1220648-2025-29853 | 22347133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22347133&pc=OZD | 06/01/2025 | 07/18/2025 |  | 06/27/2025 |  |  | 06/01/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026654712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp for mechanical circulatory support. during the implant, the easy guide lumen came out and the pump was damaged while attempting to insert it. the implant was aborted and a new pump was used. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue : the root cause of the delivery issue was most likely use related, during the implant, the easyguide lumen came out and the pump was damaged while attempting to insert it. product damage ( cannula kink) : the root cause of the product damage issue was not determined since product was not returned and insufficient clinical details provided.
7954 | 1220648-2025-29854 | 22347200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22347200&pc=OZD | 05/06/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 05/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611381 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella 5.5 for hemodynamic support while awaiting heart transplant. during support the patient had hemolysis. altavase was running through pump with normal purge pressure and purge flow. the patient's free hemoglobin dropped. echo had good position. additionally, the patient tested positive for rhinovirus and enterovirus. the patient continues support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7955 | 1220648-2025-29855 | 22347230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22347230&pc=OZD | 06/02/2025 | 07/30/2025 |  | 06/27/2025 |  |  | 06/02/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657665 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, high purge pressure was noted 48 hours after the impella 5.5 was implanted. the next day, tissue plasminogen activator (tpa) was administered. and the device remained implanted for support. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. based on the clinical details provided that tissue plasminogen activator resolved the complication and data log review confirming the pattern of biomaterial build up, the root cause of the high purge pressure was determined to most likely be biomaterial build up. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
7956 | 1220648-2025-29856 | 22347467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22347467&pc=OZD | 06/01/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/01/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660156 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Flutter | Stroke/CVA | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the heparin induced thrombocytopenia (hit) was positive. heparin was switched to argatroban for anticoagulation. there was concern of existing thrombosis in multiple areas due to hit. additionally, the patient was taken for a head computed tomography (ct) showing a new infarct. it was further reported that the patient was having atrial flutter for the last 12 hours. the device remained for support and the patient was successfully weaned. | [5 empty] | The impella device was not returned, and the investigation has been completed. the root cause of the thrombocytopenia could not be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the atrial flutter was not determined. as this clinical information was not provided, the true root cause of the infarct could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
7957 | 1220648-2025-29857 | 22347740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22347740&pc=OZD | 06/06/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 71 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and had a small hematoma at the access site and hemolysis. a femstop was applied, pressure was held, and heparin was withheld until the site was stable and the p level of the pump was decreased. furthermore, the patient had ventricular tachycardia and was defibrillated. care was eventually withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7958 | 1220648-2025-29820 | 22347891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22347891&pc=OZD | 06/12/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/12/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546121 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 103 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had atrial fibrillation (af). the patient was cardioverted. the af occurred again and the patient was chemically converted. support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7959 | 1220648-2025-29858 | 22349225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22349225&pc=OZD | 06/01/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/01/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650850 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male | 93 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the patient remained on support for 54.35 hours and was noted that the impella cp was explanted due to causing hemolysis. the patients outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7960 | 1220648-2025-29859 | 22349492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22349492&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/27/2025 |  |  | 06/06/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653677 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female | 66 KG | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an increase in impella 5.5 purge pressure was noted when the anastomosis was being done and a side clamp was placed over the aorta. a purge system blocked alarm appeared and saturated flows were observed at this time. upon removal of the clamp, the purge pressure remained and changes were noted to the flow, motor current, and left ventricle signal. the decision was made to replace the pump as a result of the noted failure. the patients outcome is stable. | [5 empty] | The investigation of saturated flow and high purge pressure has been completed. the pump was not returned for investigation. the data log shows a motor current spike while running on a steady p-level, followed by a gradual rise in purge pressure with a high purge pressure alarm, at 40 minutes into the case run. the pump flow also became saturated during the noted motor current spike. the cause of the saturated pump flow issue was most likely biomaterial ingestion, based on data log review. the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review. thrombus failure mode was added as an investigated failure mode after the data log review and the conclusion of the high purge pressure issue. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿.
7961 | 1220648-2025-29860 | 22349696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22349696&pc=OZD | 06/02/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/02/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662365 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there were several pump stops due to high motor current peaks. the pump was restarted, however, there was a strong suspicion that a clot was ingested, and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7962 | 1220648-2025-29861 | 22350045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22350045&pc=OZD | 06/02/2025 | 07/29/2025 |  | 06/27/2025 |  |  | 06/02/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597870 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was oozing at groin site with noted blood-tinged urine. the pump was weaned to p 6 and heparin was withheld and integrilin drip was discontinued. the site was redressed. additionally, the patient had runs of ventricular tachycardia (vt) and lido drip was started. | [5 empty] | The investigation of access site bleeding, hemolysis, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of access site bleeding most likely was patient condition related since oozing blood from mouth and all access sites noticed. hemolysis: the cause of hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. h6 health effect - clinical code 2095 and health effect - impact code 4644 were inadvertently omitted on the initial manufacturer device report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
7963 | 1220648-2025-29862 | 22350288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22350288&pc=OZD | 06/02/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/02/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668106 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had ventricular tachycardia (vt) during manipulation at the aortic valve. the pump position was assessed and verified. the pump also had low flows due to missing preload. the patient expired when patient went asystolic as circulation was not able to be maintained. | [5 empty] | The impella device was discarded by the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7964 | 1220648-2025-29863 | 22350403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22350403&pc=OZD | 05/30/2025 | 06/27/2025 |  | 06/27/2025 |  |  | 06/02/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1754272 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Company representation | Distributor |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The distributor reported that during training the automated impella controller (aic) had a battery failure alarm. there was no patient involvement. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issue has been completed. the product was returned, and the issue was confirmed. the console logs were partially consistent with what was reported in the complaint. one instance of battery failure (transport connection switch blown/missing) alarm was observed in the imc logs but occurred on 28-jun-2025 and not on the reported complaint date of 30-jun-2025. the console was tested after arriving in field service, and the reported battery failure was confirmed. battery test results revealed that the cells in both batteries were all below 2500mv replacing the batteries with known good ones resolved the issue. the root cause of the battery issue was identified as depleted batteries, which resulted in a battery lockout.
7965 | 1220648-2025-29864 | 22350526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22350526&pc=OZD | 06/02/2025 | 07/18/2025 |  | 06/27/2025 |  |  | 06/02/2025 | 06/27/2025 |  |  |  |  | 06/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1206663 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | Altered State of Consciousness | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility¿s biomed reported that the impella automated impella controller (aic) has a purge sprocket potentially off plane. the aic was removed from service. there was no reported patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. data review: console logs were partially consistent with what was reported in the complaint. numerous consecutive imc-ahp errors were observed in a portion of the imc logs which is indicative of the purge cassette having issues being detected by the console. device analysis: console was tested after arriving. the purge motor sprocket was determined to be properly aligned after inspection of the console and a purge cassette was able to be inserted into the purge assembly without any issues while testing on an engineering lab bench. conclusion: after inspecting the console's purge assembly, it was confirmed that the purge motor sprocket was not 'off-plane'. the purge motor sprocket was determined to be properly aligned. the cause of the issue could not be determined due to the inability to reproduce. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
7966 | 1220648-2025-29799 | 22335189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22335189&pc=OZD | 06/06/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/07/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678111 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 145 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient developed oozing from the insertion site, and one unit of red blood cells was given. the oozing stopped the next day. prior to getting the patient, the pressor needs to be increased and white blood cell level was elevated. there were also concern for sepsis. the patient was started on antibiotics. the impella 5.5 continues to support the patient. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7967 | 1220648-2025-29800 | 22335284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22335284&pc=OZD | 06/07/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/07/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616555 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female | 62 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on the second day of support, there were platelets given for concern of thrombocytopenia/heparin-induced thrombocytopenia. the pump was then explanted on day five of support and the patient survived. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7968 | 1220648-2025-29801 | 22335375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22335375&pc=OZD | 06/07/2025 | 07/07/2025 |  | 06/26/2025 |  |  | 06/07/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male | 73 KG | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhagic Stroke | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient went into atrial fibrillation. the patient was unable to receive amino, so support continued, and staff monitored. the patient also had a hemorrhagic stroke which was confirmed by computed tomography. the patient remained comatose and was made do not resuscitate. the patient expired. | [5 empty] | The investigation of vf/vt/af/at and stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7969 | 1220648-2025-29802 | 22335694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22335694&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659949 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support for 16 days, there was persistent suction alarms. the team discussed troubleshooting measures and addition of anticoagulation systemically, but due to prior gastrointestinal bleed, the anticoagulation was not restarted systemically. the pump was successfully weaned and explanted after 19 days of support. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the impella device was returned for evaluation. low pump flow: returned pump had biomaterial present within outflow cage and around impeller blade. data logs were reviewed and the logs confirm reported low flows with suction alarms. the real time (rt) log at time of initial suction alarms showed step down in motor current and pump flows with speed deviation. additionally, decoupling of placement signal and left ventricle (lv) signal occurred at time of step down in motor current and speed deviation. the decoupling resolves. the root cause of the low pump flows was most likely due to biomaterial ingestion based on data log analysis, returned product, and clinical details. thrombus: based on the returned data log analysis, an additional failure mode of "thrombus" was added to the investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added.
7970 | 1220648-2025-29803 | 22335847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22335847&pc=PYX | 06/08/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026639425 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 120 KG | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. while on support, the purge flow began to drop steadily. tissue plasminogen activator was added to the purge to address the issue and support was continued uninterrupted. echocardiogram revealed large pericardial effusion and the patient rapidly decompensated. the patient was taken for wash out. however, upon exploration of the pericardium, several large thrombi were removed. there was no evidence of active bleeding or right ventricle wall rupture. the patient had a bypass and was eventually doing well on the impella rp flex and survived the procedure. | [5 empty] | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of thrombus, high purge pressure, and optical signal issue has been completed. the impella device was returned for evaluation. high purge pressure: the returned product was inspected and a major cannula kink was observed close to the inlet cage of the pump. no biomaterial was observed. the pump was run in the lab and high purge pressure did not reproduce. the data logs from the case show no motor current spikes outside p-level changes but shows a gradual rise in purge pressure for about 27 hours before high purge pressure alarm as triggered with a corresponding decrease in purge flow. high purge pressure was sustained for about 7 hours. tpa was added to the purge and purge pressure values returned to normal values. the root cause of the high purge pressure was determined to be due to biomaterial as evidenced by data log analysis revealing gradual rise in purge pressure and clinical mention of lysis therapy resolving the issue thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. optical signal issue: the returned pump passed the light continuity test and 5s parameters revealed low signal not reliable (snr) values. the pump was run in the lab and placement signal not reliable (bad snr) alarms triggered. the logs from the field show 5s parameter patterns consistent with a fiber bend or fiber break due to excessive pull with decrease to low snr, increase in gain, lamp, and contrast (with increase in amplitude). the root cause of the optical signal issue was most likely due to a damaged fiber line, possibly a fiber bend, as evidenced by log pattern observed and issue reproduction on returned product damage (cannula kink): added failure mode due to the cannula kink observed on pump. no information on how cannula kinked provided. the root cause of the product damage (cannula kink) was due to a cannula kink as evidenced by returned product analysis but it is unknown how kink occurred. instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ h3 was erroneously selected yes on the initial manufacturer device report number 1220648-2025-29803. h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29803. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29803.
7971 | 1220648-2025-29804 | 22336199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22336199&pc=OZD | 06/08/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656839 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male | 74 KG | Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Transient Ischemic Attack | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient had left side flaccidity and upward gaze. stat head computed tomography showed right cervical internal carotid had a small area of severe narrowing. neurology was consulted and recommended keeping mean arterial pressures higher, 70-100. similar transient ischemic attack like episode occurred three hours later. the patient also had a run of ventricular tachycardia (vt) followed by a 17 second run of ventricular fibrillation. implantable cardioverter-defibrillator was shocked once. the patient had vt requiring start of amiodarone drip and bolus. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7972 | 1220648-2025-29805 | 22336338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22336338&pc=OZD | 06/08/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665816 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp and showed flat waveforms. the physician repositioned the impella cp but the patient coded. despite obtaining return of spontaneous circulation by providing cardiopulmonary resuscitation, the patient later expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7973 | 1220648-2025-29806 | 22336514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22336514&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656757 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 87 KG |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that heparin was withheld from a patient implanted with an impella 5.5 due to continuous output from the patient's jackson-pratt drain and concern for bleeding from surgical site. additionally, while repositioning, the patient experienced ventricular tachycardia. the patient was defibrillated and survived. | [5 empty] | The investigation of vf/vt/af/at, access site bleeding, and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding: the cause of the access site bleeding was most likely patient condition related as the patient was also bleeding from other sites. positioning issues: the cause of the positioning issue was not determined due to lack of clinical details. vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7974 | 1220648-2025-29807 | 22336655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22336655&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677842 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and exhibited hemolysis and gastrointestinal bleeding. a heparin drip had been used during support. correct positioning of the pump was confirmed. the patient was successfully weaned from the pump and survived. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. hemolysis: based on the clinical details the cause of the hemolysis was most likely patient condition related. vascular damage (peripheral vessel): clinical details note that there was a gi bleed on the day of cp explant. sodium bicarbonate was used in the purge composition. no other details given. data logs are not applicable and product was not returned. the root cause of the vascular damage was not determined since insufficient clinical details were provided.
7975 | 1220648-2025-29808 | 22336960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22336960&pc=OZD | 06/09/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/09/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205394 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | While performing preventive maintenance on an automated impella controller battery issues were identified. new batteries were procured to resolve the issue. patient involvement was not reported. additionally, investigation findings found multiple instances of controller failure (purge pressure sensor defective) alarms were observed in the logs. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. the console logs were consistent with what was reported in the complaint. multiple instances of controller failure (purge pressure sensor defective) alarms were observed in the logs. in addition to the controller failure alarms, multiple battery failure (transport connection blown/missing) alarms were observed in the logs. the console was tested and the reported controller error was able to be reproduced during testing. troubleshoting of the controller failure by swapping out the purge pressure sensor with a known good one resolved the issue. results of running batttest revealed that all of the cells in both batteries had voltages of below 2400mv. the root cause of the reported controller failure was likely a defective purge pressure sensor. the root cause of the observed battery issue was identified as depleted batteries, which resulted in a battery lockout.
7976 | 1220648-2025-29809 | 22337057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22337057&pc=OZD | 06/09/2025 | 08/08/2025 |  | 06/26/2025 |  |  | 06/09/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218090 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller (aic) had a controller error. this issue occurred during patient support. no patient details were provided. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (yellow alarm) has been completed. on the complaint date, june 09, 2025, the console logs recorded only a continuous powermod_1 communication issue, with no 'controller error' alarm recorded in the logs. the downloaded logs did not show any record of a 'controller error' alarm, indicating that the logs may have been overwritten. this console was tested and during bootup, the console rebooted automatically and displayed the "system self check failure" screen. visual inspection confirmed power battery manager (pbm) contamination, which has impacted a neighboring rs232 communication channel. the root cause of the controller error was determined to be pbm corrosion resulting from electrolyte leakage caused by battery failure alarm super caps. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
7977 | 1220648-2025-29810 | 22337240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22337240&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/09/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666071 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and experienced low pump flow. the impella cp position was confirmed to be correct but a kink was noted proximal to the outflow cage. the pump was successfully replaced with a new impella cp but the patient ultimately expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. low flow - the root cause of the low pump flow was most likely a kinked cannula since a cannula kink was observed under x-ray. cannula kink - the root cause of the product damage (cannula kink) was not determined since product was not returned and insufficient clinical details provided.
7978 | 1220648-2025-29811 | 22337293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22337293&pc=OZD | 04/03/2025 | 09/19/2025 |  | 06/26/2025 |  |  | 04/03/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578898 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the patient ventricular tachycardia (vt) several times. the pump was repositioned and p level was decreased. additionally, there was minimal oozing at the start of support after turning the patient. the dressing was clean and intact. support continued. furthermore, after repositioning for vt, there were issues with the impella flow. the pump was repositioned further to prevent suction alarms. moreover, a month after support was completed, the patient returned to the er with pain at the access site. us shows potential foreign body at site. used perclose with sutures which were left in place, but no notes of issues during explant. staff was going to evaluate further. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Access site bleeding: the clinical details specify that minimal oozing was noted at the access site following turning the patient. the root cause of the access site bleed is most likely patient movement. positioning issue: the root cause of why the pump was in an inadequate position is unable to be determined due to insufficient clinical details. low pump flow: low pump flow was confirmed in the data logs with suction alarms being seen on the day of the event. the suction alarms resolve, which aligns with the clinical narrative indicating that the pump was repositioned to troubleshoot the low flow and was successful. the root cause of the low pump flow was most likely a positioning issue. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined.
7979 | 1220648-2025-29812 | 22337830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22337830&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/09/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662617 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | [11 empty] | Life Threatening | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impact study reported that the patient on impella 5.5 support had multiple runs of atrial fibrillation on (b)(6) 2025 that resolved within 24 hours with amiodarone bolus. the pi relatedness assessment as probable related to the impella procedure. the impact study also reported that on (b)(6) 2025, the patient experienced hemolysis. the hemolysis was resolved on (b)(6) 2025. abiomed safety relatedness: impella device possible related. pi relatedness impella device is probable related. patient was successfully weaned. | [5 empty] | The investigation of vf/vt/af/at and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29812. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29812. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7980 | 1220648-2025-29813 | 22337970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22337970&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/26/2025 |  |  | 06/09/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611994 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a nstemi patient with an impella 5.5 was made. during support, hemolysis was noted in urine. echo done and pump was turned toward mitral. plan to reposition at bedside. additionally, lv thrombus 1cm x 1cm was noted on tte with definity. decision on holding off on pump repo away from mv given thrombus concern. on another, echo showed pump to be shallow at 3.4 cm and faced towards the mitral. lv thrombus noted. consult surgery for possible repositioning, urine continued to be dark. additional information states that care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of positioning issues, thrombus, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29813. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-29813.
7981 | 1220648-2025-29815 | 22338549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22338549&pc=OZD | 06/13/2025 | 08/07/2025 |  | 06/26/2025 |  |  | 06/13/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589199 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 64 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 and experienced intermittent suction alarms and an episode of ventricular tachycardia, which improved after multiple blood products were transfused and the pump was repositioned. the cassette tubing was changed and primed to address "air in purge" alarms. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. low pump flow: the pump was not returned. data logs were not recovered. the cause of the low pump flow was most likely improper positioning of the device as the low pump flow resolved after adjusting position from 6.7cm to 4.6cm. positioning issues: implant depth was 5.2cm. after 1 day on support, pump was noted as too deep at 6.7cm causing flow issues with no listed cause for the change in positioning. one day later, pump repositioned to a depth of 4.6cm, flow issues resolved. the cause of the positioning issue was not determined due to insufficient clinical details. vf/vt/af/at: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
7982 | 1220648-2025-29816 | 22338686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22338686&pc=OZD | 05/07/2025 | 09/24/2025 |  | 06/26/2025 |  |  | 05/07/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563093 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during impella 5.5 support. roughly twelve days into support the patient experienced ventricular tachycardia. an amnio bolus was given to treat the condition. the patient survived. | Finally, after another fifteen days of support, a placement signal alarm appeared. the staff was walked through disabling the alarm and support continued uninterrupted. |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. optical signal issue: the placement signal (ps) was stable for the case. on the day of (b)(6) 2025, the ps is stable until the placement signal goes out to 3000, while the signal-to-noise ratio (snr) drops to 0, contrast barcodes, lamp increases, gain increases, and light decreases. this is indicative that the fiber line is most likely damaged indicating some stretch or break to it. the ps is out for the remaining days. based on data log analysis, the root cause of the optical signal issue is most likely due to a damaged fiber line. vt/vf cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem b5 h6 medical device problem code
7983 | 1220648-2025-29817 | 22338703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22338703&pc=OZD | 06/10/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/10/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636176 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 77 KG | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp was placed to support the patient admitted in with cardiogenic shock. the pump supported for just over four days when removed/escalated to the left ventricle assist device. the day prior the team had observed a hematoma that prompted washout and the limb was numb from possible nerve compression. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7984 | 1220648-2025-29818 | 22338742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22338742&pc=OZD | 05/08/2025 | 10/31/2025 |  | 06/26/2025 |  |  | 05/08/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657498 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 115 KG |  |  | Death | Yes | Yes | Increase in Pressure | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The patient presented with end-stage heart failure requiring inotropic support and admitted with volume overload, and severe lower extremity edema. patient continued to decompensate and the impella 5.5 device was successfully placed for mechanical circulatory support during work up for a left ventricular assist device (lvad) or orthotopic heart transplantation (oht). approximately five days after start of the support the patient experienced a nose bleed. heparin was suspended. mechanical circulatory support (mcs) continued, and heparin was restarted at one point. heparin protocol changed for a higher therapeutic goal. the patient was denied for oht. work up continued and the patient was optimized for a durable lvad. following, nine days later, rhino rocket was replaced to help nostril oozing. impella mcs continued. however the patient became septic and was placed on antibiotics and more pressor requirements noted on day 20 of mcs. support continued for the patient. approximately 26 days later, placement signal not reliable (psnr) on console was received in addition to the automated impella controller (aic) alarmed low purge flow. the purge flow dropped below three when the patient moved. however, the flow returned to normal. there was no report or indication of aic-to-aic exchange. the issue self resolved and mcs continued. approximately 14 days later the patient expired as care was withdrawn. of note, there were no further updates noted during this 14-day period. | It was reported that a patient implanted with an impella 5.5 had a nosebleed so heparin was withheld. the patient was also septic and was treated with antibiotics and pressors. the patient remained on impella support. |  |  |  |  | The demise outcome is more related to the patient's condition given the history which includes but not limited to an ejection fraction of 20-25% and exacerbated heart failure. the patient was denied for orthotopic heart transplant showing signs of regulatory volume decrease and plans were noted for possibly cannulation for veno-arterial extracorporeal membrane oxygenation (va ecmo) for additional support while working up for left ventricular assist device (lvad). however, there is not enough evidence to rule out, dissociate the impella 5.5 from the patient's outcome. it cannot be determined if the use of anticoagulant for use of the impella device impacted/exacerbated the out of the bleed for the patient. the patient became septic treated with antibiotic and it is unknown how, if any, the impella 5.5 pump played a role. regarding the automated impella controller (aic), patient expired for reasons unrelated to the aic. there was no interruption in support as the aic does not appear to have been exchanged or issue impacted the therapy. aic device remained in service. product malfunction for the aic are documented in manufacturer device report 1220648-2025-30089. b2 revised to include death and date of death. b5 revised to include the patient outcome. h1 type of report revised to death. h6 health effect - impact code revised to include code 1802 death h6 medical device problem codes 1491 and 2917 was inadvertently omitted on manufacturer device report number 1220648-2025-29818-1. | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The medical safety officer reviewed the case and agreed with the assessment provided in 1220648-2025-29818-2. | Vascular damage - peripheral: clinical details mention bleeding through the nose on (b)(6) 2025. anticoagulant use was changed throughout the case to attempt to manage the nose bleeding but also support impella. the root cause of the vascular damage was not determined since insufficient clinical details were provided. infection/sepsis: clinical details mention that the patient was septic on (b)(6) 2025 and on antibiotics and more pressor requirements. no other details known. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations: laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. optical signal issue: clinical details state a psnr alarm on (b)(6) 2025 which led to them silencing ps alarms. data log review showed ps going out of range and contrast oscillating intermittently starting on (b)(6) 2025. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to (b)(4) for an investigation into the console. high purge pressure: clinical details state that patient¿s aic alarmed low purge flow and it dropped below 3 after patient moved on (b)(6) 2025. flow returned to normal afterwards. data logs show a "purge flow decreased" alarm on (b)(6) 2025. the purge pressure slightly increased and purge flow decreased before a de-airing procedure was performed and the purge pressure/flow went back to normal levels. this all occurred within 5 minutes. there were no spikes or trends in motor current at this time. no product was returned. the root cause of the high purge pressure was most likely a kinked purge line as the clinical details report the low purge flow after the patient moved and data logs show the rise in pp, de-airing, and return to normal levels within a 5-minute period with no motor current spikes/trends.
7985 | 1220648-2025-29819 | 22338806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22338806&pc=OZD | 08/23/2024 | 06/26/2025 |  | 06/26/2025 |  |  | 06/10/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523082 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured pericardial effusion with a "possible" relationship to the impella cp device used: ¿pericardial effusion present on echo during percutaneous coronary intervention procedure and impella cp implantation. after second pea arrest, the subcostal area was re-prepped and pericardiocentesis was performed with drainage of around 800 ml of blood.¿. the pump was explanted and later, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7986 | 1220648-2025-29821 | 22338969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22338969&pc=OZD | 01/29/2025 | 09/26/2025 |  | 06/26/2025 |  |  | 01/29/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076068 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported the 62-year-old female patient was on impella cp support and during support, there were placement signal issues. staff monitored. the investigating engineer found that this issue was likely related to the automated impella controller (aic) and that there was a controller error on the aic. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the real time (rt) logs confirm this behavior as for the whole case placement signal goes to 3000 after re-calibration and raw optical sensor drops to -3276.8 and signal not reliable (snr) is 0. the low snr may have triggered the re-calibration of the console. device analysis: the console was found to have low snr with abnormalities in the ccd array with and without a pump. replacing the optical bench with a known good resolved the issue. root cause: the cause of the placement signal not reliable alarm was a defective optical bench. h3 no was erroneously entered on the initial manufacturer device report number 1220648-2025-29821. h10 investigation results was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29821. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates. h6 medical device problem code.
7987 | 1220648-2025-29822 | 22339516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22339516&pc=OZD | 03/12/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 03/12/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625120 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Abiomed received a long-term outcome and quality indicator (loqi) database adverse event notification of hemolysis with a patient on impella cp support. relationship is noted to be "probable" related. loqi states: "she also began to have dark red urine with evidence of hemolysis on labs in regard of impella management, patient had been having a lot of hemolysis. she did respond to afterload reduction with nitroprusside and she had improvement in cardiac output with night pride. the patient was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
7988 | 1220648-2025-29823 | 22339742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22339742&pc=OZD | 06/12/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/12/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639050 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella cp was repositioned under fluoroscopy, but flows did not go above p4. it was determined then to use extracorporeal membrane oxygenation. it was further reported that the physician explanted the impella cp due to poor perfusion to distal foot. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
7989 | 1220648-2025-29824 | 22340158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340158&pc=OZD | 01/29/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 01/29/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547614 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 was implanted into the left ventricle and turned toward mitral. attempts to reposition the device were unsuccessful. the pump was removed, rewired, and successfully reinserted. the patient exhibited bleeding at the anastomosis and the graft site was surgically repaired. care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
7990 | 1220648-2025-29826 | 22340471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340471&pc=OZD | 06/14/2024 | 06/26/2025 |  | 06/26/2025 |  |  | 06/14/2024 | 06/26/2025 |  |  | 03/31/2025 |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024320136 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient experienced multiple runs of atrial fibrillation during impella 5.5 support. an amio bolus was given to help manage the condition. upon eventual explant of the pump, bleeding complications were noted for which manual pressure was applied and sutures were utilized to achieve hemostasis. once stabilized, the site was flushed with antibiotic solution and the sternum was closed with sternal wires. roughly three days following impella explant, the patient's culture came back positive for burkholderia multivorans. the patient was given bactrim to treat the infection. the patient survived the procedure. | [5 empty] | The investigation for the reported atrial fibrillation (af), access site bleeding, and sepsis has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of af, access site bleeding, and sepsis could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
7991 | 1220648-2025-29825 | 22340474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340474&pc=OZD | 06/12/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/12/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666782 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female | 77 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient had ventricular tachycardia (vt). the patient was defibrillated once and converted back to sinus. support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7992 | 1220648-2025-29827 | 22340523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340523&pc=OZD | 06/05/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/05/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567964 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male | 91 KG | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced ventricular tachycardia that required cardioversion and placement of an implantable cardioverter-defibrillator. after five hours, the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
7993 | 1220648-2025-29828 | 22340565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340565&pc=OZD | 06/09/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/09/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546159 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a status 4 patient on home milrinone and dobutamine with worsening function, presented for impella for 5.5 insertion to stabilize. during support the patient had nose bleeds and suction alarms with p level down to 7 and tolerating. labs sent for tea colored urine. additionally the impella was repositioned due to intermittent ventricular tachycardia runs. the pump inserted one cm deeper and noted no issues with pump since repositioned. it was also noted that the patient had impella ¿placement signal low¿ and ¿suction¿ alarms while standing. the patient is stable and listed as status 2. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
7994 | 1220648-2025-29829 | 22340723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340723&pc=PYX | 06/05/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/05/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026635535 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 91 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | Malposition of Device | [5 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during implantation of an impella rp flex, there was difficulty advancing the pump. during support, the pump migrated multiple times and was repositioned multiple times at bedside. additionally, the patient's blood cultures were found to be positive so all lines were replaced. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp flex with smart assist system with automated impella controller section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
7995 | 1220648-2025-29830 | 22340883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340883&pc=OZD | 06/10/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/10/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026665336 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during the impella 5.5 support the patient had a transient ischemic attack (tia). upon impella removal, ct confirmed that the patient had a stroke. the patient was weaned and survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
7996 | 1220648-2025-29831 | 22340902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22340902&pc=OZD | 06/10/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/10/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Other | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The abiomed field rep found a publication out of impella users in (b)(6) germany. the publication speaks to 5.5 pump supports in surgical/coronary artery bypass graft patients. the retrospective review of 12 patients speak to adverse events. there was on average 6 day supports and in the supports there was bleeding, thrombosis, dislodgement/positioning failures, and 4 strokes. the strokes were thromboembolic. no additional information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1- a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
7997 | 1220648-2025-29833 | 22341483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22341483&pc=OZD | 06/10/2025 | 09/04/2025 |  | 06/26/2025 |  |  | 06/10/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026659787 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump placed to support the patient with cardiogenic shock. the pump failed to deliver appropriate flows and was removed due to the suction/low flows. when explanted, the limb had thrombus occlude the radialis artery and thrombectomy was performed. patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of low pump flow and thrombus has been completed. data was not returned for review. pump was returned cut at the catheter. visual inspection found no issues on the pump. dried blood around impeller. no biomaterial observed. unable to conduct the pump run test since the pump was damaged. clinical states pumps highest p level was 5, on higher levels suction alarms. asked for a thrombus in cannula, comment of explant: no thrombus was seen on cannula. by explanation of the pump the a. radialis occluded by a thrombus. thrombectomy successful. low flow: the root cause of low flow was unable to be determined as limited clinical details were provided and no data logs returned. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. d.9. device returned to manufacturer date added g1 manufacturing site name was erroneously entered on the initial manufacturer device report number 1220648-2025-29833.
7998 | 1220648-2025-29834 | 22341603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22341603&pc=OZD | 06/12/2025 | 06/26/2025 |  | 06/26/2025 |  |  | 06/12/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1587053 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) controller has a broken purge disc sensor, and the aic was removed from service. there was no reported patient involvement or harm. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. the imc logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined after arriving in fs. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
7999 | 1220648-2025-29835 | 22341624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22341624&pc=OZD | 06/03/2025 | 07/30/2025 |  | 06/26/2025 |  |  | 06/03/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603932 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient had an impella 5.5 pump implanted while awaiting heart transplant. optical signal and suction alarms were present. the automated impella controller was replaced and patient remained on impella support. | [5 empty] | The investigation of optical signal issue and low pump flow has been completed. the impella device and data logs were received for investigation. optical signal issue: data log review shows sensor drift begin 4 months into the case. the placement signa (ps) began to decrease and the user adjusted the ps multiple times. approximately 230 hours after the ps began to drift, the ps went out of range. there was slight decrease in signal not reliable (snr) and contrast during this time, but gain and light remained fairly stable. there were also suction alarms after the ps began decreasing. these alarms were triggered due to the low ps caused by sensor wear. the product was returned cut at the catheter above the motor housing. there were no visual abnormalities with the sensor, but neither the optical parameters or the pump itself could not be tested or run since the product was returned cut. the root cause of the optical signal issue was most likely sensor wear since the ps downward trend matches the known data log pattern and also occurred at 4 months of support, which is beyond the intended design life of 60 days per design traceability matrix. low pump flow: data log review confirmed suction alarms in february and march as reported. these alarms appeared to be triggering correctly. without exact timeline of patient movement and repositioning, the cause of these alarms is unclear. the suction alarms on (b)(6) 2025 and later occurred at the same time as the decreasing ps from the sensor wear. these suction alarms were triggered due to the low ps caused by sensor wear. the product was returned cut, so it could not be run in the lab. there was a small amount of biomaterial under the impeller and a kink in the middle of the cannula. the root cause of the low pump flow was not determined since there are insufficient clinical details regarding the timeline of patient movement and repositioning related to the suction alarms. product damage (cannula kink): the failure mode product damage (cannula kink) was added since a cannula kink was found on the returned product. it is unknown how or when the kink occurred. clinical details mention some repositioning throughout the case. data log review only shows suction alarms intermittently throughout the case. the kink was found on the returned product in the middle of the cannula. the root cause of the product damage (cannula kink) was not determined since insufficient clinical details provided regarding when/how the cannula kink occurred. d6b added d9 device returned to manufacturer date added h6 medical device problem code 2917 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29835. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8000 | 1220648-2025-29836 | 22341647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22341647&pc=OZD | 06/04/2025 | 08/01/2025 |  | 06/26/2025 |  |  | 06/04/2025 | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616465 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had bloody stool due to a gastrointestinal bleed so heparin was withheld. the patient remained on impella support. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. reported having blood stools due to a gastrointestinal (gi) bleed on (b)(6) 2025. the source of the bleed was clipped on (b)(6) 2025 but bleeding occurred again on (b)(6) 2025. impella 5.5 was returned and there was biomaterial around the impeller. no kinks noted or other damages / abnormalities. the root cause of the vascular damage gi bleed was not determined due to insufficient clinical details. d6b explant date added. d9 device returned to manufacturer date added. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
8001 | MW5172064 | 22348714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22348714&pc=OZD | 06/20/2025 | 06/24/2025 |  | 06/26/2025 |  |  |  | 06/26/2025 |  |  |  |  | 06/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Yes | [6 empty] | Health Professional |  |  | Voluntary report | Yes | Physician | [6 empty] | 46 YR |  | Female |  | Hispanic |  | Injury | No | Yes | Failure to Power Up | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Post myocardial infarction (mi), status post (s/p) coronary artery bypass grafting (cabg) with extracorporeal membrane oxygenation (ecmo) and impella 5.5. ecmo decannulated. patient had impella 5.5 malposition. repositioned appropriately. however, in order to remove slack form the cable to prevent dislodgement from current position cable had to be disconnected from console and reconnected. on reconnection with console the device did not restart. new console changed but device (impella 5.5) did not restart. had to remove this impella and place a new impella device. malfunction impella 5.5 device serial number (b)(6), abiomed j and j.
8002 | 1220648-2025-29761 | 22326693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22326693&pc=OZD | 06/06/2025 | 07/16/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629136 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 77 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella cp for support during a high risk cardiac procedure. the medical doctor attempted to micropuncture to the right groin at the start of the case and developed a hematoma. the vessel was tortous so the medical doctor elected to take the micropuncture sheath out and use the left groin for impella access. in the middle of the case, the patients blood pressure (bp) started to drop. the medical doctor noticed bleeding on the right groin which was not the impella side. levophed was started and blood was given as well as fluids. an 8 french catheter was placed through the peel away sheath to access the opposite groin and a stent was placed to the right iliac artery. no issues at all with impella access site were noted. the peel away sheath was removed at end of case and a perclose was used for hemostasis. the patients outcome is critical. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The root cause of access site bleeding was most likely patient condition since the vessel was tortuous per md.
8003 | 1220648-2025-29762 | 22326784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22326784&pc=OZD | 06/03/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/03/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026647344 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient who had escalated care from an intra-aortic balloon pump to an impella cp to this 5.5 pump and ecmo the pump was supporting when signs of hemolysis were present in elevated free hemoglobin and the patient suffered a brain bleed. family withdrew care and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
8004 | 1220648-2025-29763 | 22326810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22326810&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648160 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male | 100 KG |  |  | Malfunction | No | Yes | Increase in Pressure | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the purge system pressure suddenly increased, leading to the pump stopping. attempts to restart it were unsuccessful. the medical team subsequently decided to removed the impella cp with smartassist and initiate extracorporeal membrane oxygenation (ecmo) support. there was concern that perhaps the pump had thrombosed. the patient continued on ecmo support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8005 | 1220648-2025-29764 | 22326898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22326898&pc=OZD | 06/06/2025 | 07/18/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668044 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with a impella 5.5 for support during a cardiac high-risk procedure. it was noted that there was difficulty coming off of cardiopulmonary bypass when impella 5.5 was placed via direct aortic access utilizing trans-esophageal echo for optimal position. impella was started, however failed after approximately one minute. multiple aic (console) alarms show impella stopped. retrograde flow/impella stopped. motor current high/ impella stopped. controller failure. the console was exchanged prior to removing the device. the same alarms presented on the new console. it was then the device was explanted and a new device was opened, prepped, and implanted. normal functionality with the new device. surgery successfully completed. | [5 empty] | The investigation of pump stop has been completed. thrombus was added based on the investigation. pump stop: the pump was running for about 5 minutes and there were no issues with motor current or placement signal (ps). there was a then a sudden large motor current spikes followed by a high motor current pump stop. the pump was attempted to be restarted but would not. the returned product revealed a large chunk of hardened biological material under the bottom of the impeller. the biological material was removed and the pump ran without issues. based on the data logs and returned product, the root cause of the pump stop is most likely related to biomaterial. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8006 | 1220648-2025-29765 | 22326916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22326916&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669289 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that upon inserting impella cp, flows were unable to exceed 0.3 l/min. physician did not want to use another impella cp for the current case. impella cp was removed, inspected, and suspicion of clot was thought to be cause of persistent low flows. physician elected to flush impella cp and reuse impella. after second insertion, no issues during implant or starting of impella cp. patient survived. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8007 | 1220648-2025-29766 | 22326921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22326921&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1739661 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an automated impella controller (aic) had an issue opening the purge cassette door when the button was pressed to release the door. the staff had to gently pry it open. this issue occurred during support. there was no patient harm. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. console logs from the reported day of event do not contain any information relevant to this failure mode. the console was investigated. initially, the purge assembly door was sealed, and upon closer inspection of the purge assembly, purge fluid residue was found. this residue indicated a purge fluid leakage during clinical use. the root cause of the purge door inoperable i.e. not opening easily was due to the glucose ingress. the root cause of the automated impella controller - damage during therapy found no issue found during the case. the device functioned/operated to specification. a.1-a.5 are unknown.
8008 | 1220648-2025-29767 | 22327037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327037&pc=OZD | 06/10/2025 | 07/31/2025 |  | 06/25/2025 |  |  | 06/10/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630505 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male | 88 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and after the implant, the pump was deep towards the mitral/septal wall, and touching ventricle. the pump was ramped from p-2 to p-8 quickly. the patient went into ventricular tachycardia/ventricular fibrillation. the patient was defibrillated 10 times. upon repositioning and drop to p-4, arrhythmias stopped completely. the procedural outcome was the patient survived the surgery and was stable. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. positioning: based on the clinical details the cause of the positioning issue was most likely use related as the device was delivered deep and then required repositioning. vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8009 | 1220648-2025-29768 | 22327259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327259&pc=OZD | 06/10/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/10/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660559 |  | 0048-0014 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support and developed a hematoma at the impella access site. two units of red blood cells were administered. the procedural outcome was the patient survived the surgery and is stable. additional information was received. the bleeding is resolved and the vessel did not perforate. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8010 | 1035166-2025-00040 | 22327425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327425&pc=DYB | 06/14/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/17/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | S9370060 |  | 0052-3006 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Distributor |  |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | We have been notified of a complaint involving an introducer (material #0052-3006, lot #s9370060). at the beginning during an insertion of impella 5.5 procedure it was reported that the peel away sheath leaked/kinked during implantation. patient stable. there were no additional devices used during the procedure. no medical intervention or procedure delay reported. patient was presented to physician for cardiomyopathy, patient is stable. the issue was resolved by using another device of the same product. a new sheath was used and the impella was successfully placed. device discarded. note: this report contains a code for kink as the kink was reported in the event details; however, this type of defect is not a reportable event. we are only reporting the leak. | [5 empty] | No product was provided for analysis as product was discarded by customer. the customer stated that the peel away sheath leaked / kinked during implantation. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. the instructions for use (ifu): never advance or withdraw guidewire or sheath when resistance is met. determine cause by fluoroscopy and take remedial action. when injecting or aspirating through the sheath, use the sideport only. avoid subjecting the device to unusual stresses. when assembling the sheath and dilator, care must be taken to insert the dilator tip straight through the center of the valve membrane in order to prevent inadvertent puncturing of the membrane. based on the investigation, a capa is not required. in addition, there was no new failure mode identified, and the risk remains acceptable. the event will be re-evaluated if additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may have not been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
8011 | 1220648-2025-29769 | 22327436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327436&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  |  |  | 2026631347 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Delivered as Unsterile Product | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump being placed to support the patient admitted in with plan for coronary artery bypass graft. the team observed what they thought was condensation in the axillary insertion kit. there were concerns of sterility and so the team replaced and used another axillary kit. additionally, there was ventricular tachycardia (vt) present during the pump use on the day of insertion. later during the support, the heparin systemically was stopped due to bleeding concerns and subsequently there was urine color change, but hemolysis not definitively diagnosed. | [5 empty] | The investigation of packaging issues- sterility, vf/vt/af/at, and access site bleeding has been completed. only the 8fr dilator was returned in its opened original package. the returned dilator had coating particles on it. the presence of silicone oil on the 8fr impella dilator is expected per design, is typical to visualize, and is part of the qualified system. no problems were found during the case, as the presence of silicone oil on the 8fr dilators is anticipated according to design specifications. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. the cause of the access site bleeding issue was most likely patient's high act values, based on given clinical details. h6 health effect - clinical code 1888 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29769. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
8012 | 1220648-2025-29770 | 22327528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327528&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651099 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella cp support was not completely oriented. ct was completed and showed right mca infarct. it is unknown if the infarct was old or new. additionally, lactate dehydrogenase (ldh) was up to 2037. the team lowered the impella to p7 for possible hemolysis. it was further reported that the patient had positive blood culture (bc), unsure of source of infection. the patient remains status 7 on transplant list. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
8013 | 1220648-2025-29771 | 22327532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327532&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509884 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 51 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was implanted in a patient requiring support after cardiac pulmonary bypass. the impella 5.5 was inserted wireless and advanced in slow movements under transesophageal echocariogram. it was reported that on advancement of the impella 5.5 in the left ventricle (lv), the inlet of the impella 5.5 perforated the posteriolateral lv wall. the impella 5.5 was withdrawn, sutures were placed to repair the lv wall perforation, and the impella 5.5 was readvanced into the lv. additionally, during support, suction alarms occurred with rapid fluid administration and pressors started. the patient continued on support and remained hemodynamically stable for explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8014 | 1220648-2025-29772 | 22327683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327683&pc=PYX | 06/10/2025 | 08/07/2025 |  | 06/25/2025 |  |  | 06/10/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026668049 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp-rp flex experienced suction events in both pumps. the team tried to resolve the issue by infusing albumin. after 26 hours the rp flex was explanted and the cp was escalated to an impella 5.5 pump. the patient survived. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: clinical details note that rp flex was placed during cpr. patient continued to have suction on both rp flex and 5.5 and could not tolerate above p-5 on either device. a lot of volume was administered throughout the case. data log review showed no motor current spikes or purge trends. placement signal was trending downward near the end of the case on 11-june-2025. cvp was negative most of the case and was lower on higher p-levels. no product was returned. the root cause of the suction was most likely patient condition related since the patient could not tolerate higher p-levels, required multiple rounds of volume, and had a negative cvp based on data logs. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8015 | 1220648-2025-29773 | 22327720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327720&pc=OZD | 06/11/2025 | 10/27/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640954 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male | 65 KG |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The product evaluation revealed a kink in the cannula. | It was reported that implantation of an impella cp could not be completed due to the patient's anatomy. the pump could not pass beyond the iliac in the longer 14fr sheath. no patient harm was reported. |  |  |  |  | Delivery issues. the clinical details state that the impella was unable to the iliac artery below the bifurcation of the descending aorta due to calcification. based on the clinical details, the root cause of the delivery issues is most likely due to the patients condition as they had calcified vessels and could not pass the pump. product damage (cannula kink): this fm was added based on the investigation the product evaluation revealed a kink in the cannula. the root cause of the product damage (cannula kink) is most likely due to the patients condition as there was difficulty passing the bifurcation based on the calcification. b5 updated to include cannula kink issue description. d10 concomitant medical products and therapy dates updated. | A2 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8016 | 1220648-2025-29774 | 22327728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327728&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664682 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was implanted with steelcore wire, and it was turned on to p-2 while on bypass. the impella cp was deep at 4.5 centimeters and was pulled back to 3.3 centimeters. the impella cp got pulled out of the ventricle into the aorta. the impella cp was explanted. the impella cp was attempted to be reinserted over steelcore, but unsuccessful. the wire was kinked and was stuck in the aorta. the wire was not able to go through the pump. the impella cp and steelcore was explanted. the ventricle was accessed for the third time. the impella cp was advanced over the impella wire. once in the ventricle, the wire was removed. the impella cp was pushed forward and got pulled out into the aorta. the impella cp was explanted again. a kink in the catheter was discovered. impella cp placement was aborted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of positioning issues, delivery issues, and product damage (catheter kink) has been completed. positioning issue: data logs were reviewed and the reported complication was confirmed. initially, there was a position in aorta alarm that triggered when a distinct drop in motor current (mc) modulation occurred without a change in placement signal (ps). ps monitoring was then disabled. shortly after, the same pattern was noted in mc modulation and ps again, but since ps monitoring was disabled, no alarm triggered. this was likely the instance that was reported, because the pump was shut off and pulled out of the body shortly after. returned product was investigated and there were no visual damages/abnormalities related to the repositioning unit. that being said, the reported kink in the catheter was confirmed. the location of the kink was just proximal to the motor housing. clinical details did not report that the patient had a small/rotated heart or valve disease. clinical details provided that the pump pulled back into the aorta when the physician was manipulating position. data logs confirmed that the pump pulled back into the aorta while running. based on these factors, the root cause of the positioning issue was determined to most likely be a use issue. delivery issue: clinical details report that after the pump pulled back into the aorta the first time, the team decided to explant and reinsert the pump over the same steelcore guidewire. this delivery attempt failed because the guidewire had gotten kinked and was stuck in the aorta. they were unable to advance the pump over the kinked guidewire. the decision was made to remove both and use a new guidewire. it was also reported that the patient had tortuous vessels, which made it difficult to pass the pump through the anatomy. the guidewire was not returned for investigation, but a kink in the pump catheter just proximal to the motor housing was confirmed on returned product. based on the clinical details provided that they were unable to deliver the pump due to a kinked guidewire with tortuous patient anatomy, the root cause of the delivery issue was most likely patient condition. product damage (catheter kink): clinical details report that on the third attempt to position the pump, it was in the lv, and when they pushed forward, the pump pulled out into the aorta again. upon explant a kink was noted in the catheter. returned product was investigated and the kink was confirmed in the catheter, just proximal to the motor housing. it was also reported that the guidewire had become kinked earlier in the case, and the patient had tortuous anatomy. based on the clinical details provided that the patient had difficult anatomy and returned product confirmed the kink, the root cause of the catheter kink was determined to most likely be patient condition. d9 device returned to manufacturer date added h5 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29774.
8017 | 1220648-2025-29775 | 22327832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327832&pc=OZD | 06/11/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311572 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella automated impella controller (aic) did not recognize the purge disc upon start up. the aic was removed from service. no patient harm was reported. | [5 empty] | Patient information was not provided. section a was left blank. a review of the device's data logs revealed that the console issued purge disc not detected alarm several times. the returned console was examined. the purge disc flag was observed to be broken and was determined to be the root cause of the inability to detect the purge disc alarm.
8018 | 1220648-2025-29776 | 22327871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327871&pc=OZD | 06/11/2025 | 08/05/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1182670 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A biomed engineer reported an automated impella controller (aic) had a frozen screen. the issue did not occur during patient management. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. data analysis: we were unable to identify in device logs any issues that would be consistent with report of ¿console with a frozen screen.¿ device analysis: during testing the console operated within specification and attempts to reproduce the described issue were unsuccessful. there was significant contamination of the optical pump connector, as well as the rotary push button. the optical 5s parameters were obtained and were within specification. no distortion of the ccd output was noted.
8019 | 1220648-2025-29777 | 22327961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22327961&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672723 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 75 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was placed for a patient admitted in for a high risk cardiac procedure. the pump was placed and allowed for the percutaneous intervention. the pump was running when there was a pump stop and restart. the pump was restarted successfully. the pump supported but then there was saturated flow. the flow saturation was thought to be from biomaterial ingestion. the pump was explanted after just an hour and the patient survived. | [5 empty] | The investigation of pump stop and saturated flow has been completed. the impella device was not returned for evaluation. the data logs show that the motor current was stable for 47 minutes until there was a spike in motor current followed by a high motor current pump stop. the pump was able to be restarted but saturated flows were seen after. the pump was explanted. based on the data log analysis and the clinical description, the root cause of the pump stop is most likely due to biomaterial. the root cause of the saturated flow is most likely due to biomaterial. thrombus was added based on the investigation. thrombus: e root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8020 | 1220648-2025-29778 | 22328112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22328112&pc=OZD | 06/10/2025 | 08/07/2025 |  | 06/25/2025 |  |  | 06/10/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670896 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male | 74 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and noted that they encountered high purge pressure. the pump had a high purge pressure alarm, with subsequent increase in motor currents starting. the doctor wanted to remove the pump before it stopped. the pump was weaned to p-2 and then removed. manual pressure was held for one hour. the procedural outcome was the patient survived the surgery and the patients outcome is critical. | [5 empty] | The investigation of high purge pressure (hpp) has been completed. the impella device was returned for evaluation. the data logs show that the motor current and purge pressure was stable until about 19 hours in when we see a large spike in motor current followed by a dip in ps, an increase in pump flows and a rapid rise in purge pressure ~3minutes. the product was returned and biomaterial was found wrapped around the impeller and purge gap. the pump was soaked to remove dried blood but biomaterial remained in the gap. the pump was purged and hpp reproduced. a window in the catheter was cut and no ingress was found. the catheter was cut near the motor housing and hpp resolved. the root cause of the high purge pressure is most likely due to ingested biomaterial based on the data logs and returned product. thrombus: thrombus was added based on the investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8021 | 1220648-2025-29779 | 22328280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22328280&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/25/2025 |  |  | 06/09/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601497 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 107 KG | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 placement, it was difficult crossing the innominate artery with 0.025 plat plus wire. the team performed a left axillary cutdown and implant but encountered the same problem. the procedure was aborted and an impella cp implant through initial right axillary cutdown will be done. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because pump was unable to pass through innominate either from left and right axillary artery. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8022 | 1220648-2025-29780 | 22328359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22328359&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/25/2025 |  |  | 06/09/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615877 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was admitted in for a high risk cardiac procedure. the team was able to successfully place the 14 french sheath but the elderly patient's native anatomy precluded the placement despite multiple attempts and techniques. the procedure was aborted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
8023 | 1220648-2025-29781 | 22328528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22328528&pc=OZD | 06/09/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/09/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male | 81 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient developed a large hematoma to the inner and lower thigh due to having to do a second stick while implanting. manual pressure was held and heparin was withheld. the hematoma continued to extend and the pump was explanted. the patients outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8024 | 1220648-2025-29782 | 22329244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22329244&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/25/2025 |  |  | 06/09/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651725 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received from the clinical site the impella cp had a pump stop and was explanted. patient outcome is unknown. | The user facility reported a patient was on an impella cp for support. during impella cp support, it was noted that the patient developed a hematoma to the right groin. it was noted that the plan was to take the patient to the operating room for cutdown exploration by vascular surgery. six days later, two units of platelets were transfused. the patient continues on support. |  |  |  |  | The investigation of thrombocytopenia, access site bleeding, and pump stop has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. pump stop: the data logs show motor current spikes with speed deviations after which there was a shift in pump flow as well as a decrease in purge flow. purge pressure values also elevated after motor current spikes. minutes later there were 'impella stopped - motor current high' alarms. the log pattern identified could be as a result of biomaterial interaction or bearing wear. the root cause of the pump stop was unable to be determined as no product was returned for analysis and limited clinical information related to failure was provided. b1 updated to reflect both adverse event and product problem based on the additional information received from the clinical site. b5 updated with additional information received from the clinical site. d6b added h6 medical device problem code added based on the additional information received from the clinical site. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped, ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The impella device has not been received from the customer as it continues to support the patient. . therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8025 | 1220648-2025-29783 | 22329436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22329436&pc=OZD | 06/09/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/09/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653698 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had access site bleeding. two units of packed red blood cells were given to the patient in the operating room. the plan was to monitor hemoglobin, hematocrit and platelets and for signs of bleeding. hemogloblin was to be repeated post transfusion. angiomax was held until the jp drainage slows. support was continued. additionally, an aortic thrombus was noted at the level of the diaphragm. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8026 | 1220648-2025-29784 | 22329525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22329525&pc=OZD | 06/09/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/09/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678079 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella cp was placed using best practices and left on auto. the patient had to get shocked after the wire and impella were in the ventricle. the impella was left shallow to avoid any further arrhythmias. the patients outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8027 | 1220648-2025-29785 | 22329950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22329950&pc=OZD | 06/10/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/10/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there were issues with the automated impella controller (aic). per the biomedical engineer, the cartridge door of the aic was not opening and required use of forceps to forcefully open. the aic was switched out and there was no patient harm. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. the console logs did not contain any data relevant to this complaint. console was tested after and the reported stuck purge door issue was able to be reproduced. further inspection of the console revealed that there was solidified glucose buildup on top of the purge door. the root cause of the inability to open the purge door was the observed solidified glucose buildup on top of the purge door. this buildup caused the purge door to be momentarily stuck to the front housing of the console. the root cause of the automated impella controller (aic) - damage during therapy was there was no damage observed on the aic. the device functioned/operated to specification.
8028 | 1220648-2025-29786 | 22330197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330197&pc=OZD | 06/10/2025 | 08/11/2025 |  | 06/25/2025 |  |  | 06/10/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607702 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a pump stop while on impella cp. after the patient was brought to operating room for scheduled coronary artery bypass graft, the device failure occurred. restarting was attempted three times unsuccessfully. a new console was brought to the room and the pump was unable to be restarted on new console. the device was removed into the descending aorta and removed from the patient over the wire and the 14 french sheath was inserted. the patient's outcome is unknown. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. data log review showed pump stop with three attempts to restart all failing. in the last eight hours of the case, purge trends occurred three separate times in which the purge pressure would spike up and the purge flows would trend down. the first two occurrences trending toward recovery within two hours but the third ended in a pump stop. the pump was returned and inspected. after soaking the pump in warm water and tergazyme, biomaterial was found wrapped around the impellor blades and around the purge gap. no biomaterial was found in the purge tubing nor the red pump plug. electrical testing showed open cable lines between motor phases 1, with connection 1-2 at 72.2k ohms and 1-3 at 74.9k ohms. cables coming into and out of the red pump plug appeared twisted and one of the motor cables going into the catheter was necked and frayed. the cause of the pump stop was an open electrical line found on one of the motor cable lines that run from the red pump plug into the patient cable.
8029 | 1220648-2025-29787 | 22330479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330479&pc=PYX | 03/22/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  | 11/30/2025 |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024425642 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 94 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right saphenofemoral artery occlusion with a " probable" relationship to the impella device used: it was reported as ¿decreased temperature, increased tension, and decreased pulses were appreciated in the lower right extremity. the patient underwent ultrasound and computed tomography (ct) imaging of the right leg. ultrasound revealed arterial occlusion of the superficial femoral and popliteal arteries. ct revealed that the lumen of the right external iliac and proximal femoral artery was completely filled by the impella cp catheter (placed at outside facility). vascular surgery assessed the leg and determined that the limb was unsalvageable and would otherwise need to be amputated, however, given the patient's state and concern for developing worsening metabolic problems, the patient was too unstable to undergo surgery. vascular surgery also explained that the patient likely had chronic right superior femoral artery (sfa) occlusion which was exacerbated by occlusion of the external iliac from the impella cp placement.¿ ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
8030 | 1220648-2025-29789 | 22330572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330572&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657725 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male | 100 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, there was oozing around the jp drain site. the staff stated that they may add a stitch, but it was never confirmed. additionally, the patient was ambulating, and the pump was repositioned, which triggered the pump in aorta alarm due to new position. the pump was attempted to be repositioned without success and the pump was explanted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8031 | 1220648-2025-29788 | 22330573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330573&pc=OZD | 03/07/2025 | 07/31/2025 |  | 06/25/2025 |  |  | 05/28/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302148 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 68 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while the patient was implanted with an impella cp medical staff pushed the automated impella controller (aic) and the patient in a bed over a bump. subsequently, the impella alarmed and stopped. the perfusionist dropped the p level and the impella restarted. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - pump stop has been completed. the aic was received for investigation. at 23:10:03, the user opens up the menu for p-level and chooses p-0 at 23:10:12, which triggers impella stopped alarm at 23:10:22. p-level is set to p-5 at 23:10:48, and pump starts back up resolving pump stopped alarm and delivering hemodynamic support. the real time (rt) logs at this time confirm the user manually stopped the pump. at 11:39:30 on (b)(6), the user again opens up the menu and sets p level to 0 triggering a pump stop. the console is then put into surgical mode and attempted to be restarted but there is high motor current resulting in a pump stop. this happens a total of 8 times over the next 3 minutes before the console is left at p-0 for the next hour. the console is then set to p-1 and the pump successfully starts back up. finally, on 17:24:11 on (b)(6), the log entry "pump motor driver (pmd)" ok indicates that the pmd is reinitialized (though there is not a log entry for what the impetus of this was). this is followed by eep communication reinitializing. there is an event 115 impella stopped alarm which triggers because the pump speed goes to 0 despite p-level of 7 as the impellatronic board is reset. the pmd completes reset and the pump is started back up returning hemodynamic support. the case is completed and the pump is disconnected on (b)(6). device analysis: during testing, the console was ran continuously for 3 days with no interruption in pump performance. console could be manually set to p-0 mimicking the log entries from the (b)(6) and (b)(6) that triggered the event 18 impella stopped alarms but did not have difficulty restarting as occurred on (b)(6). connections between the pbm and impellatronic board were confirmed to be properly connected and shaken during run (to reproduce going over bump on (b)(6)) with no effect on console performance. failures were not reproduced during testing. root cause: the cause of the aic to not be able to restart on (b)(6) after the console was set to p-0 is undetermined as the failure mode was not reproduced and the logs did not indicate what was preventing restarting the pump. b3. date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29788. d4. serial number was erroneously entered on the initial manufacturer device report number 1220648-2025-29788. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8032 | 1220648-2025-29790 | 22330757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330757&pc=OZD | 06/11/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667213 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there were concerns of bleeding at the access site despite use of femostop and best practices, as well as there were signs of hemolysis. the pump was removed due to these reasons after just 16.5 hours. prior to the explant the purge cassette also failed, but caused no harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8033 | 1220648-2025-29791 | 22330759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330759&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673611 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced pulseless electrical activity arrest and cardiopulmonary resuscitation was in progress as the impella cp was being placed. it was noted that the 14 french peeled away with continuous oozing at the site. stitches were placed and pressure applied. the patient was stable, and the plan was to start continuous renal replacement therapy (crrt). it was reported that the patient was having hemolysis. it was advised to reposition the pump since it measured at 4.7 centimeters. however, the decision was made to remove the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system - section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter - ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle - ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis - ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction - ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8034 | 1220648-2025-29792 | 22330821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330821&pc=OZD | 06/11/2025 | 08/13/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic was unplugged and the battery life indicator immediately dropped from 100% to 50% and alarmed for low battery. the team quickly changed to backup aic. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. the impella device and data logs were returned for evaluation. the root cause of the "battery level low" alarm was most likely an erroneous trigger resulting from a single invalid data sample in the smbus communication between the power battery manager and the batteries. however, the specific component responsible was not determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29792. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8035 | 1220648-2025-29793 | 22330972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330972&pc=OZD | 06/11/2025 | 08/07/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026662802 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | Infusion or Flow Problem |  |  |  |  | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a run of ventricular tachycardia. echo was performed and the pump was slightly deep. the pump was pulled back 1 cm. after repositioning, purge pressure decreased with motor current and left ventricle (lv) waveform increased. purge flow decreased as well along with flow saturation. the team wanted to increase heparin due to possible thrombus ingestion. nothing was visualized on echo and support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of saturated flow, high purge pressure, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. saturated flow: data log review showed a motor current spike and immediate rise in purge pressure. the motor current also stepped up with the motor current spike. pump flows were saturated after the motor current spike. the root cause of the saturated flow was most likely biomaterial ingestion since there was a motor current spike followed by saturated flows. high purge pressure: data log review showing a motor current spike at the start of a 1-minute purge pressure rise from 400 to 1000 mmhg. no purge pressure high alarms triggered. the higher purge pressure resolved within an hour. heparin was then started in the purge fluid. no product was returned. the root cause of the high purge pressure was most likely biomaterial ingestion since there was a motor current spike prior to the 1-minute purge pressure rise. vf/vt/af/at: data logs are not applicable. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. positioning issues: data log review showed various positioning alarms throughout the case including an impella in aorta alarm prior to the purge pressure rise on (b)(6) 2025. the root cause of the positioning issues was not determined since insufficient clinical details were provided and product was not returned. thrombus: the failure mode thrombus was added since data log review indicated biomaterial ingestion related to high purge pressure and saturated flow. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
8036 | 1220648-2025-29794 | 22330995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22330995&pc=OZD | 06/06/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490836 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female | 100 KG |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient developed a hematoma at the impella access site. the patient was taken to the operating room and the hematoma was evacuated. two units of packed red blood cells were administered and heparin was held. additionally, the patient had mild discomfort at site but felt good otherwise. the patient was sitting in a chair. overnight, the patient had atrial fibrillation with rapid with rapid ventricular response which resolved with amiodarone drip. the bag and cassette were changed at one in the morning. increased in purge pressure was noted from 500/600's to 900/1000's. the purge system blocked alarm. alarm was disabled. a discussion was had and a decision was made to switch the purge solution to tissue plasminogen activator. no changes in flow or motor current were noted. echocardiogram was showing optimal impella position. another echocardiogram was done showing good position and no evidence of thrombus. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8037 | 1220648-2025-29795 | 22331057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22331057&pc=OZD | 03/03/2022 | 06/25/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  | 11/30/2023 |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169034 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Legal forwarded documents to the complaint team on that speaks to the impella cp support in which there is an allegation/litigation. the complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, after arriving in the icu unit, suction alarms were present, and bedside echo was ordered to verify positioning. impella was pulled back and proper positioning was verified. impella position continued to need adjustment during support. the patient¿s arterial blood pressure started trending down to maps of 50 and then 30. as expected, the patient did not have a pulse, however, the impella suggested decreased flow with low pressure readings. patient's norepinephrine was increased to 30 without effect. given that the patient had dropping maps to 30s, a code was called. epinephrine 1mg iv was administered, the impella pump was reduced to p2, and chest compressions were performed; patient had vt on the monitor and 1 shock was delivered following which patient had recovery of the maps to >70s. an epinephrine drip was started in addition to the norepinephrine gtt. patient was already on amiodarone gtt and a lidocaine gtt was also initiated. patient's impella was repositioned. a bedside echo was performed by the ccu fellow which showed poor ventricular contractility with an ef of 5-10%. patient continued to have drop in his maps (20s) while remaining pulseless and requiring another 1.5 epinephrine boluses as well as shock x 2. on exam, the patient had left mid-dilated pupil. the family decided for comfort care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8038 | 1220648-2025-29796 | 22331097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22331097&pc=OZD | 06/11/2025 | 12/04/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1723047 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) was undergoing preventive maintenance and found to have a red controller failure alarm. the issue was found outside of patient use. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: on 11th june 2025, though the epm hw revision was correct (sau_sens_eep_1 hw rev. 0 version c 0 epm_v6_0) as a known good epm hw revision, the console still received the controller failure (epm1 sw corruption) - alarm, event #1029. the console was turned off and rebooted four more times, and event #1029 was persistent. note: event #1029 could occur due to epm software corruption or the incompatible hw revision. device analysis: the console (b)(4) was returned to field service for further investigation. the reported failure mode was able to reproduced. upon booting, the console received the controller failure alarm. further, tried to access the firmware file through communicating with the micro, but was not able to even tried multiple times. the impellatronic pca was replaced with a known good one, after which the controller failure alarm not triggered. additionally, the console was run with the known good pump and purge cassette, and no issues were observed. root cause: the root cause for the controller failure (epm 1 sw corruption #1029) was due to the failure of the impellatronic pca.
8039 | 1220648-2025-29797 | 22331138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22331138&pc=OZD | 06/06/2025 | 07/16/2025 |  | 06/25/2025 |  |  | 06/06/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642007 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after an impella cp was implanted, the physician attempted to use perclose on the access site which resulted in perforation and dissection of the right iliac artery. the bleeding could not be managed with manual pressure so the impella cp was explanted and vascular surgery was performed to repair the artery. the physician elected not to exchange the peel-away sheath given the patient's vascular complications and repair. the patient expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Vascular damage peripheral vessel: the root cause of vascular damage peripheral vessel was most likely use issue since md attempted to perclose site x3 which resulted in r iliac artery perforation & dissection and bleed. access site bleeding: the root cause of access site bleeding was most likely use issue since md attempted to perclose site x3 which resulted in r iliac artery perforation & dissection and bleed.
8040 | 1220648-2025-29798 | 22331319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22331319&pc=OZD | 06/11/2025 | 06/25/2025 |  | 06/25/2025 |  |  | 06/11/2025 | 06/25/2025 |  |  |  |  | 06/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675573 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was a red alarm in ic "impella position in aorta" issue that resolved independently. additionally, the physician reported a suction alarm with a spike that was visible in the mc, and since then the mc had been flat, the impella was repositioned several times without success. heparin was paused for several hours due to the spike. it was noted of a clot being suspected. however, after several days of support, the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8041 | 1220648-2025-29717 | 22317654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22317654&pc=OZD | 06/16/2025 | 08/21/2025 |  | 06/24/2025 |  |  | 06/17/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605289 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death |  |  |  |  | NON-ABIOMED PCI GUIDEWIRE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that an impella cp was inserted after percutaneous coronary intervention (pci), but perforation by the 0.014 inch guidewire had occurred duringthe pci. the perforation was near the tip of left anterior descending artery. the treatment for the perforation was using a micro catheter and hemostatis performed with microthrombus and sponge. the guidewire was not damaged. the perforation was the main cause of deterioration in hemodynamics, leading to death due to multiple organ failure. | [5 empty] | Additional information was received stating that during the pci procedure, a perforation to left anterior descending artery (lad) was noted. this perforation was caused by a pci guidewire(non-abiomed), leading to death due to multiple organ failure. there was no allegation of a injury or malfunction attributed to the use of the impella. the death is unrelated. d10-concomitant device added. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | Additional information was received stating that during the pci procedure, a perforation to left anterior descending artery (lad) was noted. this perforation was caused by a pci guidewire (non-abiomed), leading to death due to multiple organ failure. there was no allegation of a injury or malfunction attributed to the use of the impella. the death is unrelated. there is currently no allegation of a malfunction or serious injury associated with the implanted pump. a regulatory report is no longer necessary. previously submitted manufacturer device report 1220648-2025-29717 is no longer applicable. h1 changed to na.
8042 | 1220648-2025-29718 | 22317717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22317717&pc=OZD | 06/19/2025 | 07/15/2025 |  | 06/24/2025 |  |  | 06/19/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1604236 |  | 0042-0040-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The patient expired due to preexisting cardiac condition and multiple organ failure. however, it is unclear if the automated impella controller (aic) issue contributed to the cause of death. | The user facility reported the automated impella controller (aic) had a "controller error alarm (yellow)" alarm sounded temporarily after using the impella and while returning to the intensive care unit. the position waveform display disappeared. even after taking measures such as re-inserting the connection cable, the alarm continued to appear frequently. the aic was replaced. the patient's vital signs remained stable during replacment. |  |  |  |  | Device returned on jul-14-2025 and a device evaluation is underway. should any new/relevant information become available, we will update the record b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29718. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29718. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29718. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29718. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-29718. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29718. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8043 | 1220648-2025-29719 | 22317765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22317765&pc=OZD | 05/27/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 05/27/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678020 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella cp was needing mechanical circulatory support. during impella cp support, it was noted that due to decreased distal flow to the right lower extremity and need for escalation of care, the patient was taken to the operating room for discontinuation of the impella cp and insertion of an impella 5.5. the patient was also peripherally cannulated for venoarterial extracorporeal membrane oxygenation. the procedure was successfully completed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8044 | 1220648-2025-29720 | 22317892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22317892&pc=OZD | 05/31/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 05/31/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640479 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male | 70 KG |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient had access site bleeding that was managed with tabotamp. however, the patient expired while on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8045 | 1220648-2025-29721 | 22317995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22317995&pc=OZD | 06/15/2025 | 09/03/2025 |  | 06/24/2025 |  |  | 06/18/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP- JP |  |  | ABIOMED, INC. |  |  | US |  | 2026616996 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after ecmo insertion, the impella cp was inserted for a cardiac case. the impella was inserted up to the pigtail, however the left ventricular cavity was narrow. although the pump was advanced along the 0.018 inch guidewire, it could not be advanced to the apex and got stuck halfway. while inserting and removing the pump, the guidewire got kinked. a new 0.018 inch guidewire was taken out and another attempt was made to insert it, however it did not work. the impella insertion was aborted. the patient was returned to the icu with ecmo and intra-aortic balloon pump (iabp). | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for delivery issues and product damage has been completed. the device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the impella was inserted up to the pigtail, but the left ventricular cavity was narrow and the impella didn't pass through. the root cause of the product damage issue was most likely patient condition related since the pump was advanced along the 0.018 inch guidewire, it could not be advanced to the apex and got stuck halfway. while inserting and removing the pump, the guidewire got kinked.
8046 | 1220648-2025-29722 | 22318034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318034&pc=OZD | 06/05/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/05/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656980 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a st-segment elevation myocardial infarction patient. the pump was placed to the left ventricle (lv) but the patient had ventricular tachycardia. the physician was concerned this could be from the impella placement. the impella cp was removed to rewire and implant for better position. an echocardiogram was performed and revealed a small lv cavity as well as a clot. cardiogenic shock (cgs) was believed to be coming from the right sided failure due to suspected pulmonary embolism. thrombectomy was performed and the physician elected to not replaced the impella cp as cgs was right ventricle not lv. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.".
8047 | 1220648-2025-29723 | 22318052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318052&pc=OZD | 06/03/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/03/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618558 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and was malpositioned by the mitral valve. the impella cp device was removed, the left ventricle was rewired, but the staff was unable to readvance the pump into position. the impella cp device was explanted; a plan to reevaluate the following day was noted. the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer as the site indicated the device was discarded. therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8048 | 1220648-2025-29724 | 22318126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318126&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/24/2025 |  |  | 06/05/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657887 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | [6 empty] | Fever | Unspecified Infection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella 5.5 support. during support the patient became febrile and the source of the infection remains unclear. at the same time, high purge pressure was experienced for which tissue plasminogen activator was used. later it was reported that care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis and high purge pressure has been completed. the impella device was not returned for evaluation. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. high purge pressure: data logs were reviewed and a purge pressure rise of approximately 7 hours was observed before "high purge pressure" alarms were triggered. no motor current spikes observed prior to purge pressure rise. purge pressure remained high for approximately 7 hours before slowly resolving. clinical details noted that purge pressure high alarms were triggered and lysis therapy was added to purge and purge flows improved. the root cause of the high purge pressure was most likely due to biomaterial build-up based on returned data log analysis and clinical details.
8049 | 1220648-2025-29725 | 22318154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318154&pc=OZD | 06/07/2025 | 07/16/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613601 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support due to microvascular disease and severe mitral regurgitation and was admitted for premature atrial contraction guided diuresis with low dose milrinone. during impella 5.5 support, the patient experienced a suction alarm followed by low pulsitility followed by an impella in aorta alarm. echocardiogram confirmed position of inlet in left ventricle. the console was reviewed which showed a motor current spike just prior to initial suction alarm. initially, the pump had near matching systolic and diastolic flows, and fluctuated between that or diastolic flows of 0.1 until the pump was explanted. upon explant, large biomaterial was noted in the outlet cage. the procedure was successful with this device. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of low pump flow and thrombus has been completed. low pump flow: the returned pump was severed along the catheter and only a partial pump was returned. the returned cannula was inspected and biomaterial was observed on and beneath the impeller. reproduction testing could not be performed as pump was cut. the data logs from the case show a motor current spike with speed devaluation and a shift in placement signal characteristic of biomaterial ingestion with suction alarms. the investigation aligns with the provided clinical information of biomaterial observed on explant and motor current spike observed before suction alarms. the root cause of the low pump flow was determined to be biomaterial as evidenced by returned product analysis revealing biomaterial on impeller and data log analysis showing ingestion pattern. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added
8050 | 1220648-2025-29726 | 22318250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318250&pc=OZD | 06/07/2025 | 08/05/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643508 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp was placed for support of a patient admitted in with known cardiac history and possible peripartum cardiomyopathy. the team placed an extracorporeal membrane oxygenation and then the impella cp was placed despite knowledge of possible apical thrombus. the pump was placed but there was pump saturation that prompted pump explant and replacement. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of saturated flow has been completed. the impella device was not returned for evaluation. thrombus was added based on review of the data logs. the cause of the issue was not determined since product was not returned. the clinical states that "echo showed left ventricle (lv) thrombus but medical care team made decision that risk versus benefit of impella placement would be best to go ahead and place impella due to concerns for lv distention. local team not called prior to placement. upon cp impella placement metrics showed pump saturation." the data logs show that there were spikes in the motor current immediately after starting the pump followed by saturated flow. based on the data log analysis and clinical information, the root cause of the saturated flow is most likely due to ingested biomaterial. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8051 | 1220648-2025-29727 | 22318305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318305&pc=OZD | 06/06/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/06/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681516 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Great Vessel Perforation | Cardiac Tamponade | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, the patient had a tamponade with a residual hematoma in the pericardium. heparin was used in the purge of the impella and may have contributed to the patient harm. the patient was taken for a washout and the impella cp was removed. after removal, the patient suffered severe mitral valve regurgitation. it was stated that the pump was pointed towards the mitral valve when placed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8052 | 1220648-2025-29728 | 22318376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318376&pc=OZD | 06/07/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580393 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Fever | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented to the cardiac catheterization laboratory to undergo a percutaneous coronary intervention and was subsequently admitted to the intensive care unit (icu). while in the icu, the patient began to require increasing amounts of vasopressor support and decision for impella 5.5 was made. it was noted that the impella was implanted without complication. during support, it was noted that the patient decompensated. ventricular tachycardia arrest x2, febrile, loss of pulsatility, lactate and creatinine trending up. decision was made to place protek duo for right ventricle support. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8053 | 1220648-2025-29729 | 22318385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318385&pc=OZD | 06/07/2025 | 08/05/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573447 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male | 137 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Defective Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, after peeling away the peel away sheath, a clear ring remained at the insertion site around the repositioning sheath. the decision was made by the doctor to leave the clear ring portion. the staff stated that they made attempts to remove the clear ring but were unsuccessful. the doctor stated that it could remain and the patient was transferred to intensive care unit with this in place. support was continued. additionally, the pump became malpositioned. the impella was retracted four centimeters with improvement of positioning alarms and the patient having ectopy. support was continued again. furthermore, the patient had access site bleeding. forward tension sutures were utilized and heparin was withheld. the procedural outcome was the patient survived the surgery. | [5 empty] | Additional information has been provided in h6 and h10. the investigation into the reported issues has been completed. no device was received for evaluation. visual and functional testing was not possible. the device passed all post sterile inspection checks. based on the clinical details, the root cause of the access site bleeding - major is most likely due to use as the forward tension sutures resolved the bleeding. the root cause of the ectopy was unable to be determined due to insufficient clinical details. due to lack of sufficient clinical information and no product returned, the root cause of the introducer damage - mechanical cannot be determined. due to lack of clincial details, it is unknown how the pump became in the incorrect position so the root cause of the positioning issues cannot fully be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8054 | 1220648-2025-29730 | 22318618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318618&pc=OZD | 06/07/2025 | 07/15/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665332 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Pulmonary Edema | Tachycardia | Ventricular Fibrillation | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented with chest pains and on electrocardiogram showed st-segment elevation myocardial infarction. the patient was taken to the cardiac catheterization laboratory for emergent percutaneous coronary intervention (pci). it was reported that shortly after arriving to catheterization laboratory, the patient went into ventricular fibrillation arrest and required cardiopulmonary resuscitation (cpr) and was defibrillated five times. the patient was intubated and pci was performed to left anterior descending artery. right heart catheterization performed and the impella cp inserted to right femoral artery. during support, pulses were confirmed and urology consult ordered for foley placement. heparin was held on arrival due to right groin hematoma. the pump was pulled back under echocardiogram guidance. it was noted that they were unable to get the foley catheter in place despite multiple attempts. the patient had visible blood at the urethral meatus. however, the next day, urology placed a foley with scope at bedside. also noted was that the critical care team repositioned the device with echocardiogram. ultimately, the patient decompensated severely into flash pulmonary edema upon induction for intubation, experience two rounds of ventricular tachycardia requiring defibrillation and cpr for nine minutes. return of spontaneous circulation was obtained. however, the patient was not responding to chemical or mechanical therapy. the decision was made to discontinue care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding: clinical states the hematoma to impella groin site on arrival to hillcrest medical center and patient was coughing a lot. the cause of the access site bleeding was most likely patient movement related as the hematoma occurred after patient transferred hospitals and was coughing a lot. positioning: clinical states the pump was pulled back under echo guidance to 3.9 cm on day 1 of support and again critical care repositioned the device bedside with echo 2 days later. the cause of the positioning issue was not determined due to lack of clinical details. vt: clinical states that the pt decompensated severely into flash pulmonary edema upon induction for intubation, experience two rounds of vt requiring defibrillation, cpr for nine minutes. in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
8055 | 1220648-2025-29731 | 22318809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22318809&pc=OZD | 06/07/2025 | 08/05/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665197 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 167 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella cp implanted and after the pump started, flows stopped at 1.4 liters per minute. the position was checked and volume was given, but suction alarms were noted at p-2/p-3 and had to go to p-1 to break suction. blood products were given to the patient. the cardiologist decided to remove impella cp and place intra-aortic balloon pump instead. no clot was noted in the impella when removed. there was no patient harm. | [5 empty] | Additional information provided in h3, h6 nd h10. the investigation into the reported issue has been resolved. impella cp was returned and biomaterial was found around the impeller inside the cannula. the pump was run in the lab with normal flow once the biomaterial was kicked off. data logs show suction throughout the case and low flows when the pump was started. there is a small motor current spike and spike in flows to 3l/min from when the pump is started indicating biomaterial may have been pulled in upon activation resulting in low flows. the root cause of the low pump flow was most likely biomaterial since biomaterial was found on the returned pump and data logs show an ingestion pattern. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8056 | 1220648-2025-29732 | 22319049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22319049&pc=OZD | 05/26/2025 | 07/21/2025 |  | 06/24/2025 |  |  | 05/26/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610813 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that thirty six days into impella 5.5 support, a patient's blood culture came back positive for an unspecified infection. the source of the infection was unknown. the patient was treated with intravenous antibiotics in response to the condition. roughly a month later, it was noted that ongoing placement signal unreliable alarms began to appear. both the waveforms and hemodynamics remained unchanged despite the noted signal issue. the patient remains on support. | [5 empty] | The investigation of infection/sepsis and optical signal issue has been completed. pump was not returned for investigation. the cause of the placement signal issue was sensor silicone wear beyond the intended design life of 60 days, based on data log review. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. d6b explant date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8057 | 1220648-2025-29733 | 22319309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22319309&pc=OZD | 05/31/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 05/31/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601542 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male | 128 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemolysis | Fungal Infection | Ischemia Stroke | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with the first impella 5.5 for mechanical circulatory support (mcs) and very soon thereafter, the patient had a pump stop and consequently, the impella 5.5 pump 1 was explant and replaced with another impella 5.5 pump 2 and cannulated for extracorporeal membrane oxygenation (ecmo). during this time, it was noted the patient experienced an ischemic stroke. clot ingestion was suspected. further information noted it is unknown exactly when the stroke occurred. sedation had not been weaned at this time, so neurological changes had not been noted. the physician noted likely clot was in left ventricle and undetected. following, a computed tomography scan revealed that the patient had experienced thromboembolic strokes leading to some hemorrhagic conversion. additionally, it was noted the patient developed hemolysis. the patient was taken off anticoagulants in response to the conditions. support was continued. now approximately 14 days later, it was noted the patient was running a fever for a few days. blood cultures were positive for fungus. all lines were changed with plan to re-send cultures. nephrology changed dialysis rate from 2.5 to 2. the patient continued on antifungal medication for the infection. latest update reflected the patient remained on impella mcs and noted to be in stable condition. | [5 empty] | This is one of two devices associated with the event. refer to manufacturing report number for the report on the impella 5.5 pump 1. the device remains implanted supporting the patient so therefore, evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8058 | 1220648-2025-29734 | 22319582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22319582&pc=OZD | 06/03/2025 | 07/18/2025 |  | 06/24/2025 |  |  | 06/03/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620869 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and when peeling away the 14 french sheath, vascular damage occurred. the impella cp was placed via the left femoral artery for left main percutaneous coronary intervention. the impella peel away sheath was going to be removed. however, it was peeled away in the artery. the device was removed for bleeding. the hematoma was stopped with a 16 french gore sheath. significant right arm hematoma and bleeding at the a line site as well was noted. the patient required and was administered four units of blood. it was noted that the patient responded well and was reported to be in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The root cause of the vascular damage issue was most likely use related, the peel away technique used by the md caused vascular damage and bleed/hematoma.
8059 | 1220648-2025-29709 | 22319965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22319965&pc=OZD | 06/09/2025 | 09/26/2025 |  | 06/24/2025 |  |  | 06/09/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1798913 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that shortly following impella cp startup, the placement and lv signal were lost. a yellow controller error subsequently appeared and the decision was made to swap out the automated impella controller (aic). both the motor current and flows were not affected by the noted failure. upon successful transfer to a new aic the signals appeared and no further issues were encountered. | [5 empty] | A5, race, is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of optical signal issue and automated impella controller (aic) - controller error (yellow alarm) been completed. this console was tested upon arrival and the root cause of the placement signal issue and the controller error was due to an aic software issue. this was entered into jama with defect/ bug id gid-dft-2927090. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
8060 | 1220648-2025-29735 | 22320161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320161&pc=OZD | 06/03/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/03/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642089 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and during support, the device malpositioned requiring explant and resulted in access site bleeding complication. after the impella cp was implanted via the left femoral artery, the patient was prepped and transferred via medevac to an outside facility for higher level of care. the patient arrived at the receiving hospital awake and following commands. the plan was to start epinephrine to wean down levophed, complete a point of care bedside ultrasound, send off for labs and evaluate with cardiothoracic surgery for an impella 5.5 upgrade. upon a bedside echocardiogram, the impella cp was found to be a little deep with the pigtail touching the wall. the resident verbalized discomfort in repositioning the impella cp and waited for the attending physician (who is at the time was involved in other emergencies) to assist. additionally, noted was the patient's left leg was cooler and dusky. the decision was made later that morning to escalate to the impella 5.5. the impella cp was removed from the access artery and perclose was attempted and held for approximately 10 minutes. however, the groin started bleeding. the decision was then made to do an open surgical repair for removal which was completed. the patient was stabilized and returned back to the intensive care unit. the status of the patient following the event was noted as unknown. however, plan of care updates indicated the patient remained on impella mcs for at least seven days. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8061 | 1220648-2025-29736 | 22320320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320320&pc=OZD | 06/07/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657792 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support labs were hemolyzed and the team were discussing options. the next day the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8062 | 1220648-2025-29737 | 22320354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320354&pc=OZD | 06/07/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666816 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed for the patient admitted in with acute myocardial infarction and cardiogenic shock. the pump was placed but the patient was exhibiting signs of hemolysis within first hours of support. the team called for air transport to send patient to tertiary center but, the patient expired while coding prior to air transport. the family stopped the code and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8063 | 1220648-2025-29738 | 22320442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320442&pc=OZD | 06/07/2025 | 07/16/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had oozing from the access site. a stitch was placed and support continued. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of access site bleeding was most likely use issue since hemostasis was achieved by adding additional stitch in the icu.
8064 | 1220648-2025-29739 | 22320500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320500&pc=OZD | 06/05/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/05/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572039 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female | 73 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced bleeding at the access site, atrial fibrillation and subsequently expired. the patient with left ventricular ejection fraction of approximately 25% and aortic valve area of 0.4 centimeters squared was admitted to the hospital and experienced cardiogenic shock alongside hypotension and respiratory failure. two days later, the patient underwent aortic valvuloplasty and impella cp insertion. the following day, bleeding at the impella insertion site was noted. the patient received one unit of red blood cells due to a drop in hemoglobin. the dressing was noted to have been changed several times and hemostatic dressing was applied. the next day, it was noted that a decision was made by the patient's son to not escalate care and withdraw care if no improvement were made. the pump was repositioned with hopes of improvement. support was continued and the patient remained critically ill. the patient's blood pressure was extremely labile, remained on high doses of pressors and performance level p-9. the plan was to stay course until the son made a final decision. the patient was taken off levophed over the weekend with improvement but went into atrial fibrillation with rapid ventricular response the following morning and needed levophed support. the physician noted if support was continued the patient would need cardioverted and be placed on continuous renal replacement therapy. however, this day the patient would expire. care was withdrawn. | [5 empty] | Although the patient's family member declined to continue with escalation in care and subsequently withdrew care, there is not enough evidence to exclude the impella device used as an associated factor to the overall outcome (given the impella related complications). the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8065 | 1220648-2025-29740 | 22320514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320514&pc=PYX | 06/07/2025 | 08/05/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026671812 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was taken to cardiac catheterization laboratory for placement of impella rp flex due to worsening right ventricle failure status post extracorporeal membrane oxygenation (ecmo) cannulation via tandem heart. decision was made to place the impella rp flex via right ij due to current occupied cannulation in the groin via ecmo. however it was noted that the patient had hemolysis and the decision was made to explant the impella rp flex to avoid further complications per team. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. the pump was not returned for investigation. the data log shows a motor current spike followed by a drop in pump flow and hike in the placement signal. central venous pressures and optical parameters were not affected after the motor current spike incident. pump flow and placement signal remained the same after the motor current spike until the end of the case run. the cause of the hemolysis issue was most likely biomaterial ingestion, based on data log review. the root cause of the thrombus was unable to be determined due to insufficient clinical details.. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8066 | 1220648-2025-29741 | 22320714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320714&pc=OZD | 06/07/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622847 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with shortness of breath and was diagnosed with atrial fibrillation with rapid ventricular response, cardioverted and admitted to intensive care unit. the patient began decompensating and a pa catheter was placed immediately. decision for impella 5.5 was made for biopsy to recovery or advanced heart failure therapies. the impella 5.5 was placed without complication. during support it was noted that echocardiogram showed inlet slightly flipped towards mitral valve. also, low ci, attempted to increase to p-9 however had suction alarm. the doctor attempted to reposition 5.5 and the patient went into ventricular tachycardia requiring cardioversion x 3. a second attempt to reposition and required cardioversion x2. the patient was critically ill while supported on ventilator and extracorporeal membrane oxygenation. the patient remains on support awaiting liver work up. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8067 | 1220648-2025-29742 | 22320727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320727&pc=OZD | 06/01/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/01/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655387 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Atrial Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient experienced atrial fibrillation during support. no intervention was noted in response to the condition. the following day, the pump stopped unexpectedly. the device was restarted, but stopped a second time. upon restart again, the motor current was no longer pulsatile and persistent suction / low pump flow occurred. the pump was explanted in response to the failure and biomaterial was visualized within the pump outlet. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8068 | 1220648-2025-29743 | 22320784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320784&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/24/2025 |  |  | 06/08/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578353 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. implant was attempted, however unable to be delivered due to the patient¿s anatomy. the case was aborted, and pump was thrown away post case. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to resistance in femoral artery. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8069 | 1220648-2025-29744 | 22320945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320945&pc=OZD | 06/08/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/08/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666192 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was explanted after 44 hours due to limb ischemia and was escalated to an impella 5.5 pump. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8070 | 1220648-2025-29745 | 22320958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22320958&pc=OZD | 06/01/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/01/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026688988 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. at the end of support, the patient experienced limb ischemia. the pump was removed due to the ischemia. the patients outcome was noted as stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8071 | 1220648-2025-29746 | 22321067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22321067&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/24/2025 |  |  | 06/08/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662059 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was oozing at the access site due to patient movement. pressure was held and the sheath was pushed in further to slow the bleed. the patient moved again, and the site continued to bleed. two units of packed red blood cells were given to the patient. a right femoral artery angiogram was performed and found a clot around the sheath access site, but patient found to have great distal flows down their leg past repo sheath site. the access site was redressed. heparin was withheld. support continued. additionally, the pump was noted to have suction. the p level was reduced which resolved the suction. however, two units of blood products were given shortly after, and it is unknown if this was related to the suction. support continued. | [5 empty] | The investigation of thrombus, access site bleeding, and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: access site bleeding complication was achieved through surgery. the cause of the access site bleeding was most likely patient movement due to the bleeding occurring during the patients inability to remain still despite being instructed to do so and subsequent sedation efforts. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. low pump flow: data log review showed suction alarms occurring but stopped after p-level was reduced down from p-7. an motor current (mc) spike was noted later on in the case with no noticeable impact to flows or purge. the cause of the low pump flow was most likely patient condition related as the suction alarms resolved by giving volume and reducing the p-level. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8072 | 1220648-2025-29747 | 22321174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22321174&pc=OZD | 06/08/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/08/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678531 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, bilateral cold and mottled feet were observed. no flow visualized on ultrasound. contrast through reaccess port shows restricted flow. the decision was made to place fem-fem bypass, utilizing a braided antegrade sheath. the patient¿s feet are now warm and normal color. good distal pulses and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8073 | 1220648-2025-29748 | 22321262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22321262&pc=OZD | 06/08/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/08/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675708 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had ventricular fibrillation, and coded. chest compressions were performed and resuscitation was achieved and support continued. the patient eventually expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8074 | 1220648-2025-29749 | 22321317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22321317&pc=OZD | 06/08/2025 | 08/06/2025 |  | 06/24/2025 |  |  | 06/08/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639615 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp and rp flex devices for mechanical circulatory support. while on support, the team escalated to a va ecmo from the rp. the cp supported despite bleeding from all access sites and concerns of limb ischemia for 12 hours before the family made the choice to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. access site bleeding : bleeding from all access sites. the root cause of the access site bleeding issue was most likely patient condition related since there was bleeding from all access sites. limb ischemia : the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. h6 health effect - impact code 1802 and component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29749.
8075 | 1220648-2025-29750 | 22321877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22321877&pc=PYX | 06/07/2025 | 10/31/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1726714 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 94 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA RP FLEX | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella rp had placement signal not reliable alarm after patient movement. the patient was monitored closely until the pump was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: the logs align with the reported issue. the case originated on (b)(6) 2025, and shortly after the pump was started, multiple suction alarms were triggered. the pump was then stopped and restarted several times. during the most recent pump start, alarms for placement signal not reliable (dp signal out of range #1051) and placement signal not reliable (epm sensor failure #1052) were observed. despite disconnecting and reconnecting the pump multiple times, these alarms persisted. furthermore, the lt_logs and rt_logs consistently showed repeated instances of the same alarms: placement signal not reliable (dp signal out of range #1051) and placement signal not reliable (epm sensor failure #1052). additionally, the second aic used, displayed the same alarms as well. device analysis: the console (b)(6) was returned to field service for further investigation. a visual inspection of the internal circuitry and cable connections of the console did not reveal any issues. the console was tested with a known good pump and purge cassette, and no anomalies were observed during testing. additionally, the analog output from the ccd array of the optical bench (fringe pattern) was seen without any distortion, and the "5s" parameter using 2 different optical cavity found that the values were within the specified values as per the pm procedure (0042-7309). root cause: the root cause of the "placement signal not reliable (rp dual sensor, signal out of range for 2 minutes # 1051)" and "placement signal not reliable (rp dual sensor, sensor failure or signal out of range for 2 minutes # 1052)" alarms was not related to any issue with the aic. no issues were identified with the aic. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h6 updated. d10 concomitant medical products and therapy dates updated.
8076 | 1220648-2025-29751 | 22321897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22321897&pc=OZD | 06/07/2025 | 09/03/2025 |  | 06/24/2025 |  |  | 06/07/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484677 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male | 86 KG |  |  | Injury | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that a patient was placed on an impella 5.5 as venting and long term support solution due to dilatative cardiomyopathy. during 5.5 support, it was noted that when the patient was on dialysis therapy, a suction alarm on the automated impella controller occurred. however, this was good to handle for the team. prolonged partial thromboplastin time was carefully increased into theraupic range due to bleeding issues (not impella related). however, the patient was systemically heparinized. it was also reported that the impella had slipped into the aorta and stated the pump dislodged due to trying to reposition the pump without pressing the yellow repositioning button. this damaged the catheter and it was no longer possible to pull it forward or backward. the device was removed. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues, vascular damage - peripheral vessel, and product damage (catheter kink) has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issues was most likely use-related since clinical details state that pump was attempted to be repositioned without pressing the repositioning button. product damage (catheter kink): the cause of the product damage (catheter kink) was most likely use-related since clinical details state that pump was attempted to be repositioned without pressing the repositioning button, which damaged the catheter and prevented it from being moved forward/backward. vascular damage: the cause of the vascular damage was not determined since insufficient clinical details were provided and product was not returned. h6 medical device problem code 2889 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29751.
8077 | 1220648-2025-29753 | 22322127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322127&pc=OZD | 06/10/2025 | 08/07/2025 |  | 06/24/2025 |  |  | 06/10/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655562 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, there were suction alarms. albumin was administered. additionally, an infection was noted at the access site. the doctor intentionally left the vessel loops at the access site and this caused an infection. antibiotics were administered. impella support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of low pump flow and infection/sepsis has been completed. the impella device was not returned for evaluation. low pump flow: data log reviewed, no motor current spike outside of p-level changes observed. no positioning issue observed. many suctions occurred on the reported issue date. less suctions observed a few days later until it solved at p4. placement signal was trending low at the time of the suction event on 6/10. purge spikes occurred. the cause of low pump flow issue most likely was patient condition related since suctions decreased after blood administrated and lowered p-level. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. d6b explant date added.
8078 | 1220648-2025-29752 | 22322128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322128&pc=OZD | 06/01/2025 | 08/15/2025 |  | 06/24/2025 |  |  | 06/01/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661768 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 63 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a right groin bleed. heparin had been administered leading up to the event. the patient was administered one unit of blood in response to the bleed. the patient continues on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
8079 | 1220648-2025-29754 | 22322263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322263&pc=OZD | 06/01/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/01/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653684 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female | 108 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient experienced frequent runs of ventricular tachycardia. the pump was repositioned and the p-level was dropped in response to the noted condition. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8080 | 1220648-2025-29755 | 22322399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322399&pc=OZD | 06/09/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/09/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613433 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 82 KG | Non Hispanic | White | Death | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Product damage was reported by a customer implanting an impella cp. the sterile sleeve tore off the distal end of the device and the pump became unsterile. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8081 | 1220648-2025-29756 | 22322443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322443&pc=OZD | 06/01/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/01/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641099 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. it was reported that after the sheath was exchanged, the patient developed a large hematoma under the skin. the impella cp was removed due to the hematoma. an intra-aortic balloon pump was placed on the left groin. femstop was placed on right groin after perclose was deployed. the patients outcome was noted to be stable. | [5 empty] | The impella device was discarded, and the investigation has been completed. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8082 | 1220648-2025-29757 | 22322476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322476&pc=OZD | 06/09/2025 | 08/07/2025 |  | 06/24/2025 |  |  | 06/09/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661732 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 81 KG | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue and the patient experienced vascular damage while attempting to implant an impella 5.5. the impella 5.5 was unable to be placed due to axillary vessel size. multiple wires and guides were used but were unable to be placed. the decision was made to implant an impella cp instead. after implantation of the impella cp it was discovered on echo that the subclavian artery had been dissected to the aortic root. the surgeon believed that the dissection occurred when having difficulty placing the impella 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: & impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. the device was discarded. the root cause of delivery issue was determined to be patient condition because the pump was unable to pass through left ventricle due to small vessel size. the root cause of the vascular damage was determined to be patient condition because the patient experienced vascular damage while attempting to insert the pump through the small vessels. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ b.1 revised as product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29757. b.7 added medical history as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29757. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-29757. h.6 code 4114 was reported incorrectly on the initial manufacturer device report number 1220648-2025-29757. code 4115 has been added as the device was discarded. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the initial manufacturer device report number 1220648-2025-29757 stated that the device was not returned but the device was discarded.
8083 | 1220648-2025-29758 | 22322590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322590&pc=OZD | 06/01/2025 | 07/21/2025 |  | 06/24/2025 |  |  | 06/01/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 84 KG | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during implantation, the pump was noted to be shallow and the inlet was against the inferior wall. the doctor attempted to reposition and accidentally pulled the pump back across the valve. the doctor was unable to advance the pump back across the valve. the pump was explanted and a new axillary insertion kit was used. the doctor attempted to reaccess the left ventricle for over 1.5 hours without success. the patient remained stable and the decision was made to abort the case at this time. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was pulled back to aorta accidentally while repositioning of the pump.
8084 | 1220648-2025-29759 | 22322619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322619&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/24/2025 |  |  | 06/09/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female | 81 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella cp encountered a thrombus, low pump flow and ventricular fibrillation (vf). after 10 minutes of support, the impella began to suction with no flow. the patient became hypotensive and arrested. a thrombus was found occluding the aortic root, for which chest compression was performed. the thrombus migrated and pump flows returned. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of thrombus, low pump flow, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the data log shows several suction alarms and low pump flow during the case start, with an observed trendy pump flow and placement signal trend. pump flow was recovered and remained within specifications for the rest of the case run. there was no indication of biomaterial interaction with the impeller during the case. according to the clinical notes, flow returned to normal after a large thrombus migrated from the aortic root. the cause of the low pump flow issue was not determined without returned product investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29759.
8085 | 1220648-2025-29760 | 22322806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22322806&pc=OZD | 06/06/2025 | 06/24/2025 |  | 06/24/2025 |  |  | 06/06/2025 | 06/24/2025 |  |  |  |  | 06/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665241 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp pump migrated across the aortic valve during patient support. after multiple failed attempts to recross the valve, the decision was made to leave the pump in the descending aorta for explant the following morning. patient outcome is unknown. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8086 | 1220648-2025-29698 | 22306499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22306499&pc=OZD | 06/06/2025 | 07/16/2025 |  | 06/23/2025 |  |  | 06/05/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653705 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to vent the extracorporeal membrane oxygenation for the patient admitted in with cardiogenic shock and known valve disease. it was reported that after placing the 10mm graft on the right axillary artery, the surgeon was unable to advance the 5.5. after multiple dilations, he was still unable to advance with the 5.5. patient¿s anatomy would not support the 5.5. decision to use a impella cp which the cp advanced with no issues. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29698. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8087 | 1220648-2025-29699 | 22306755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22306755&pc=OZD | 06/03/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/02/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601143 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support (mcs) experienced ventricular fibrillation and then subsequently limb ischemia resulting in explant of the device. the patient was admitted with st-segment elevation myocardial infarction and two days later was taken for to the catheterization lab for percutaneous coronary intervention to the ostial left anterior descending artery (lad) and left main artery with two drug-eluting stents. the patient was transferred to the unit and developed "severe" chest pain followed by cardiac arrest treated with cardiopulmonary resuscitation and epinephrine with return of spontaneous circulation. the patient was awake and alert post arrest and electively intubated and emergently taken back to catheterization lab where a coronary angiography revealed acute thrombosis of proximal lad. the decision was made at this time to establish mcs with an impella cp prior to any intervention. the impella was successfully placed via the right femoral artery, and the thrombosis was treated with mechanical thrombectomy and new stenting. during the procedure with impella mcs, the patient experienced three episodes of ventricular fibrillation requiring defibrillation. the patient was returned to the unit remaining with impella mcs. however, the following day, approximately 11 hours later, no doppler pulses was noted on the right foot. the physician made the decision to explant the impella cp device at bedside, which was successfully completed with a survived outcome. the patient remained on two pressors. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
8088 | 1220648-2025-29700 | 22306841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22306841&pc=OZD | 06/04/2025 | 08/26/2025 |  | 06/23/2025 |  |  | 06/05/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273761 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, looking at the automated impella controller (aic) data, it was noticed that the aic date was not updated. the date was updated. the aic was turned off and then reactivated, but the date did not update. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data analysis: evaluation of the logs confirms the customer complaint of updating the date and time but the changes not persisting after a reboot. device analysis: the console was returned. the coin cell battery that provides power to the rtc while the console is unpowered was tested with a multimeter and measured 0v. the failure of the aic to hold the correct date and time settings was due to a depleted rtc battery. d9 device returned to manufacturer date added g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29700. g1 manufacturing site name, address, country, and fax number was erroneously entered on the initial manufacturer device report number 1220648-2025-29700. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29700.
8089 | 1220648-2025-29701 | 22307959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22307959&pc=OZD | 06/03/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/03/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump support ongoing for 7 days for the patient awaiting heart transplant when there was purge priming issues with the aic (console). the aic was swapped out and replaced. there was no harm or injury noted from the interruption in support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8090 | 1220648-2025-29634 | 22308151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22308151&pc=OZD | 04/02/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/05/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625955 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the packaging of an introducer required excessive force to open. it was noted that the use of two clamps and significant force was needed to finally open the packaging. there was no patient harm. upon return of the noted introducer, an evaluation revealed evidence of improper material separation. some residual fragments of molded plastic remained adhered to one side of the introducer hub, while the opposite side showed a corresponding cavity or defect indicating incomplete parting. there was no patient harm. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. all introducer from batch #s9241349 passed all inspection checks. the 14fr x 13cm introducer was returned for investigation. as per the clinical, peel-away sheath could not be opened manually. significant force and the use of two clamps were required. the returned introducer was already peeled away. further investigation of the returned product reveals evidence of improper material separation. a residual fragment of molded plastic remains adhered to one side of the introducer hub. the opposite side shows a corresponding cavity or defect indicating incomplete parting. preliminary investigation attributes the excessive peel-away force to a molding defect on the returned product. the root cause of the introducer damage issue was difficulty peeling away the sheath due to molding defect.
8091 | 1220648-2025-29702 | 22308523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22308523&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | ABIOMED, INC. |  |  | ABIOMED, INC. |  |  | US |  | 2025574040 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, access site bleed prompted intervention. the team achieved hemostasis by redressing and re-suturing the site and the pump remained on for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related since re sutured, and angle matched, re-dressed, oozing was resolved.
8092 | 1220648-2025-29703 | 22308696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22308696&pc=OZD | 06/06/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612723 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had ventricular tachycardia (vt)/ventricular fibrillation (vf). the patient had six (6) episodes, all requiring shocks and cardiopulmonary resuscitation. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8093 | 1220648-2025-29704 | 22308861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22308861&pc=OZD | 06/06/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026667311 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient coded for about one minute. cpr started and one of epi given before return of spontaneous circulation (rosc) was achieved. left main received 1 drug-eluting stent (des) and cx 1 des. the patient's urine was red and all labs hemolyzing. the pump was repositioned, and p level was reduced. additionally, the patient went back into afib with rapid ventricular response (rvr), requiring increased impella support. the patient required increased pressor, having oxygenation issues. the care was ultimately withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8094 | 1220648-2025-29705 | 22308919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22308919&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted in a patient an impella cp generated suction and high purge pressure alarms that were resolved infusing fluids and lasix. the position of the impella was confirmed to be correct. no patient harm was reported. | [5 empty] | The investigation of high purge pressure and product damage (catheter kink) has been completed. the impella device was not returned for evaluation. product damage (catheter kink): clinical stated brief high purge pressure (hpp) alarm due to a kink, resolved quickly. unknown location and how the kink occurred. the cause of the catheter kink was not determined due to insufficient clinical details. high purge pressure: complaint device not returned for analyzing. data log reviewed, several brief hpp events suddenly occur and suddenly resolve which are indicative of a purge line kink corresponding with the clinical details that the cause of the hpp was a kink in the pump. two motor current spike events were observed but did not correlate with any trend in pp or hpp event. the cause of the hpp was a kinked purge tubing as reported in the clinical and evidenced in data log review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8095 | 1220648-2025-29706 | 22309089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309089&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632492 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male | 75 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced a gastrointestinal bleed. bivalirudin was withheld. impella support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
8096 | 1220648-2025-29707 | 22309116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309116&pc=OZD | 06/06/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668061 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient attempted to stand up and suffered sustained ventricular tachycardia (vt) which required a total of 3 cardioversions. initial echo showed inlet 7cm from av and pump was pulled back. vt continued at p-levels greater than 2. a decision was made to leave at p2 to reduce irritability for several hours as the patient was hemodynamically stable. then the p-level was raised to p-5 and ultimately to p-8 without return of ventricular arrythmia. echo showed that the pump appeared in appropriate position. the patient then assisted to the chair without ectopy | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8097 | 1220648-2025-29708 | 22309198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309198&pc=OZD | 06/07/2025 | 07/21/2025 |  | 06/23/2025 |  |  | 06/07/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675663 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the patient went into sustained ventricular tachycardia/atrial fibrillation. the patient was given amio bolus. | It was reported that a patient implanted with an impella 5.5 experienced oozing with a hematoma at the access site. one unit of blood products was transfused to the patient. additionally, there was high purge pressure noted. tissue plasma activator was added to the purge. impella support continued. |  |  |  |  | Access site bleed: clinical details report that the patient was noted to have oozing at the rax access site on (b)(6) 2025. the patient reported that they sleep on their right side with their arm elevated to cradle their head. additionally, they were noted to be young, restless and mobile. the oozing was noted to have subsided the following day. product was not returned for investigation. based on the clinical detail provided that the patient was restless, mobile, and would raise their access site arm over their head to sleep, the root cause of the access site bleed was determined to most likely be patient movement. high purge pressure: clinical details reported that the high purge pressure alarm triggered on (b)(6) 2025. tpa was added to the purge, and by the second bag, the alarm had resolved and purge pressure values were decreasing steadily. there were no further reports of purge issues. data logs were reviewed the high purge pressure event was confirmed. the evening of (b)(6) 2025, purge pressure began trending up gradually over 7 hours leading up to the alarm. there were no mc spikes or long bag changes prior to the rise. data logs showed sodium bicarbonate being taken out of the purge during the rise in purge pressure, likely indicating this was when tpa was implemented. 2 hours after the alarm triggered, purge pressure started to decrease gradually, and within 30 hours, the pressure values stabilized between 700-800 mmhg. product was not returned for investigation. based on the gradual rise in purge pressure over a number of hours without a mc spike in data logs that resolved gradually after tpa was added to the purge, the root cause of the high purge pressure was most likely biomaterial buildup. vt/vf: clinical details reported that the patient experienced svt-afib on (b)(6) /2025. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29708. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29708. h6 health effect - clinical codes 1729, 2095, and health effect - impact code 4644 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29708. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8098 | 1220648-2025-29710 | 22309508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309508&pc=OZD | 06/09/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/09/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675476 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient encountered limb ischemia while implanted with an impella cp. the impella cp was explanted with plans to implant an impella 5.5 instead. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8099 | 1220648-2025-29711 | 22309633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309633&pc=OZD | 06/04/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/04/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653446 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 support experienced hemolysis. the patient started having progressive dark red urine and labs were hemolyzed. the 5.5 was explanted. the patient continued on ecmo support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8100 | 1220648-2025-29712 | 22309688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309688&pc=OZD | 06/06/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026671038 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was successfully weaned, explanted, and perclose deployed. bleeding was still noted after perclose deployment, so 8fr angioseal used and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8101 | 1220648-2025-29713 | 22309851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22309851&pc=OZD | 06/06/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/06/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026678206 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.5 device for mechanical circulatory support. during the implant, the valve would not open. the implant was aborted, and the patient was placed on extracorporeal membrane oxygenation (ecmo). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8102 | 1220648-2025-29714 | 22310117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22310117&pc=PYX | 06/04/2025 | 08/01/2025 |  | 06/23/2025 |  |  | 06/04/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026647412 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella rp flex support and during support, there were elevated pa placement signal and decreased impella flows. it was unknown if there was possible sensor damage or a clot ingestion. the clot was never visualized or confirmed. blood volume was given to the patient. support continued. the pump was explanted due to elevated ldh and decreased platelet count. patient survived the procedure. | [5 empty] | The investigation of low pump flow, thrombocytopenia, hemolysis, and optical signal issue has been completed. the impella device and data logs were received for investigation. low pump flow: product was returned and evaluated. no cannula kinks were found. there was biomaterial found wrapped around the base of the impeller and some biomaterial on the blades. no other abnormalities were found. the pump was run in a bucket and pump flow was slightly low. the pump was soaked to remove biomaterial and was then hemolysis tested. pump flow was low in hemolysis testing and also the subsequent flow loop testing. after the hemolysis and flow loop test, borescope images showed a clean impeller and cannula. the root cause of the low pump flow was not determined since the low flow reproduced after clearing the impeller of biomaterial, there were no other abnormalities found, and data logs were insufficient. placement signal issue: the failure mode was updated from optical signal issue to placement signal issue since data log review showed no abnormalities with the optical parameters but rather the dp sensor. the pump was returned and evaluated. the sensor passed electrical testing. there was biomaterial found wrapped around the base of the impeller and some biomaterial on the blades. the pump was run in a bucket and pump flow was slightly low. there were no clear abnormalities with the sensor face. the pump was then run in the hemolysis and flow loop and it did not show signs of sensor drift. the root cause of the placement signal issue was not determined since the issue did not reproduce in the lab and data logs were insufficient. thrombocytopenia: clinical details state that the pump was removed due to concerns for elevated ldh and decreased platelet count. data logs are not applicable and product was not returned. the cause of the thrombocytopenia was not determined due to insufficient clinical details. hemolysis: data log review showed suction shortly after pump startup. on the night of 04-june-2025, there was a slight step down in mc and speed deviations. there were no clear motor current spikes throughout the case; however, rt logs were unavailable in the beginning of the case. the root cause of the hemolysis was most likely biomaterial as data log review showed suction and motor speed deviations the day before hemolysis was noted, the returned product had biomaterial wrapped around the base of the impeller, and the pump passed hemolysis testing. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
8103 | 1220648-2025-29715 | 22310275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22310275&pc=OZD | 06/04/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/04/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026676842 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed ischemia in their right foot. it was noted that the distal flow beyond the sheath was present, however pedal pulses were not dopplerable. the patient had an unknown underlying peripheral arterial disease (pad) and their right foot became mottled and blistered. vascular surgery was on standby to monitor the situation. additionally, the patient coded during support, around 9:30 pm. currently in junctional rhythm in the 50s. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8104 | 1220648-2025-29716 | 22310474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22310474&pc=OZD | 06/09/2025 | 06/23/2025 |  | 06/23/2025 |  |  | 06/09/2025 | 06/23/2025 |  |  |  |  | 06/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638694 |  | 1000080 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 107 KG | Non Hispanic | White | Injury | Yes | No | Malposition of Device | [6 empty] | Arrhythmia | Hemolysis | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with an implanted impella cp. the patient developed anterior st changes and an angiogram revealed severe left main disease and total occlusion of the right coronary artery. the patient experienced an episode of ventricular tachycardia for which defibrillation was administered. pulses in the distal leg were lost and echo showed the pump was found to be positioned deeper than expected. the patient's urine was red in color. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist system lists ventricular arrythmia as an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8105 | 1220648-2025-29674 | 22304400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304400&pc=OZD | 06/03/2025 | 07/16/2025 |  | 06/22/2025 |  |  | 06/03/2025 | 06/22/2025 |  |  | 06/30/2025 |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430740 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during and post pci operation, the flows of an implanted impella cp pump would not increase beyond 1.8 lpm. around that time, purge flows decreased and a high purge pressure alarm appeared. the pump was subsequently replaced to resolve the noted issues. | [5 empty] | The investigation of low pump flow and high purge pressure has been completed. the pump was returned for investigation. no physical damage was noted on the returned product, and no biomaterial was observed. the pump was then tested in a bucket of water and was able to run according to specifications. low pump flow and high purge pressure did not reproduce on the returned product. total case run time was 1.5 hours. the data log shows several motor current spikes during the start of the case while the pump was set to steady p-level p9, followed by low pump flow and a gradual rise in purge pressure. purge pressure was seen to resolve later during the case run. the cause of the low pump flow issue was most likely biomaterial, based on returned product investigation and data log review. the cause of the high purge pressure issue was most likely biomaterial, based on returned product investigation and data log review. thrombus failure mode was added as an investigated failure mode based on the data log investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
8106 | 1220648-2025-29675 | 22304401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304401&pc=PYX | 06/04/2025 | 07/09/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026622720 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on bipella support developed hemolysis. the rp was exchanged for protek duo. the next day the labs were no longer hemolyzing. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of hemolysis has been completed. the returned product was inspected and there was no biomaterial observed. there were no pump abnormalities and pump passed hemolysis test. clinical mentioned hemolysis occurred leading to pump explant on (b)(6). the data logs show several minor motor current spikes with shifts in placement signal and effects on pump flow in the morning of (b)(6), a pattern likely occurring as a result of biomaterial interacting with the impeller. the root cause of the hemolysis was most likely biomaterial as evidenced by the data log pattern observed and pump passing hemolysis testing.
8107 | 1220648-2025-29676 | 22304402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304402&pc=OZD | 06/04/2025 | 07/02/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610541 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, there was a new onset of atrial tachycardia, patient given antiarrhythmics. additionally, there was a hematoma at the access site. patient was given red blood cells and site monitored for any changes. anticoagulation was on hold. support continued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D6: explant date not available. access site bleeding: the root cause of access site bleeding was most likely use issue since non-abiomed product (vascutek gelweave 10 mm) was used which is not recommended per ifu. vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
8108 | 1220648-2025-29677 | 22304405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304405&pc=OZD | 06/04/2025 | 08/06/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 -EU |  |  | ABIOMED, INC. |  |  | US |  | 2026628540 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Failure to Pump | [5 empty] | Anemia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Wound Infection | Thrombocytopenia | [7 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported an impella 5.5 pump support ongoing for 1 day for a high-risk surgical patient. the pump position was poor throughout the support and was ingesting biomaterial as it layed at the ventricle wall. the pump stopped and was explanted. patient survived coming off support. 2 weeks later again the impact study updated that there was also thrombocytopenia attributed to the pump and patient was administered "substitute" for the platelet drop. also, there was a stroke that lead to limb paralysis on the right side attributed to pump use. patient outcome is unknown. | Since the original report was filed, there were additional adverse events reported by the study. the patient required long term ventilation and developed a wound infection at the impella site that prompted antibiotics and wound vacuum pump, additional developed elevation in bilirubin count. |  |  |  |  | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | B5 updated with additional information received from the clinical site h6 updated with additional failure modes received from the clinical site the investigation of positioning issues, pump stop, stroke/cva, thrombocytopenia, and infection/sepsis has been completed. the impella device was returned for evaluation. positioning issues: repositioning done multiple times and it always adhered to the heart wall. return product has no major physical abnormalities apart from biomaterial. the root cause of the positioning issue was not determined due to insufficient clinical details. pump stop: in the morning, during ongoing support, emergency procedure to remove impella because myocardium was in the inlet cage. further confirmed impella was running on p3-p4 and had a lot of suction issues. biomaterial was observed at outflow during explant. return product had biomaterial wrapped around impeller at outflow and biomaterial ingested all over cannula/inflow cage. no open electrical lines or other physical abnormalities were found. the pump was run in bucket of water and low flows reproduced with flows fluctuating between 0.1 to 3 l/min after initial high motor current (mc) pump stop. the root cause of the pump stop issue was most likely biomaterial based on clinical and return product analysis. stroke/cva: as per clinical, there was a stroke that lead to limb paralysis on the right side attributed to pump use. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. thrombocytopenia: the root cause of the thrombocytopenia issue was not determined due to insufficient clinical details. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. thrombus: failure mode added based on product evaluation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b2 disability was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29677. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29677. h6 health effect - impact code 4641 and 4618 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29677. h10 narrative erroneously stated that the device was not received on the initial manufacturer device report number 1220648-2025-29677.
8109 | 1220648-2025-29678 | 22304409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304409&pc=OZD | 06/04/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637138 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that prior to repositioning sheath placement, a hematoma began to develop around the access site. the perclose were tightened to manage the condition and the peel-away sheath was swapped with the repositioning sheath. despite best efforts, some oozing was still noted at the site. that same day, the patient endured significant arrhythmias and was in and out of junctional rhythm. support continued with the implanted pump throughout the event. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8110 | 1220648-2025-29679 | 22304411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304411&pc=OZD | 06/04/2025 | 07/18/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629876 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on the list for transplant and developed a fever and had hypotension was listed as status 7. physician decided to place intra-aortic balloon pump to help with worsening cardiogenic shock. however, due to the patient's continual decompensation decision was made for impella 5.5 placement. it was reported that the impella 5.5 was then primed and initial access was gained into left ventricle (lv), however some difficulty noted crossing with the 0.035 j wire. contrast shot and showed some stenosis distal on the axillary artery. they elected to try insertion with 0.018 steel core wire and impella 5.5. upon insertion the pump crossed anastomosis however was unable to pass stenosis on the distal part of the artery. after multiple attempts the 5.5 was removed and place an axillary impella cp instead. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following : section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature despite multiple attempts with many trouble shooting techniques.
8111 | 1220648-2025-29680 | 22304428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304428&pc=OZD | 06/04/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636373 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient had a run of ventricular tachycardia during impella cp support. an amio drip was initiated in response to the issue. support continued uninterrupted. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8112 | 1220648-2025-29681 | 22304440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304440&pc=OZD | 06/04/2025 | 07/31/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during the implant, the implant was on the patient's right side with short sheath, but the impella kinked as they were trying to insert it, so they took impella out and reinserted with the long sheath. even with the long sheath, it was hard to pass with standard j and pig into the left ventricle so doctor exchanged for a stiffer wire and was able to pass with ease. when the companion sheath was inserted, it helped straighten out the sheath and the rest of the pci was completed. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ | The investigation of product damage (cannula kink) and delivery issues has been completed. the impella device was not returned for evaluation. product damage (cannula kink): product was not returned for investigation. based on the clinical details provided that the patient had tortuous iliac arteries, the root cause of the product damage (cannula kink) was determined to most likely be patient condition. delivery issue: product was not returned for investigation. based on the clinical details provided that the patient had tortuous iliac arteries and all best practices were followed, the root cause of the delivery issue was most likely due to patient condition.
8113 | 1220648-2025-29682 | 22304447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304447&pc=OZD | 06/04/2025 | 07/31/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp became malposition during weaning. when flow decreased from p-7 to p-2 the impella migrated into the aorta which was confirmed with echo. patient hemodynamically stable. the patient was taken for successful removal. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the pump was mispositioned during the weaning process. the cause of the positioning issue was most likely use related. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8114 | 1220648-2025-29683 | 22304455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304455&pc=PYX | 06/04/2025 | 07/17/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026655491 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella rp flex pump was unable to advance into the right ventricle during attempted delivery. the implant was subsequently aborted as a result of the complication. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of delivery issue was most likely patient condition since impella cannula kept getting caught up in the rv.
8115 | 1220648-2025-29684 | 22304462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304462&pc=PYX | 06/04/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/04/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479846 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella rp flex support and during support, the patient had multiple episodes of ventricular tachycardia (vt). no further information was provided. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8116 | 1220648-2025-29685 | 22304479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304479&pc=OZD | 06/06/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664161 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced access site bleeding and hematoma resulting in explant of the device. the patient transported via emergency medical services for chest pain and electrocardiogram was completed and indicated an anterior myocardial infarction. the patient went into ventricular fibrillation on arrival to the emergency department and was shocked three times. the patient was taken to the catheterization lab left heart catheterization which revealed 100$ occlusion of the proximal left anterior descending artery and impella implant which was placed via the right femoral artery. percutaneous coronary intervention was completed with one stent and post dilation noted with "nice" result. the impella peel away was removed from body and peeled away. repositioning sheath was advanced into place and there was oozing around site. preclose sutures were pulled to tighten down around sheath. however, oozing continued. angle was adjusted steeper with slowing of the oozing. the catheter was secured and appropriately dressed. a growing hematoma was then noted, and the site was still oozing despite best practices. the patient was hemodynamically stable; therefore, the physician decided to explant the impella cp device. the device was weaned and explanted without issue. perclose sutures were tightened down and pressure was held for additional 20 minutes. the patient was then transported to the unit noted to be in stable condition. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
8117 | 1220648-2025-29686 | 22304483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304483&pc=OZD | 06/05/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/05/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support (mcs) experienced ventricular tachycardia and would subsequently expire. the patient underwent a left and right heart catheterization, and the decision was made to consult surgery. as the patient waited for surgery consult, the patient coded. cardiopulmonary resuscitation was initiated, and one shock was delivered. the patient was transported to the catheterization lab for the impella cp device implant via the left femoral artery. performance level was ramped up to p-5. the patient had two episodes of ventricular tachycardia necessitating shocks. a pulmonary catheter was place and the patient was return to the intensive care unit in stable condition. however, some time thereafter the patient expired. further details were unnoted. total impella mcs time was 18 hours. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8118 | 1220648-2025-29687 | 22304497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304497&pc=OZD | 06/06/2025 | 08/05/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658648 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was unable to be placed. the aorta only measured 4.5 cm from root to innominate. left axillary has a sharp ¿s¿ curve and ¿u¿ bend anatomy, therefore, the 5.5 placement was aborted and an impella cp was implanted via femoral artery. | [5 empty] | Additional information provided in d9, h3, h6 and h10. the investigation into the reported issue has been completed. the pump was returned and no abnormalities were found. the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to aorta only measuring 4.5 cm from root to innominate along with left axillary having sharp s curve and u bend anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8119 | 1220648-2025-29688 | 22304511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304511&pc=OZD | 06/06/2025 | 09/02/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656236 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella had less flow that decreased over time. the position was verified. the intensivist decided to stop the impella and replace it with an intra-aortic balloon pump (iabp) for unloading. upon explant, biomaterial was observed. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. low pump flow: complaint device not returned for investigation. data log reviewed: many suction alarms occurred during support. suction alarm followed by drop in flows and motor current (mc)spike and flows continued in dropping on nigh of 06-jun-2025. no placement signal issue seen. purge spike and impella position unknown occurred. clinical details stated, after moving the patient the impella had less flow and flows decreased more over time. repositioning and left ventricle (lv) filling didn't solve the issue, lv thrombus and clot inside the canula observed. the cause of low pump flow issue most likely was biomaterial ingestion as biomaterial and lv thrombus observed and log review confirmed with occurrence of mc spike. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8120 | 1220648-2025-29689 | 22304520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304520&pc=OZD | 06/06/2025 | 07/16/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664105 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had hemolyzed labs and red-tinged urine despite repositioning the pump. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis and thrombus has been completed. the impella device was not returned for evaluation. hemolysis: no product was returned for analysis. clinical mention hemolyzed labs which continued despite pump repositioning. no patient abnormalities were noted. the data logs show a motor current spike on 6/6 a little after midnight before hemolyzed labs were reported later in the morning. there was a motor current spike with speed deviation characteristic of an ingestion event. the root cause of the hemolysis was most likely biomaterial as evidenced by the ingestion pattern observed. the failure mode thrombus was added as a finding during the investigation. thrombus: the root cause of thrombus was unable to be determined due to insufficient clinical details.
8121 | 1220648-2025-29690 | 22304536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304536&pc=OZD | 06/06/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674550 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited low flow and pain in the right leg. an external fem-fem bypass was performed to resolve the limb ischemia. later, the patient complained of numbness, pain, and loss of sensation in the limb. the impella was explanted to resolve the issue. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8122 | 1220648-2025-29691 | 22304551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304551&pc=OZD | 06/06/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492233 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female | 92 KG | Non Hispanic | Black Or African American | Death | Yes | No | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that delivery issues were encountered when attempting to implant an impella 5.5. resistance was met at anastamosis so the pump was removed, flushed, and implantation was successful. the patient also exhibited bleeding at the surgical site. the site was surgically repaired. the patient's family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
8123 | 1220648-2025-29692 | 22304561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304561&pc=OZD | 06/06/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670925 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male | 68 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported after explantation of an impella cp the patient lost pedal pulses. the common femoral artery was found to be occluded. it was ballooned and stented to resolve the limb ischemia. the patient survived. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8124 | 1220648-2025-29693 | 22304569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304569&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662055 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female | 57 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp support experienced bleeding from the radiofrequency ablation site, and the patient's hemoglobin dropped from 12 to 5. additionally, the pump repositioning sheath had backed out despite being sutured in place. four units of red blood cells were transfused and the sheath was repositioned to resolve the issues. the next day the patient experienced a retroperitoneal bleed and additional blood products were transfused. care was withdrawn and the patient expired. | [5 empty] | The device was not returned for evaluation. should the device or any additional information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding is determined to be use issue because bleeding was resolved by advancing the repo sheath back into place and angle matching. vascular damage - peripheral vessel: the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-29693. h6 medical device problem code 2616 was erroneously entered on the initial manufacturer device report number 1220648-2025-29693.
8125 | 1220648-2025-29694 | 22304571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304571&pc=OZD | 06/06/2025 | 08/01/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the implanted impella cp pump was pulled back into the aorta after repositioning the patient. the pump was explanted and replaced. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was most likely patient movement based on the clinical narrative indicating that the pump became dislodged after rolling the patient. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8126 | 1220648-2025-29695 | 22304575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304575&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673819 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 84 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced access site bleeding. the bleeding would not stop so the pump was explanted and a new impella cp was implanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was determined to be most likely use issue because the pump was placed without placing the perclose.
8127 | 1220648-2025-29696 | 22304577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304577&pc=OZD | 06/06/2025 | 08/04/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659289 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male | 159 KG | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced bleeding as the impella peel away sheath was accidently pulled back. a large hematoma developed as a result. the patient's family withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation positioning issue: the cause of the positioning issue most likely use related as the peel-away was accidentally pulled back during cardiopulmonary resuscitation (cpr). access site bleeding: the cause of the access site bleeding issue most likely was use related as the peel-away was accidentally pulled back during cpr leading to hematoma. b2 intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29696. h6 health effect - impact code 4641 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29696.
8128 | 1220648-2025-29697 | 22304588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22304588&pc=OZD | 06/06/2025 | 06/22/2025 |  | 06/22/2025 |  |  | 06/06/2025 | 06/22/2025 |  |  |  |  | 06/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669993 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male | 102 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp developed a hematoma above the access site. a fem stop was placed and heparin was withheld to address the hematoma. the patient then experienced ventricular tachycardia requiring defibrillation and administration of lidocaine and amiodarone. after the patient was transferred the patient's urine became red. the pump's position was verified and the groin was redressed. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8129 | 1220648-2025-29637 | 22303541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303541&pc=OZD | 06/04/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671774 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a high purge pressure/purge system blocked alarm occurred on an implanted impella 5.5. the high purge pressure spontaneously resolved without intervention. impella support was continued. additional information was received noting unrelated to the impella. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8130 | 1220648-2025-29638 | 22303543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303543&pc=OZD | 06/01/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/01/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658218 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. three days after start of the support a patient update noted the patient's status as critical. due to suction the performance level is at p-6, and the patient was septic (no further information regarding this complication was noted). partial thromboplastin time: 47 seconds (high dosage heparin need was noted) and access side was dry. escalation or transfer was not possible according to the nurse. the next day, the impella cp was successfully weaned and explanted with a survived outcome. further status is unknown. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8131 | 1220648-2025-29639 | 22303546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303546&pc=OZD | 06/04/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640496 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 83 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and exhibited a mottled and cool lower extremity. the limb lost doppler pulse and showed weak non-pulsatile flow. the pump was weaned to p4 to restore doppler pulses, color, and temperature. due to the patient's poor prognosis, the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8132 | 1220648-2025-29640 | 22303548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303548&pc=OZD | 06/01/2025 | 08/28/2025 |  | 06/21/2025 |  |  | 06/01/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600691 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and at start, the impella cp device showed zero flow. re were suction alarms. the impella cp pump was repositioned, and the performance level was decreased to no avail. echocardiogram was completed in addition to an x-ray to verify the pump position. on echocardiogram, a something was seen near the mitral valve and was thought to be thrombus. as a result, the impella cp device was explanted. the pump was placed in water, the pump worked well, and no thrombus was noted under optical observation. the patient was reported to be in stable condition following the event. it is unknown if the patient was placed back on mcs as the site was uncertain if another impella cp device was in stock. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. impella cp pump was returned. pump did not have any cannula kinks found during visual inspection and no damaged blades. borescope view of cannula obtained and biomaterial wrapped around impeller was found. no data logs were returned for investigation. the clinical reported pump was experiencing low flows after impella pump start. position was confirmed to be good and no cannula kinks observed. the root cause of the low pump flow issue was most likely biomaterial based on return product investigation. the root cause of the thrombus could not be determined. | The impella device was received from the customer for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8133 | 1220648-2025-29641 | 22303552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303552&pc=OZD | 06/01/2025 | 08/28/2025 |  | 06/21/2025 |  |  | 06/01/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026639124 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and approximately 34 days after start of the support, there was a drop in purge flow and increase of purge pressure. catheter shaft and puncture area were checked, and no abnormalities identified. anticoagulation was in upper therapeutic range; there was no issue with the purge line, and there were no spikes or rise in the motor current. the plan was to explant the impella 5.5 for bridge to left ventricular assist device the following day. later this same day, it was noted the purge pressure issue was unchanged. purge lyse was discussed but was not an option given the plan. the following day the impella 5.5 device was successfully weaned and explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure (hpp) has been completed. the impella device was not returned for evaluation. clinical details report that a high purge pressure alarm suddenly triggered on 1/jun/2025. there were no changes in motor current (mc) noted, no kinks in the purge line/catheter, and anticoagulation was in therapeutic range. tissue plasminogen activator (tpa) could not be administered due to a scheduled left ventricular assist device (lvad) placement. no further details were provided. data logs were reviewed and the hpp alarm was confirmed on 1/jun/2025. leading up to that point, purge values were stable and there were no long bag changes. at the time of a p-level drop, purge pressure started spike up quickly and triggered the alarm 2 minutes later. when the p-level dropped, mc remained constant, as if there had been no change in motor speed. finally, it was noted that the purge flow tick stalled out when purge pressure spiked. it appeared that it was unable/struggled to advance forward. purge pressure remained elevated for 22 hours. during this time, there was a purge cassette change, but it didn't resolve the issue. while the purge pressure was elevated, there was a period that the nature of the values were jagged: there were small spikes and drops in the pressure. on 2/jun/2025, purge pressure suddenly trended back down to normal values in a 10 minute period. there was some variability in mc and ms that appeared abnormal during the resolution, but the relationship between the trends could not be confirmed. the patterns seen in data logs could be indicative of biomaterial and/or a kinked purge line causing the complication. details from the clinical narrative to not support one cause over another. product was not returned for evaluation. since product wasn't returned for evaluation, the signatures seen in data logs could be indicative of multiple different causes, and the clinical narrative doesn't support either, the root cause of the high purge pressure could not be determined.
8134 | 1220648-2025-29642 | 22303556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303556&pc=OZD | 06/02/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/02/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650230 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed a gastrointestinal bleed. anticoagulation was withheld and support continued. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8135 | 1220648-2025-29643 | 22303557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303557&pc=OZD | 06/04/2025 | 07/18/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661920 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 127 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed a hematomaat the insertion site after the repositioning sheath was advanced. the peel away sheath was removed from the body completely prior to cracking and peeling away. the patient was not on any systemic anticoagulation. the last known activated clotting time was 260. the patient received integrillin bolus after stent placement. forward tension sutures were placed and angle matching dressing was applied. manual pressure was held for 20 minutes. five percent dextrose in water was in the purge solution. right lower extremity was mottled and pulses were absent. hematoma was extending per the nurse. the impella cp was removed and the artery was repaired and the hematoma evacuated. the patient was to be escalated to an impella 5.5. four units of packed red blood cells were also transfused. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel, limb ischemia, and access site bleeding has been completed. the impella device was not returned for evaluation. vascular damage peripheral : the cp was removed and vascular repair the artery. the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical details provided. ¿limb ischemia: as all the necessary information was not provided, the root cause of the limb ischemia was not determined." access site bleeding major: hematoma was noted at impella site. the root cause of the access site bleeding issue was most likely use related, forward tension sutures. and angle matching dressing applied to stop the bleeding. instructions for use impella cp with smart assist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ h6 health effect - clinical code 1888 and health effect - impact code 4624 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29643. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-29643.
8136 | 1220648-2025-29644 | 22303574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303574&pc=OZD | 06/02/2025 | 07/21/2025 |  | 06/21/2025 |  |  | 06/02/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644902 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. during the implant, after multiple attempts to pass the impella cp through the descending aorta, the pump was removed. there was kinking of the pump below inlet of device. a new impella was placed successfully. | [5 empty] | The investigation of delivery issues and product damage (cannula kink) has been completed. reported unable to pass to pass descending aorta and the kinking below the inlet of the device. patient has vessel calcification and tortuosity. impella cp was returned and a cannula kink was found below the inlet cage. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient had vessel calcification and tortuosity. cannula kink - the root cause of the product damage (cannula kink) was patient condition related since the patient had vessel calcification and tortuosity. | The impella device was received from the customer for evaluation. upon closure of the investigation, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
8137 | 1220648-2025-29645 | 22303598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303598&pc=OZD | 06/03/2025 | 10/30/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665339 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female | 67 KG |  |  | Injury | Yes | Yes | Increase in Pressure | Deformation Due to Compressive Stress | [5 empty] | Stroke/CVA | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an impella 5.5 could not be repositioned because of a cable problem inside the repositioning unit outside the body. there was a poor processing in the cable so that it did not fit into the touhy burst. the pump was secured but there was no chance to reposition the pump. the patient was still on support. the pump's position was verified and was ok. no patient harm was reported. | High purge pressure and cerebral hemorrhage was reported under manufacturer device report 1220648-2025-30206. 25-45798 is a duplicate of this complaint. investigating results will be submitted under this manufacturer device report 1220648-2025-29645. the complainant reported that nineteen days into impella 5.5 support, the purge flow dropped while the purge pressure increased to 1000 mmhg. lysis therapy was discussed but ultimately declined due to the patient's existing cerebral hemorrhage. information pertaining to the bleed and its origin was not provided. the staff was advised to replace the pump and decisions regarding further treatment were pending. |  |  |  |  | The pump and purge cassette were returned for investigation. no data logs were returned for analysis. product damage (catheter kink): clinical details report that the catheter was kinked on (b)(6) 2025. the team was unable to move the catheter through the touhy-borst valve due to the kink. for this reason, they could not reposition the pump. that being said, position was checked via imaging and confirmed to be good. no excessive force was reportedly used. no further details were provided. an image was returned in addition to the product. in the image it appears there is a kink in the catheter just outside the touhy-borst valve within the sterile sleeve, which was confirmed during product evaluation. data logs are not relevant to the reported issue. the root cause of the catheter kink was not determined due to lack of sufficient clinical details. high purge pressure: the root cause of the high purge pressure was most likely biomaterial in the purge gap based on the product evaluation. since no data logs were returned for analysis, it could not be confirmed whether it was biomaterial buildup or ingestion. the pump was further investigated for thrombus in case of ingestion. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (catheter kink), high purge pressure, and stroke has been completed. the impella device was not returned for evaluation. product damage (catheter kink): product was not returned for investigation, however, an image was. in the image it appears there is a kink in the cathether just outside the touhy-borst valve within the sterile sleeve. however, without returned product to confirm the cath-anchor is functioning properly, the root cause of the catheter kink could not be determined definitively. data logs are not relevant to the reported issue. since product was not returned for investigation, the root cause of the catheter kink could not be determined. high purge pressure: data logs were not available for investigation and product wasn't returned for investigation. since neither data logs nor product were available for review, the root cause of the high purge pressure could not be determined. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ it was discovered that high purge pressure and stroke were reported in a duplicate manufacturer device report number 1220648-2025-30206. the following corrections were made to manufacturer device report number 1220648-2025-29645 to include high purge pressure and stroke and the investigation results. b1 adverse event was inadvertently omitted b5 updated with corrective statement and patient narrative. h1 type of report was erroneously selected. h6 health effect - clinical code 4582 and health effect - impact code 2199 were erroneously entered. h6 component code 925 was inadvertently omitted. h10 instructions for use were inadvertently omitted.
8138 | 1220648-2025-29646 | 22303600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303600&pc=OZD | 06/13/2025 | 09/03/2025 |  | 06/21/2025 |  |  | 06/13/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP -JP |  |  | ABIOMED, INC. |  |  | US |  | 2025595604 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella cp placement was aborted due to the patient¿s vessel condition which were calcified. no patient harm reported. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation has been completed. no product was returned. the root cause of delivery issue was most likely patient condition since there was a severe calcified lesion in the ao bifurcation.
8139 | 1220648-2025-29647 | 22303614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303614&pc=OZD | 06/02/2025 | 07/02/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643340 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs). later that evening, the patient experienced bleeding and hematoma at the access site. consequently, the impella site was redressed, and a stitch was placed. the patient continued on mcs that remained for approximately two days. it was noted the patient tolerated the impella mcs support well, inotropes were decreased, and the device was successfully weaned and explanted with a survived outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since adding an additional stitch resolved the issue. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8140 | 1220648-2025-29648 | 22303619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303619&pc=OZD | 06/16/2025 | 09/05/2025 |  | 06/21/2025 |  |  | 06/16/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638628 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Decrease in Pressure | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complainant reported the patient was on impella cp support and during support, there was a decrease in purge flow. a blood clot was suspected and the pump was explanted. a fibrin-like material was found to have adhered to the pump. there were no patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of high purge pressure and thrombus has been completed. the impella device was not returned for evaluation. high purge pressure: data logs were reviewed and purge pressure/flow were stable through (b)(6) 2025 around 12:00. at this time, there was a sudden step down in the motor current (mc) without any p-level change. this correlated with the first drop in purge flow and a minor increase in purge pressure from 450 to 550 mmhg. then, roughly 2 hours later, there was another significant drop in purge flow and step up in purge pressure from 550 to 800 mmhg that correlated with a spike in the placement signal (ps) to 150 mmhg. there were no motor current spikes, but mc min values did drop just before this event without a change in p-level. then, 1 hour later, just before the pump was explanted, there was a clear motor current spike to 800 ma with a speed deviation. this was followed by an 8 minute rise in purge pressure above 1000 mmhg. the pump was explanted before purge pressure rose high enough to trigger an alarm. the initial two spikes/steps up in purge pressure did not correlate with mc spikes but still happened quickly. it is unclear if a kinked purge line could have caused this or biomaterial of some sort (ingestion or buildup). there was an ingestion signature just before pump explant that caused the purge pressure to continue to trend higher, however, it cannot be determined if this biomaterial was responsible for the initial rise as well. product was not returned for investigation. since it is unclear based on data logs what caused the initial onset of increased purge pressure and product wasn't returned for investigation, the root cause of the high purge pressure could not be determined. thrombus: it was reported that thrombus was visualized on the pump outflow cage on explant. the root cause of the thrombus was unable to be determined due to insufficient clinical details. h6 medical device problem code 1490 was erroneously entered on the initial manufacturer device report number 1220648-2025-29648.
8141 | 1220648-2025-29649 | 22303624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303624&pc=OZD | 06/03/2025 | 07/18/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604698 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful. the impella 5.5 was unable to pass through the anatomy. the physician attempted many troubleshooting techniques including rotoglide, repositioning of the arm, removing neck roll, not tunneling, and attempted a buddy wire technique all with no success. the decision was made to implant an impella cp instead which supported the patient for approximately seven days before being weaned and explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature despite multiple attempts with many trouble shooting techniques.
8142 | 1220648-2025-29650 | 22303636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303636&pc=OZD | 05/24/2025 | 07/21/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635440 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was also reported that during the impella 5.5 implantation the impella was advanced over the .018-inch guidewire however there was a kink and curl noticed in the aorta. the wire was removed and placed with a v18 wire. impella 5.5 was placed over v18 with no difficulty. | The user facility reported that the patient on impella 5.5 support developed hemolysis. plasma free hemoglobin was 140. blood products were administered, and hemolysis was resolved. patient remained on support. |  |  |  |  | The investigation of hemolysis and product damage (guidewire kink) has been completed. the pump was not returned for investigation. the data log did not record any abnormalities related to positioning issues, motor current spikes, or the patient condition at the time of the reported hemolysis event. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided. d6b explant date added b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29650. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-29650. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29650. h6 medical device problem code 2981 and component code 463 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29650. h6 medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2025-29650. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8143 | 1220648-2025-29651 | 22303639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303639&pc=OZD | 06/03/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620428 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of an impella 5.5 to a patient with cardiomyopathy was unsuccessful. it was not possible to advance the pump over the guidewire. the patient was reported to in stable condition following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8144 | 1220648-2025-29652 | 22303646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303646&pc=OZD | 05/29/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 05/29/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605321 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed an expanding right groin hematoma following impella cp implant. the patient was taken to the operating room for urgent escalation to impella 5.5 for mixed cardiogenic and hemorrhagic shock. the bleed was originating from around the repositioning sheath. the condition was controlled via vascular surgery and the hematoma was evacuated. the patient continued with impella 5.5 support following the intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8145 | 1220648-2025-29653 | 22303654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303654&pc=PYX | 06/03/2025 | 07/29/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529184 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella rp flex implant and there was kinking and buckling at the inlet and motor housing. not able to advance catheter any further without this issue. decision made to start over with a new catheter with fear of damage from the kinking and buckling. further details state that once the bi-pella was placed, the patient expired on the table. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage(cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be use issue because the new rp flex was inserted successfully. product damage(cannula kink): the root cause of product damage, cannula kink is determined to be use issue because the new rp flex was implanted successfully.
8146 | 1220648-2025-29654 | 22303657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303657&pc=OZD | 05/31/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 05/31/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666726 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs), and during support, the patient experienced an intracranial hemorrhage. the patient arrived at the emergency department and was diagnosed with urinary tract infection. the patient was sent home, returned and was admitted for shortness of breath. twelve hours later, the patient was intubated and septic and multi-organ failure. the decision was made to implant an impella cp and cannulate for venovenous extracorporeal membrane oxygenation (vv ecmo). the patient was tolerating the impella mcs well and on the fourth day of support, a plan to go to the operating room for ecmo decannulation and determine whether the patient needed to upgrade to an impella 5.5 or not as noted. the physician explained to the family the outcome was grim but will use transesophageal echocardiogram to determine flow and function. head scan later this day showed intracranial hemorrhage. the patient was unarousable with sedation off for approximately 24 hours. heparin was discontinued due to the bleed. due to the intracranial hemorrhage, the family was deciding on comfort care. the following day, it was noted the impella cp was explanted. it is unclear if the patient was placed into comfort care. based on the information at hand the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; it was reported after the device was explanted, it was subsequently thrown away. therefore, an evaluation/analysis of the device was not possible. should any new information be received, a supplemental manufactuer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8147 | 1220648-2025-29655 | 22303673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303673&pc=OZD | 06/03/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263045 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during prepping the impella aic would not detect the purge disc. the aic was exchanged. | [5 empty] | The investigation for the purge disc/flags issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) multiple times. the console was returned for evaluation. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
8148 | 1220648-2025-29656 | 22303677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303677&pc=OZD | 06/02/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/02/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026675613 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon removal of the 14 french dilator, bleed back occurred through the 14 french x 13 centimeter sheath. the dilator was placed back through the sheath to stop the bleed and then the entire sheath was replaced with a new introducer. impella cp implant was successful following the swap. upon removal, a split was noted at the top of the sheath. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8149 | 1220648-2025-29657 | 22303680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303680&pc=PYX | 06/03/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661849 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella rp flex support and during support, labs were showing signs of hemolysis. the pump appeared deep and was repositioned and pulled back 2 cm. support continued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8150 | 1220648-2025-29658 | 22303690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303690&pc=OZD | 06/03/2025 | 07/18/2025 |  | 06/21/2025 |  |  | 06/03/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672056 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump was unable to advance through the aortic valve during attempted delivery. upon reaching the level of the aortic valve, a kink in the catheter occurred due to heavy calcification of the subclavian artery. the implant was aborted and an impella cp pump was placed for planned support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to heavy calcification of the subclavian artery. product damage(cannula kink): the root cause of product damage is determined to be catheter kink because the catheter got kinked during the insertion while attempting to pass impella through heavy calcified subclavian artery.
8151 | 1220648-2025-29659 | 22303694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303694&pc=OZD | 11/24/2024 | 08/29/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Literature | Health Professional | User facility | Company representation | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Difficult to Insert | Difficult to Remove | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a literature publication from krakow found entitled: "how to unfold a folded large-bore sheath after impella-supported percutaneous coronary intervention". the publication discussed a patient who had impella support and during the implant, the impella sheath could not deliver, the 14f isleeve sheath by boston scientific was implanted. support continued. during the removal of the pump, the pump was stuck within the isleeve sheath and could not be retrieved. staff delivered a long shaft peripheral 4.0 × 40 balloon, which was inflated at the site of resistance. then, after partial removal of the sheath, the impella was successfully removed. furthermore, angiography showed a minor dissection of the right external iliac artery. | [5 empty] | The investigation has been completed. no product was returned. the cause of the introducer insertion issues was most likely patient condition related due to the noted tortuosity and calcification of the iliac vessels. the cause of the pump removal issues was the use of a sheath that was not indicated for use with impella. the cause of the issue was not determined since product was not returned and insufficient clinical details. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d4: model number, serial number, catalog number and lot number are unknown. d6a & d6b implant and explant dates are unknown. h4: device manufacture date unknown.
8152 | 1220648-2025-29660 | 22303708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303708&pc=OZD | 06/04/2025 | 07/02/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658499 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma with oozing at the access site. one unit of blood products was given to the patient and pressure was applied. the following day, the patient had nausea and vomiting over night that resulted in bleeding from impella site, resolved with pressure and dressing change. | [5 empty] | The root cause of access site bleeding was most likely patient movement since the patient developed a large hematoma after being log rolled having nausea and vomiting over night that resulted in bleeding from impella site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8153 | 1220648-2025-29661 | 22303718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303718&pc=OZD | 06/04/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  | 11/30/2025 |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430058 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 82 KG |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella position in aorta alarms on and off and was repositioned at bedside and alarms corrected. lactic was trending up. oliguric with rose colored urine noted. dialysis catheter was placed. the heparin was restarted and was off related to right brachial hematoma after a-line attempt. based on increasing lactic, vasoactive support, declining kidney function, possible hemolysis. in addition, it was noted that impella positioning remained difficult to maintain and hemolysis was present. hemodynamics and ef were improving and impella was being weaned down in effort to decrease hemolysis. low h/h with unknown etiology and 2 units of blood were given. the patient also went into vt for 1 minute, however, did not require cpr/defib. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high-risk pci. instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position. ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8154 | 1220648-2025-29662 | 22303730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303730&pc=OZD | 06/04/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508813 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, neuro status changes prompted imaging. limb movement, and pupil changes were observed. the pump remains on despite the infarct while awaiting high risk percutaneous cardiac insertion (hrpci). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8155 | 1220648-2025-29663 | 22303762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303762&pc=OZD | 06/04/2025 | 08/01/2025 |  | 06/21/2025 |  |  | 06/04/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026672061 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male | 64 KG | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, echo showed impella migration to the sept wall. the physician does not want to reposition with normal hemolytic indices and absence of suction alarms. two days later, the patient had runs of ventricular tachycardia (vtach). the patient was taken to or for placement evaluation due to vtach. the patient developed bleeding from the access site. heparin was paused. no further instances since. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: impella access site bleeding reported around the sheath after insertion, bleeding continued to jun 8th. heparin paused. unknow if bleeding stopped and how bleeding started. the cause of bleeding issue cannot be determined due to insufficient clinical details. positioning issues: post implant echo showed impella migration to septal wall. physician does not want to reposition with normal hemolytic indices and absence of suction alarms. patient had severe heart failure and underwent heart transplant surgery. the cause of positioning issue most likely was due to patient condition related. vf/vt/af/at: patient had v tach event. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details.
8156 | 1220648-2025-29664 | 22303788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303788&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026648713 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the heart team decided to perform high risk percutaneous cardiac insertion (hrpci). once pci was completed with hemodynamic stability, the decision was made to leave the impella in overnight for myocardial rest. it was noted that right groin hematoma developed shortly after arriving at the intensive care unit. pressure was held to the site and integrillin was stopped. additionally, they noted a spontaneous leaking sidearm at the luer lock. the purge cassette was changed, and a hair line crack was noted on the bottom of check valve luer during reservoir prime. dermabond was applied to the site of the leak and purge pressure was reestablished. the impella was successfully weaned, removed and pressure held to achieve hemostasis. groin dressing applied post with the site intact without hematoma. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and purge leak ¿ pump has been completed. access site bleeding: the cause of the access site bleeding was most likely high patient activated clotting time (act) values at 268. purge leak - pump: data log review showed purge pressure drop to zero and purge flow increases from 15 to 30ml/hr, indicating a leak. a bag change took place during this event. the pump was returned and inspected. the leak reproduced near the yellow luer on the sidearm. foreign material thought to be the field applied dermabond coated the sidearm nozzle that connects to the yellow luer. further cleaning and imaging showed a crack down the length of the nozzle where the pr luer connects with the sidearm. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc., per supplier evaluation.
8157 | 1220648-2025-29665 | 22303814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303814&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666130 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma at the access site. the patient was transfused with four (4) units of packed red blood cells. the patient was very volume dependent with suction alarms at p2. a sandbag was placed on the hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: patient arrived from satellite hospital with right groin hematoma at impella site. the cause of access site bleeding issue most likely was patient movement related as the hematoma was noted after patient transferred hospitals. low pump: the cause of low pump flow issue most likely patient condition related since suction alarms resolve and flows in range after blood is given.
8158 | 1220648-2025-29666 | 22303820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303820&pc=OZD | 06/05/2025 | 12/02/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351867 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) failed system check upon startup. the aic was removed from service. patient involvement was not reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported issue has been completed. data analysis: on june 11, 2025, the console logs recorded a continuous powermod_1 communication issue. subsequently, the "system self check failed" screen was displayed. device analysis: this console was tested after arriving at fs. during bootup, the console rebooted automatically and displayed the "system self check failure" screen. visual inspection confirmed pbm contamination, which has impacted a neighboring (b)(6) communication channel. root cause: the root cause of the controller error was determined to be pbm corrosion resulting from electrolyte leakage caused by battery failure alarm super caps. h6 investigation type, findings and conclusions. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
8159 | 1220648-2025-29667 | 22303834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303834&pc=OZD | 06/05/2025 | 07/17/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572770 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient presented with worsening shortness of breath and exercise intolerance. decision was made to place impella 5.5 pre-op for scheduled coronary artery bypass grafting. however, impella 5.5 was placed through 10 mm hemashield platinum graft attached to the right axillary artery. multiple attempts made to insert impella 5.5 over .018 wire but unable to track and pass the motor housing around curve back in to ascending aorta. decision then made to abort the 5.5 and insert impella cp through the existing 10mm graft. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The root cause of delivery issue was most likely patient condition since patient had severe multivessel cad and ischemic cardiomyopathy.
8160 | 1220648-2025-29668 | 22303838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303838&pc=OZD | 06/05/2025 | 07/18/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666906 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was unable to be implanted due to the patient¿s anatomy. impella placement was aborted. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8161 | 1220648-2025-29669 | 22303841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303841&pc=OZD | 06/05/2025 | 07/18/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660925 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was attempted to be inserted into the axillary artery but unable to traverse the anatomy into the left ventricle. the 5.5 was removed and cp was placed. the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the impella 5.5 was attempted to be inserted into the axillary artery but unable to traverse the anatomy into the left ventricle.
8162 | 1220648-2025-29670 | 22303854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303854&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657595 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump began to experience ongoing suction issues following implant. the pump was pulled back in an attempt to reposition the device, however the physician was unable to recross the aortic valve. the decision was ultimately made to replace the device due to the noted issue. | [5 empty] | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. positioning issues: the cause of positioning issues most likely was use related due to improper technique since impella was pulled back into aorta during repositioning. physician was unable to advance impella across av. low pump flow; data log reviewed, no motor current spike or placement signal issue observed. low flows observed when pump ran on p7. suction occurred. impella position in aorta alarms seen most likely related to repositioning attempts by physician, it looks flow dropped for while and then back to flow range at p2. drop in flows observed most of impella support period. patient had mitral valve disease (mvd). the cause of low pump flow most likely was patient condition related due to mitral valve disease. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8163 | 1220648-2025-29671 | 22303868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303868&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678651 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented for a planned percutaneous coronary intervention (pci), however during pci perforation was noted to mid left circumflex artery. the patient decompensated requiring cardiopulmonary resuscitation, epi, pericardiocentesis and impella cp placement. covered stent was placed to lm. the impella flows were suboptimal and impella in ventricle alarm was noted. decision was made to remove impella. the patient outcome is survived. | [5 empty] | The investigation of low pump flow has been completed. no product was returned. log review found the placement signal trend declining since the start of the case and low triggering ps low alarm. since the pump start few suction alarms were seen with relative to the event timeline. also, impella in ventricle alarms noted few minutes after suction and pump run. no other abnormalities to motor current trend was seen. clinical also states that the patient was experiencing poor condition after perforation which occurred prior to impella placement and the patient requiring pericardiocentesis procedure which likely contributed to suction issues during impella support. the root cause of the issue was most likely patient condition related since placement signal was low when the pump was started and there was a perforation event that caused blood loss. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8164 | 1220648-2025-29672 | 22303874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303874&pc=OZD | 06/05/2025 | 06/21/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642084 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient presented with chest pain & pressure. the patient acutely decompensated and decision to place impella cp was made, followed by peripheral veno-arterial extracorporeal membrane oxygenation (va ecmo). it was reported that after impella insertion, the patient went into ventricular tachycardia arrested requiring multiple shocks and peripheral va ecmo cannulation. during support it was noted of a rp bleed on a ct scan that was related to ecmo cannulation. the iliac vein was stented due to venous bleed. the patient was heparinized. in addition, the patient was turned down to p1 due to hemolysis. the cp was explanted and as a escalation to the 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8165 | 1220648-2025-29673 | 22303882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22303882&pc=OZD | 06/05/2025 | 07/18/2025 |  | 06/21/2025 |  |  | 06/05/2025 | 06/21/2025 |  |  |  |  | 06/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664311 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had oozing at the impella site. new forward tension done and existing perclose was adjusted. on day two of impella support, impella site was sutured, and site was re-dressed. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be use issue because hemostasis was achieved through new forward tension at the bleeding site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8166 | 1220648-2025-29597 | 22296199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296199&pc=OZD | 06/14/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/14/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592344 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a physician had difficulty getting the impella 5.5 across the aortic valve. the physician initially attempted wireless, then tried multiple wires and catheters. however, eventually went back to wireless and was able to get across the valve. the impella went straight across the valve into the left ventricle (lv) lateral wall. at this point, the physician noticed a lv perforation and removed the impella 5.5. the perforation was repaired. the decision was made not to go back in with the impella. the doctor is unsure if one of the guide wires caused the perforation and the impella made it worse or if it was the impella. the patients outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8167 | 1220648-2025-29598 | 22296240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296240&pc=OZD | 06/16/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/16/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611682 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella cp. there was a purge pressure low alarm immediately after the pump was started. the purge system was replaced to no avail. the pump was replaced after percutaneous coronary intervention was completed. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿.
8168 | 1220648-2025-29599 | 22296372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296372&pc=PYX | 05/24/2025 | 07/21/2025 |  | 06/20/2025 |  |  | 05/25/2025 | 06/20/2025 |  |  | 01/31/2026 |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025458224 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was suffering from failure and had a ross surgery performed and impella 5.5 implanted. the patient one day post operative was found to be having right ventricle dysfunction coupled with the patients continued increasing pressor requirement. the team opted to escalate to an impella rp flex. the patient currently was at p-7 on the impella 5.5 rp and was being weaned for explant. small cerebral hemorrhaging was noted on imaging so heparin was stopped and it was decided that impella rp flex would be explanted. the patient remained on levophed and vasopression with no issues overnight. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the patient's outcome was reviewed by abiomed's medical safety officer (mso). the patient on bipella support (impella rp flex and impella 5.5) and experienced complications. the impella rp flex event is a serious injury. however, there is not enough evidence to exclude the impella 5.5 device as an associated factor to the patient's outcome given related events of bleeding and low pump flow. the demise outcome was captured under the impella 5.5 manufacturer mdr report #1220648-2025-29415. | The investigation of cva/stroke has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 health effect - clinical code 1770 and health effect - impact code 4627 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29599.
8169 | 1220648-2025-29600 | 22296418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296418&pc=OZD | 05/29/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/29/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657701 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient needed several rounds of defibrillation and pacing wires were placed. additionally, the patient developed a clot in their chest cavity. the clot was found in the back of heart that caused patient decompensation. the clinical representative confirmed that this was not an impella related issue and support was continued. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8170 | 1220648-2025-29601 | 22296465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296465&pc=OZD | 05/29/2025 | 07/23/2025 |  | 06/20/2025 |  |  | 05/29/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654213 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | Fever | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient became febrile with sepsis and had a new onset of atrial fibrillation. the patient was treated with amino. additionally, there was a decrease in purge flow and an increase in purge pressure. tissue plasma activator was infused in the purge system and support was continued. | [5 empty] | The investigation of high purge pressure, vf/vt/af/at, and infection/sepsis has been completed. the impella device was not returned for evaluation. high purge pressure: data logs were reviewed and a increase in purge pressure over 12 hours before "high purge pressure" alarms were triggered. purge pressure remained high for approximately 6 hours before purge pressure began to return to normal values. clinical details noted that purge flows were decreased with an increase in purge pressure were noted on support and tpa protocol was initiated. the root cause of the high purge pressure was most likely due to biomaterial buildup based on data log analysis and clinical details noting that tissue plasminogen activator (tpa) resolved the alarms. vf/vt: the root cause of the tachycardia and a-fib was unable to be determined due to insufficient clinical details. infection/sepsis: the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. h6 medical device problem code 1491 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29601. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8171 | 1220648-2025-29602 | 22296532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296532&pc=OZD | 05/29/2025 | 07/17/2025 |  | 06/20/2025 |  |  | 05/29/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639743 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient had access site bleeding and several other sites including the patient's mouth and nose. a femstop was placed at the access site and blood products were provided to the patient. heparin was used in the purge and therefore, the impella support could be related to the patients bleeding from non-impella sites. the patients outcome is stable. | [5 empty] | Vascular damage- peripheral vessel : the root cause of the vascular damage issue was most likely patient condition related, the patient had access site bleeding and several other sites including the patient's mouth and nose. access site bleeding : the root cause of the access site bleeding issue was most likely patient condition related, the patient had access site bleeding and several other sites including the patient's mouth and nose. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8172 | 1220648-2025-29603 | 22296587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296587&pc=OZD | 05/29/2025 | 07/23/2025 |  | 06/20/2025 |  |  | 05/29/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemolysis | Perforation of Vessels | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp insertion, there was initially difficulty crossing the aortic valve (aov) with the wire into the left ventricle (lv) due to severe aortic stenosis. as the physician attempted to push across, the impella cannula kinked and could not move further across the aov. on fluoroscopy, the outflow was just across the valve. the physician then performed single access through the14 french peel-away to get a wire across and attempted valvuloplasty for better delivery of the impella across the valve. as the physician managed to get balloon across the aov, it pushed the impella cp in further to a better position and unkinked the cannula. wires and balloon were removed and the impella was started with good flows. additionally, upon arrival at the intenstive care unit, hemolysis was noted in the foley catheter. the impella p-level was turned down. the next day, the impella was repositioned and urine started to clear. it was further reported that the patient started displaying signs and symptoms of a gastrointestinal bleed. the family made the decision to withdraw care, and the patient expired. | [5 empty] | The investigation of product damage (cannula kink), hemolysis, and vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. the data log did not record any abnormalities related to pump flow, motor current, or positioning issues. some discrepancies in the placement signal were noted without affecting the pump's performance during the hemolysis event. the data log was not conclusive without the returned product investigation and sufficient clinical information to derive the root cause of the hemolysis. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the cannula kink issue was most likely patient condition related to severe aortic stenosis. the cause of the vascular damage issue was not determined since sufficient product was not returned and sufficient clinical information was not provided. the relation to impella and gastrointestinal bleed is unknown. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8173 | 1220648-2025-29604 | 22296645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296645&pc=OZD | 05/29/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/29/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622567 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient came in with a flutter and an ejection fraction of 5%. the impella 5.5 was supporting when ventricular tachycardia runs occurred when the patient leaned forward or sometimes turned and torsades occurred as well. shocking was required. the pump supported for over ten days when care was withdrawn and the pump was explanted. the patient went on comfort care and expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8174 | 1220648-2025-29605 | 22296711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296711&pc=OZD | 06/15/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/17/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026623400 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed to support an acute myocardial infarction patient. the pump failed when priming/introducing the pump for support. the team confirmed leaks from purge and so the pump was replaced. a different cp was used for support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
8175 | 1220648-2025-29606 | 22296712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296712&pc=OZD | 05/31/2025 | 07/18/2025 |  | 06/20/2025 |  |  | 05/31/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674836 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, following placement of an impella cp repositioning sheath, hemostasis could not be achieved due to the patient's body habitus. attempts to tighten the previously placed perclose sutures were unsuccessful and a hematoma developed at the site. the pump was removed in response to the condition and a 14 french cook sheath was placed to achieve hemostasis. against advisement, the physician re-inserted the pump through the newly placed sheath for patient support. the patients outcome is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of the access site bleeding issue was most likely use related, 14fr cook sheath was placed to achieve hemostasis., the md re-inserted the pump through the newly placed sheath for patient support successfully.
8176 | 1220648-2025-29607 | 22296753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296753&pc=OZD | 06/07/2025 | 09/18/2025 |  | 06/20/2025 |  |  | 06/10/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661727 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complainant reported the patient was on impella 5.5 support and during support, the patient had multiple cerebral infarctions. the relationship between the cerebral infractions and the impella is unknown. the patient remains on support. | It was further reported that the patient expired from multiple organ failure. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | B2 death and date of death added based on additional information received. b5 patient outcome added. h1 type of report updated to death. h6 health effect - impact code revised to death.
8177 | 1220648-2025-29608 | 22296845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296845&pc=OZD | 05/31/2025 | 07/17/2025 |  | 06/20/2025 |  |  | 05/31/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669211 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, there was high purge pressure. the purge cassette was changed to no avail. it was advised to use tissue plasminogen activator and support continued. eventually the patient was escalated to an impella 5.5. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Thrombus added due to ingestion pattern observed in log analysis. high purge pressure: no product was returned for analysis. the data logs show a motor current spike on 5/31 at about 6:25am about 69 hours into support. at time of spike there is a motor speed deviation with a rise in purge pressure and drop in purge flows. purge pressures do not meet alarm trigger but remain elevated for about 10 hours before decreasing and returning to prior values. it is unknown if tpa was added to the purge. the pattern observed in the logs is characteristic of biomaterial ingestion. the root cause of the high purge pressure was most likely biomaterial ingestion as evidenced by data log pattern thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc.. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
8178 | 1220648-2025-29609 | 22296895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296895&pc=OZD | 04/29/2025 | 07/08/2025 |  | 06/20/2025 |  |  | 04/29/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544988 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 75 KG | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, there were placement signal not reliable alarms. audio was disabled and the staff monitored. additionally, a pump stop occurred with motor spikes. the automated impella controller (aic) was switched out and the pump stopped several times, but was successfully restarted. | The user facility reported the patient was on impella 5.5 support and after the pump was implanted, the patient had ectopy. the patient also had ventricular tachycardia after the pump was adjusted to p 8. the patient self converted and was eventually put on amio. support continued. |  |  |  |  | B.1 product problem was added based on the additional information received. b.5 additional information was received. d.6b explant date was added. h.6 medical device problem code was added to reflect additional failure mode reported. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smart assist) section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.4 weight was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-29609. e.1 email address was incorrectly reported on the initial manufacturer device report number 1220648-2025-29609 and the phone number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-29609. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8179 | 1220648-2025-29610 | 22296920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296920&pc=OZD | 05/31/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/31/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1285251 |  | 0042-0000-US-L |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) error alert with cockle prior to procedure. the aic was replaced with another aic. no patient involvement was noted. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8180 | 1220648-2025-29611 | 22296930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22296930&pc=OZD | 05/28/2025 | 07/15/2025 |  | 06/20/2025 |  |  | 05/28/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645700 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis during impella cp support. treatment for the condition was unknown and a resolution of the event was not confirmed. additionally, there were alarms of high purge pressure. the purge cassette was changed but the alarms continued. the patient survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of high purge pressure and hemolysis has been completed. the impella device was not returned for evaluation. hemolysis: the clinical states that there was hemolysis reported on may 28th, and treatment and resolution of the event was not confirmed. the data logs show that on 5/25 there was a spike in motor current (mc) and an increase in purge pressure (pp) and pump flows. there was also an increase in placement signal (ps) and suction alarms. the purge pressure and mc remain elevated for about 42 hours leading up to the 28th, until there is a drop in mc and flows with no p-level change. based on the data logs, the root cause of the hemolysis is most likely due to ingested biomaterial. high purge pressure: the data logs show that on 5/25 there was a spike in mc and an increase in pp and pump flows. there was also an increase in ps and suction alarms. the purge pressure and mc remain elevated for about 42 hours leading up to the 28th, until there is a drop in mc and flows with no p-level change. purge pressure remains stable for about 145 hours until on 6/2, there is a 2 minute rise in purge pressure to trigger "high purge pressure" alarms. right before the rise there is a slight dip in motor current and ps but no motor current spikes. the pump is explanted shortly after. due to lack of clinical information and no product returned, the root cause of the high purge pressure cannot be determined. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined g1 reporting contact first and last name were erroneously entered on the initial manufacturer device report number 1220648-2025-29611. g1 manufacturing site, name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-29611.
8181 | 1220648-2025-29612 | 22297091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297091&pc=OZD | 05/31/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/31/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636020 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient had ischemia on the impella leg. an extracorporeal femorofemoral bypass was performed and support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8182 | 1220648-2025-29614 | 22297245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297245&pc=PYX | 05/31/2025 | 07/30/2025 |  | 06/20/2025 |  |  | 05/31/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025602752 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Sense | Incomplete or Inadequate Priming | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while attempting to setup and prime an impella rp flex, the pump failed to automatically zero itself appropriately and presented pressures of 90s/90s. the case was canceled and restarted with the console skipping a process of the setup and the insertion by going directly to the placement screen. the impella rp flex was aborted and was subsequently replaced in response to the issue. no patient harm was reported. | [5 empty] | The investigation of priming issue has been completed. the impella device was not returned for evaluation.based on the following investigation, the failure mode of "priming issues" was changed to "placement signal issue" as the pump did not have any issues with priming, issue was only related to the pump sensor. no product was returned for evaluation. data logs were reviewed and rt log of initial pump connection showed all optical 5s parameters were within a normal range, no hard failures of the optical sensor were observed. ps readings outside of the body were noted to be around 20 with pap readings of around 30. motor current and pump flow remained 0 through log. clinical details noted that when priming pump, the pump would not zero itself appropriately and the pressures were reading 90s/90s outside of the body. the root cause of the placement signal issue could not be definitively determined as no product was returned for evaluation and no hard failures of the pump were observed on returned data logs. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8183 | 1220648-2025-29613 | 22297246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297246&pc=OZD | 05/28/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/28/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1631552 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) had a battery failure. there was no patient involvement. | [5 empty] | The console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. additionally, no battery related issue were seen in the logs. visual inspection confirmed pbm contamination. the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. after the pbm and the optical bench were replaced preventative maintenance and a full functional check were performed. the console was then returned to the customer.
8184 | 1220648-2025-29615 | 22297287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297287&pc=OZD | 05/28/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/28/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600035 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp placed to support a patient with known cardiac history and new st-segment elevation myocardial infarction. while on impella cp support, the patient had dark red emesis and a drop in hemoglobin. one unit of packed red blood cells was given. the procedural outcome was the patient survived the surgery and was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8185 | 1220648-2025-29504 | 22297300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297300&pc=OZD | 05/30/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 05/30/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669020 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male | 87 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp developed a hematoma at the access site. the pump was removed from the right femoral artery and replaced on the left femoral artery. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8186 | 1220648-2025-29616 | 22297524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297524&pc=OZD | 06/03/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/03/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1658143 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller (aic) would not recognize the purge cassette. this issue occurred during patient use. no further details were provided. there was no reported patient harm. | [5 empty] | The investigation for the aic purge disc/flag issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) several times. the console was returned for evaluation. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
8187 | 1220648-2025-29617 | 22297754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22297754&pc=OZD | 06/04/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/04/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026677158 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient implanted with an impella cp experienced distal ischemia in the right leg. a distal reperfusion catheter was placed in the right leg to reestablish flow. the patient expired on support. there is not enough evidence to exclude the impella cp device as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8188 | 1220648-2025-29618 | 22298163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22298163&pc=OZD | 06/01/2025 | 07/18/2025 |  | 06/20/2025 |  |  | 06/01/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607130 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and prior to leaving the procedural area a hematoma was noted. the patient was transferred to a higher-level care facility, and once settled, the hematoma worsened. consequently, the impella cp device was explanted, and the patient received three units of red blood cells for the blood loss. a repeat echocardiogram was completed and noted the left ventricle recovered "significantly" in comparison, and the decision was made to hold off on escalation to an impella 5.5 post explant of the impella cp. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient movement because the hematoma the stabilized hematoma became worse after transfer by helicopter.
8189 | 1220648-2025-29619 | 22298993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22298993&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665418 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male | 109 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a plasma free hemoglobin value of 340. the pump was noted to be in the papillary and was repositioned to resolve the issue. impella support continued. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. hemolysis: the cause of hemolysis was most likely improper positioning as plasma free decreased after first pump repositioning and then days later was elevated again with pump in incorrect position. positioning issues: the cause of positioning issues was not determined due to lack of clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
8190 | 1220648-2025-29620 | 22299167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22299167&pc=OZD | 05/07/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/17/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564653 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complainant reported the patient was on impella 5.5 support and during support, the pump became mispositioned so the pump was removed and readjusted, but there were many blood clots in the pump. the pump was replaced. the patient expired on support of the replacement pump. the patient died without any improvement in circulatory dynamics. no problems with the operation of the replacement device. the patient died due to multiple organ failure. there is not enough evidence to exclude the impella 5.5 device used as an associated factor to the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8191 | 1220648-2025-29621 | 22299372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22299372&pc=OZD | 06/05/2025 | 09/03/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622208 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male | 85 KG |  | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp had a mitraclip procedure performed. following the procedure, the flow of the impella dropped and the pump stopped. blood volume was given and the introducer was flushed however, the issue did not resolve. the impella cp was explanted and replaced with a new impella cp. | [5 empty] | The investigation of low pump flow and pump stop has been completed. data logs were not available. the pump was returned and evaluated. the pump passed electrical testing of the motor phases. there was some biomaterial found on the impeller of the returned pump, but no other abnormalities were found. the pump ran for ~20 minutes in the lab and did not reproduce low pump flow or a pump stop. there were no other abnormalities found with the pump. low pump flow: the cause of the low pump flow was unable to be determined since data logs were not available. pump stop: the cause of the pump stop was unable to be determined since data logs were not available. | The impella device was received from the customer. upon completion of the investigation, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock. section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
8192 | 1220648-2025-29622 | 22299527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22299527&pc=PYX | 06/05/2025 | 07/22/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026619218 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male | 100 KG | Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella rp flex exhibited low pump flow and suction events that could not be resolved. additionally, there were high purge pressure alarms. the pump was explanted and a protek duo was used to support the patient. | [5 empty] | The investigation of low pump flow and high purge pressure has been completed. product damage (catheter kink) was added based on the returned product. low pump flow: the data logs show that from about 5/25 there were many suction alarms and the p-level was changed many times to try to combat the suction. the ps was relatively flat and the cvp and pap was trending down from 5/25 till about 6/5. there were no spikes in motor current or significant changes in purge pressure seen during this time. the pump was returned and there were no kinks in the cannula seen or thrombus. low pump flow was not able to be reproduced. there was a significant kink in the catheter near the motor housing. based on the clinical details and data log analysis, the root cause of the low pump flow is most likely due to the patients condition. high purge pressure: the data logs show that there were three instances of "high purge pressure" alarms triggered. the first occurrence was on 5/16, where the purge pressure spiked to 1100 over the course of two minutes and then resolved almost instantly. there were no changes to motor current or flow during this time. the next occurrences that had the same pattern were on 5/22 around almost midnight. throughout support there were many instances of these spikes but they never spiked high enough to trigger more alarms. these instances did not occur around the times of bag/cassette changes. the product was returned and there was no biomaterial in the purge gap. the cassette was seen to have two knots in it which were likely due to packaging of it, as purge flow can not flow through it if there is tied. the pump was purged with the cassette and high purge pressure was not reproduced. a significant kink was seen in the catheter near the motor housing. likely, the high purge pressure that occurred during the case was due to the kink in the catheter and dependent on it being held in a kinked position, or during times of repositioning or movement. based on the data logs and returned product, the root cause of the high purge pressure is most likely due to a kinked purge line although the cause of the kinked line is unknown. product damage (catheter kink): the pump was returned with a kink in the catheter near the motor housing. it is unknown how the kink occurred. the root cause of the product damage (catheter kink) cannot be determined due to lack of clinical details. a2 patient age was erroneously entered on the initial manufacturer device report number 1220648-2025-29622. a4 patient weight was erroneously entered on the initial manufacturer device report number 1220648-2025-29622. a5 ethnicity and race were erroneously selected on the initial manufacturer device report number 1220648-2025-29622. d6a implant date was erroneously entered on the initial manufacturer device report number 1220648-2025-29622. g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2025-29622. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
8193 | 1220648-2025-29623 | 22299545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22299545&pc=OZD | 05/31/2025 | 07/18/2025 |  | 06/20/2025 |  |  | 05/31/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608642 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Malposition of Device |  |  |  |  | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device (pump 1) for mechanical circulatory support and very soon thereafter, the patient had the pump stop twice with the pump "diving" in the ventricle. subsequently, there was continuous suction and lower than expected pump flow. consequently, the patient was taken back to the operating room for the impella 5.5 (pump 1) explant and replacement with a new impella 5.5 (pump 2) and extracorporeal membrane oxygenation (ecmo) cannulation. around this time, it was noted the patient experienced an ischemic stroke. clot ingestion was suspected. the physician noted likely clot was in left ventricle and undetected. latest update noted the patient remains on impella 5.5 (pump 2) mechanical circulatory support. | [5 empty] | The investigation of positioning issues, temporary pump stop, low pump flow, thrombus, and stroke/cva has been completed. the impella device was not returned for evaluation. clinical details report that the pump stopped twice on the first day of support, shortly after moving the patient. when the patient was being moved, the pump dove into the ventricle. at the time of the pump stop, the patient's activated clotting times (act) were reportedly not in the target range, though no value was provided. the pump was able to be restarted, but flows were lower than expected and the pump experienced continuous suction, so it was explanted. at explant, it was reported that biomaterial was visualized at the outflow cage. temporary pump stop: data logs were reviewed, and the temporary pump stop was confirmed less than 1 hour into support. leading up to the pump stop, there were several spikes/disturbances in the ps that aligned with drops in mcmin values along with an active suction alarm. then, several rapid motor current (mc) spikes to 1200 ma along with ms deviations occurred, resulting in the pump stop. the pump was able to be restarted, but flows were below specification, and there was another instance of a mc spike with ms deviation later in the case. this second mc spike did not trigger a pump stop, but was followed by an impella position in aorta alarm. all of these signatures in the data logs indicate that biomaterial had gotten into the cannula, initially interacting with the optical sensor and causing low mc min values (diastolic suction) leading up to the pump stop. then, the biomaterial impacted the impeller, spiking the mc required to maintain constant rpms, and stopped the pump. when the pump was restarted, the biomaterial was still occluding the cannula/outflow cage, so flows remained below specification. the second mc spike event is an indication that the biomaterial impeded the impeller's rotation a second time, and the trigger of the position in aorta alarm was likely a false trigger due to how the biomaterial interacted with the impeller. this is all supported by the clinical details provided that the patient's acts weren't in target range and a clot was visualized in the outflow cage of the pump upon explant. product was not returned for investigation. based on the clinical details provided and the signatures noted in data logs, the root cause of the temporary pump stop was most likely biomaterial ingestion. low pump flow: data logs were reviewed and the reported low pump flows were confirmed. there was a high motor current pump stop midway through the case that exhibited the typical signatures of biomaterial ingestion. shortly before the pump stop, pump flows had already become low and a suction alarm had triggered. during this time, there were several spikes in the ps that aligned with drops in mc minimum values, indicating that biomaterial was likely occluding flows and interacting with the optical sensor prior to the pump stop. when the pump was restarted, flows remained low and suction alarms triggered again because the mc minimum values were low relative to the ps, indicating a suction condition. shortly after restarting the pump, another mc spike with speed deviation occurred. this all supports that the biomaterial was likely still in the cannula/outflow cage, occluding flows. clinical details further support this conclusion, reporting that thrombus was noted in the outflow cage of the pump upon explant. product was not returned for investigation. based on the clinical details provided and the signatures noted in data log review, the root cause of the low pump flows was most likely biomaterial. positioning issue: since clinical details reported that the patient was being moved when the pump went too deep into the ventricle, the root cause of the positioning issue was determined to most likely be a use issue. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. | Clarification, additional information was obtained regarding the timing of the stroke and noted the following: it is unknown exactly when the stroke occurred. the impella 5.5 pump 1 was exchanged out with impella 5.5 pump 2 and ecmo just a couple hours later. when the patient was moved the impella 5.5 pump 1 dove into the ventricle as noted on impella connect. sedation had not been weaned at this time; therefore, neurological changes had not been noted. this is one of two devices associated with the event. refer to initial medical device report for impella 5.5 pump 2 serial number (b)(6) with date of event 31-may-2025 and identifier ending in (b)(6). the impella 5.5 pump 1 device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8194 | 1220648-2025-29624 | 22299727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22299727&pc=OZD | 06/05/2025 | 07/16/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661928 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that following impella cp implant, the motor current on the implanted pump was over 1100 in p-auto with a mean flow of 4.3lpm and about 200ma above normal when weaning the device. in addition, a high purge pressure alarm appeared during support at which point the purge cassette was changed to resolve the issue. the motor current remained high despite the purge cassette replacement and approached the pump failure / stop threshold. the patient was successfully weaned from the device and survived. | [5 empty] | H6 medical device problem code 1491 was inadvertently omitted on manufacturer device report number 1220648-2025-29624-1. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Saturated flow: the returned product was inspected and biomaterial was observed blocking the purge gap, by the impeller, and around the outlet. the pump was soaked in tergazyme and after clearing biomaterial, the pump was run in the lab. the saturated flow did not reproduce. the data logs from the case show a gradual rise in purge pressure with a simultaneous decrease in purge flow. the logs show the pump flow average was 4.3l/min, higher than expected at p9. purge pressure continued to rise to reach 1000s before reducing to normal levels. the root cause of the saturated flow was determined to be biomaterial as evidenced by biomaterial observed on returned product. high purge pressure: the returned product was inspected and biomaterial was observed blocking the purge gap, by the impeller, and around the outlet. the pump was soaked in tergazyme and after clearing biomaterial, the pump was run in the lab. the high purge pressure did not reproduce. the data logs from the case show a gradual rise in purge pressure with a simultaneous decrease in purge flow. there were no motor current spikes. purge pressure continued to rise to reach 1000s before reducing to normal levels. clinical reported purge cassette was swapped and issue resolved. act reported to be 250. the root cause of the high purge pressure was biomaterial as evidenced by biomaterial blocking purge gap on returned product and issue not reproducing after clearing biomaterial. thrombus: in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. priming issues: purge cassette changed in an attempt to resolve high purge pressure. the returned purge cassettes had no visual abnormalities. the purge cassettes were both de-aired without issue and were able to purge fluid through the pump with no issue. the data logs do not show any abnormalities. no problem was found as the root cause of the priming issue as product functioned as expected and no log abnormalities were observed.
8195 | 1220648-2025-29625 | 22299764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22299764&pc=OZD | 06/04/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/04/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665530 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male | 64 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited hemolysis and hematuria. correct positioning of the pump was verified and decreasing the p level of the pump did not resolve the issue. the pump was explanted to resolve the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
8196 | 1220648-2025-29626 | 22300064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22300064&pc=OZD | 06/05/2025 | 07/02/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664328 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 64 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed oozing at the impella access site. the patient was transfused one unit of blood products and a mattress suture was applied. the patient remains on impella support. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D6: no explant date available the root cause of access site bleeding was most likely use issue since hemostasis was achieved by placing mattress suture.
8197 | 1220648-2025-29627 | 22300422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22300422&pc=OZD | 06/05/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 111 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp experienced a decrease in platelet count. the patient had just begun to receive heparin. the heparin was replaced with argatroban to resolve the heparin induced thrombocytopenia. the patient decompensated shortly after and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8198 | 1220648-2025-29628 | 22301279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301279&pc=OZD | 06/05/2025 | 07/18/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 163 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during implantation of an impella cp the pump did not cross the innominate artery secondary to the 90 degree bend. the implant was aborted and a new impella 5.5 was implanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the pump did not cross innominate secondary to 90 degree bend.
8199 | 1220648-2025-29629 | 22301419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301419&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656171 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 80 KG | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while explanting an impella cp from a patient the standard j-wire was present in the peel-away sheath which made it difficult to remove. the explant was completed with the wire still in peel-away sheath. the patient survived. | [5 empty] | The investigation of difficulty inserting/removing pump through introducer and product damage (guidewire damage - peripheral vessel) has been completed. difficulty removing pump through introducer: product was not returned for investigation. since clinical details reported that the diagnostic guidewire was still in the introducer at the end of the procedure and was the cause for resistance when attempting to remove the pump, the root cause of the difficulty removing pump through introducer was determined to most likely be a use issue. product damage (guidewire damage): product was not returned for investigation. clinical details reported that the guidewire was still in the introducer when attempting to remove the pump. when trying to then remove the guidewire, it unraveled inside the introducer sheath. based on the clinical details provided that the guidewire unraveled when trying to remove the pump while the guidewire was still in the introducer, the root cause of the guidewire damage was most likely a use issue. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during impella section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8200 | 1220648-2025-29630 | 22301459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301459&pc=OZD | 06/04/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/04/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640233 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male | 70 KG | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was advanced over the wire and ran in auto. the performance dropped and the patient went into ventricular tachycardia (vt) arrest requiring cpr and shock. impella suction alarms are present. the surgeon attempted to reposition, however pulled the device across atrioventricular (av). the impella successfully advanced across valve, however required repositioning in cardiac care lab (ccl). the physician successfully repositioned the device toward apex with no issues. the decision was made to remove impella after successfully weaning over 24 hrs. the patient was responsive. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position. ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8201 | 1220648-2025-29631 | 22301503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301503&pc=OZD | 06/04/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/04/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674571 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there were false impella in aorta alarms. the position was confirmed on echo. placement signal monitoring alarm disabled. additionally, the patient had a large intercranial hematoma causing a midline line shift and active hemorrhaging. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8202 | 1220648-2025-29632 | 22301591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301591&pc=OZD | 06/05/2025 | 08/04/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678018 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male | 90 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following a supported procedure, the impella cp pump was inadvertently pulled back across the aortic valve. the pump was fully removed from the patient as a result of the issue and the patient was monitored before a decision was made on the next steps to be taken. it was concluded that support was still necessary, and as such, the pump was reinserted into the patient. the staff was advised on the risks of re-using the pump, but the decision was ultimately made to continue support with the current setup. the patient was successfully weaned from the pump and the patient survived. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. positioning issue: the root cause of positioning issue was most likely use related since the pump was accidentally pulled back into aorta during reposition of the pump. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8203 | 1220648-2025-29633 | 22301656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301656&pc=PYX | 06/05/2025 | 08/04/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659309 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that low pump flow occurred on an implanted impella rp flex. the pump was explanted and biomaterial was observed. the patient survived. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. low pump flow: the clinical states that there was low pump flow at p-8 right after starting the case. the act was 251, the pump was repositioned and there was biomaterial observed. the data logs show that right after starting impella support there were spikes in the motor current and ps followed by lower flows. the pump was run approximately for 13 minutes. based on the data logs, the root cause of the low pump flow is most likely due to ingested biomaterial but the cause of the biomaterial is unknown. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29633. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8204 | 1220648-2025-29635 | 22301763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301763&pc=OZD | 06/05/2025 | 10/01/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026659375 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported an impella rp was implanted in a patient for circulatory support. after the patient arrived at the icu unit, a couple of mattress sutures were added. it was further reported that the impella placement signal was not reliable. the team used the motor current to determine the flows. the patient was successfully weaned from the pump and the device was explanted. | The complainant reported that after the impella rp flex was implanted and the patient arrived at the icu unit, a couple of mattress sutures were added, impella site is within normal limits. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the fm optical signal issue was changed to ps issue based on the investigation indicating it is the dp sensor not the optical sensor. access site bleeding: the root cause of the access site bleeding is most likely due to use as adding mattress sutures resolved the bleeding. placement signal issue: the ps was stable around 40mmhg for around 17 hours until the ps increases to 3000 and flows drops. it then returns to normal. there are increases to unrealistic values while flows drops and comes back throughout the whole case. raw electrical shows the same. the same trend is present for pap and cvp, but all optical 5s parameters are stable as well as the raw optical signal. during these times alarms #1051 and #1052 are triggered. due to lack of product returned, the root cause of the placement signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding is most likely due to use as adding mattress sutures resolved the bleeding. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
8205 | 1220648-2025-29636 | 22301876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22301876&pc=OZD | 06/05/2025 | 06/20/2025 |  | 06/20/2025 |  |  | 06/05/2025 | 06/20/2025 |  |  |  |  | 06/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666762 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had a left ventricle (lv) thrombus. additionally, the pump measured at 2.7cm and was rotated just under the mitral valve. the doctor did not want to reposition due to the lv thrombus. care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8206 | 1220648-2025-29510 | 22289080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289080&pc=OZD | 05/23/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/23/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631341 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient experienced five minutes of ventricular tachycardia (vt)/ventricular fibrillation (vf). the patient was defibrillated six times and was given amiodarone and lidocaine boluses. support was continued for approximately a week or so. care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8207 | 1220648-2025-29511 | 22289092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289092&pc=OZD | 06/03/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626853 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that while implanted with an impella 5.5, the patient experienced several dysrhythmias and was defibrillated three times. several days later, the patient experience ventricular tachycardia (vt) requiring being shocked. four days later, the patient had repeat vt and was cardioverted. the impella 5.5 continues to support the patient. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8208 | 1220648-2025-29515 | 22289153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289153&pc=OZD | 06/02/2025 | 07/30/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647614 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | EXTRACORPEAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the implanted impella cp position was shallow and was difficult to reposition. the patient developed hemolysis which was believed to be related to the position. the patient was successfully weaned from the pump and survived. extracorporeal membrane oxygenation was continued for oxygenation support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and positioning issues has been completed. the pump was not returned for investigation. the total cassette run time was 81 hours. the data log shows several suction alarms and an impella position in aorta alarm until the middle of the case run, when placement signal monitoring was disabled. pump flow and motor current exhibited a fluctuating upward trend. according to the clinical notes, the pump position was also compromised during the hemolysis event. the cause of the hemolysis issue was most likely improper positioning of the device, based on given clinical details and data log review. the cause of the positing issue was not determined since product was not returned and sufficient clinical information was not provided.
8209 | 1220648-2025-29518 | 22289199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289199&pc=OZD | 05/26/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663350 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and that the pump was positioned deep with the pigtail folding on itself. the pump was retracted to the aortic valve. the patient developed hemolysis overnight and one liter bolus was given which allowed the pump flow to be increased. the hemolysis continued to linger from the poor positioning that was corrected. the patient continued to deteriorate which lead to a discussion with the family deciding to transition to comfort care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
8210 | 1220648-2025-29520 | 22289220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289220&pc=OZD | 04/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a med science liaison forwarded poster abstract presented on 4/29/2025 at the ishlt. the poster is entitled "mini oral 07: impact of impella support on the progression of aortic insufficiency following lvad implantation" with author dr. (b)(6). following 270 heartmate3 patients between 2015 and 2023, 35 of which had impella prior to left ventricular assist device (lvad). the study found that impella support is associated with aortic insufficiency and may warrant consideration for prophylactic aortic valve intervention at the time of lvad implantation. | [5 empty] | D4: model number, serial number, catalog number and lot number are unknown. d6a & d6b implant and explant dates are unknown. h4: device manufacture date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8211 | 1220648-2025-29521 | 22289247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289247&pc=OZD | 05/30/2025 | 07/26/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656927 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was supporting despite all underlying surgical bleeding and transfer to the tertiary center. after nine (9) days the pump stopped and was explanted. | [5 empty] | Temporary pump stop: no product was returned for evaluation. data logs were reviewed and two instances of "impella stopped - motor current high" alarms were triggered on last day of support. the first pump stop instance showed a motor current spike with speed deviation, fully saturated flows, and a shift down in placement signal. purge pressures decrease slightly and purge flows increase. 5 minutes after the motor current spike occurs, an "impella stopped - motor current high" alarm is triggered. purge pressure increases rapidly with a "purge pressure high" alarm being triggered and drop in purge flows. pump is attempted to be restarted multiple times before pump is able to be restarted without immediate pump stop. purge pressures and purge flows normalized prior to pump being able to restart. at the time of the second instance of the pump stop alarms, motor current and pump flow lost all pulsatility with an "impella in aorta" alarm just before a motor current spike and speed deviation occurred and then a pump stop that was able to be be immediately restarted. pump was run for approximately 5 more minutes before being turned off and explanted. clinical details noted that the pump stop occurred on day 9 of support and the patient had not been on anticoagulation for several days. the cause of the temporary pump stop was most likely biomaterial since saturated flows and false position in aorta alarms resulted from the high motor current event. refer to es2020-082 for additional biomaterial pattern details. fm saturated flow was added since high flows were observed in the data logs following the temporary pump stop event. the cause of the saturated flow was most likely biomaterial since a temporary pump stop and false position in aorta alarm occurred concurrently with the saturated flow. refer to es2020-082 for additional biomaterial pattern details. fm optical signal issue was added since position in aorta alarms were observed in the data logs following the temporary pump stop. the cause of the false position in aorta issue was most likely biomaterial since a high motor current pump stop occurred prior to the issue. refer to es2020-082 for additional biomaterial pattern details. fm thrombus was added since the root cause of the temporary pump stop was ingested biomaterial. in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
8212 | 1220648-2025-29524 | 22289257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289257&pc=OZD | 05/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026673859 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had oozing and a hematoma at the access site. the dressing was changed, the site angle matched, forward tension suture still intact and femstop. heparin was withheld. the plan was to remove cp prior to notice as patient was to receive an automatic implantable cardioverter defibrillator (aicd). the impella was successfully weaned and explant. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8213 | 1220648-2025-29527 | 22289275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289275&pc=OZD | 05/29/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665441 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella 5.5 support and during support, the patient developed hemolysis. the patient had hemolysis in the effluent bag and their plasma free hemoglobin was trending up. the source of the hemolysis was not confirmed. position was verified on echocardiogram. support continued. care was withdrawn and the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8214 | 1220648-2025-29528 | 22289277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289277&pc=OZD | 06/09/2025 | 08/06/2025 |  | 06/19/2025 |  |  | 06/09/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella cp support due to acute myocardial infarction. during support, it was reported that the patient sat up to vomit and caused some groin bleed. manual pressure was held, in addition to the patient receiving two units of blood. a hematoma was noticed at the site of the impella cp but had not changed since earlier that day. hemolysis was also noted earlier. however, the urine was clearing up. the patient was successfully weaned from impella cp without complications and was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient movement. the patient had sat up to vomit over the side of bed and caused some groin bleeding.
8215 | 1220648-2025-29529 | 22289292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289292&pc=OZD | 05/31/2025 | 07/30/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656913 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, a high purge pressure / purge pressure blocked alarms began to appear. the purge cassette was replaced to resolve the failure. upon planned explant of the pump, no signs of a thrombus or biomaterial were noted in the device. that night, the patient suffered a large, right sided embolic stroke with complete left sided deficits and a craniotomy was scheduled to treat the condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of stroke/cva and high pressure purge has been completed. high purge pressure (hpp) : data log review showed purge pressure high alarms on (b)(6) 2025 around 1430-1709. bag change occurs 22 hours after last hpp alarm. no motor current (mc) spikes noted. pulsatile flow and ps noted. the purge cassette was not returned. the pump was returned and inspected. the pump was returned cut so no purge pressure investigation could be completed. the purge gap, purge line, and impellor blades were free from biomaterial. the cause of the high purge pressure was not determined since returned product was cut and could not be tested to reproduced the hpp. stroke/cva: the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b1 product problem was inadvertently omitted for high pressure purge on the initial manufacturer device report number 1220648-2025-29529. h6 medical device problem code 1248 was erroneously entered on the initial manufacturer device report number 1220648-2025-29529. h10 instructions for use were inadvertently omitted for high pressure purge on the initial manufacturer device report number 1220648-2025-29529.
8216 | 1220648-2025-29531 | 22289303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289303&pc=OZD | 06/02/2025 | 07/21/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1400984 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that fluid might have entered the automed impella controller during use. there was no harm to the patient. however, the console has been isolated. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: cp pump (b)(6) was connected at 6:58 on 6/2 with purge cassette sn (b)(6). at 14:00 the user starts a purge cassette change and sn (b)(6) is swapped for sn (b)(6). the purge cassette swap is completed but then 20 minutes later the purge cassette was removed and re-inserted twice in quick succession before being removed three minutes later along with the pump ((b)(6) imclog.pdf). the rtlogs show the purge cassette change but there was no notable performance drop off ((b)(6) rtlog.png). device analysis: the console was returned for investigation and the motor was found to be unable to rotate. additionally, there were signs of glucose around the purge motor shaft ((b)(6) glucose purge shaft.png). after cleaning the shaft, it was able to rotate again. root cause: the spilled fluid that the customer said entered the console most likely caused the inability to rotate. repair: please clean (b)(6)'s purge unit, then confirm motor is able to rotate. then perform sections 10-12 of pm (0042-7309) and functional check (0042-7316).
8217 | 1220648-2025-29536 | 22289347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289347&pc=OZD | 06/02/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660170 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Delivered as Unsterile Product | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was being escalated to a impella 5.5 after an impella cp was placed. it was noted that the impella 5.5 was placed without complications. however, shortly after placement, right ventricle function worsened and the decision made to add veno-arterial extracorporeal membrane oxygenation. the impella cp was removed. it was also noted that there was a question about sterile packaging of the impella 5.5 introducer and the physician requested another one. in addition it was reported that during impella 5.5 support, the patient had runs of ventricular tachycardia (vt) requiring cardioversion. after cardioversion, the physician came to the bedside and repositioned the impella. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
8218 | 1220648-2025-29537 | 22289378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289378&pc=OZD | 05/22/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568818 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Abscess | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on a impella 5.5 support. it was reported that after explant an abscess formation was believed to have formed at the impella explant site scar. no noted intervention was taken for the reported issue. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8219 | 1220648-2025-29538 | 22289379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289379&pc=OZD | 05/27/2025 | 07/17/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654592 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was being prepped for an implant. when the pump was connected to the automated impella controller (aic) a controller error alarm occurred. the staff exchanged the aic for a new one and attempted to prime the impella, but the alarm persisted and the impella pump could not be primed successfully for insertion. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The fm optical signal issue was added based on the investigation as the controller errors prevented the user from being able to prime the pump. optical signal issue. the cause of the optical signal issue was a kinked optical fiber line resulting from tubing lines within the patient cable sliding out and into the red handle due to an improper manufacturing sequence of tubing line assembly into the patient cable. no component related defect. priming issue. there is no problem found with the priming issue as the purge system was unaffected and it was the controller errors from the optical signal issue preventing the pump to be primed.
8220 | 1220648-2025-29543 | 22289410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289410&pc=OZD | 05/29/2025 | 08/25/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1504720 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) shutoff unexpectedly during patient support. an "unexpected controller shutdown" message appeared and the nurse subsequently swapped the aic. started on second console without any issues. patient remained stable throughout console switch. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. the automated impella controller (aic) was returned for investigation. the console logs for the event revealing multiple partial restarts starting at 00:32:04. the first restart occurred approximately 13 days after the case was initiated. following the second restart, an "unexpected controller shutdown alarm" was triggered. shortly after, the pump was automatically restarted at previous p-level. moments after, both "controller error (loss of comm (pbm1))" and "controller error (loss of comm (pbm2))" alarms were triggered. the console experienced two more restarts, with similar alarms observed before ultimately shutting down and being replaced with a new console. additionally, multiple communication issues with various subsystems within the aic were noted just before the first restart, and these issues persisted until the console was shut down. the console ic8273 was sent back to field service for further investigation and while testing, the console failed to complete the boot-up process, and console was than forcefully shutdown. upon further investigation, it was determined that the 4-in-1 pca was defective and causing communication issues within the console. after replacing the faulty 4-in-1 pca with a known good one, the console was able to boot completely. further the console was tested with multiple power cycles as well as with a known good pump and purge cassette and no issues were found. additionally, no interruptions in the power supply were detected between the pbm and the replaced 4-in-1 pca. the root cause of the reported issue, console shut off without any warning, was due to the 4-in-1 pca malfunction.
8221 | 1220648-2025-29544 | 22289415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289415&pc=OZD | 06/01/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/01/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597503 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following the implant, same day, the patient's leg became marbled/mottled during the mcs. a fem-fem bypass was recommended; however, the decision was made to not intervene. the next day the patient expired. according to head physician, there was no isolated ischemia in the leg where the impella was located in the femoral artery; there was a generalized capillary hypoperfusion in all terminal vascular areas during maximally dosed vasopressor therapy. | [5 empty] | There is not enough evidence to exclude the impella cp device used as an associated factor given the limb ischemia. however, it should be noted according to the physician, the patient's cause of death was prolonged cardiogenic shock (scai shock stage e) following a myocardial infarction with pre-existing multivessel coronary artery disease. neither the hypoperfusion nor the fatal outcome can be causally attributed to the impella, but rather to the severe cardiogenic shock. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8222 | 1220648-2025-29547 | 22289440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289440&pc=OZD | 06/01/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/01/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675700 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and following the implant the pump stopped multiple times. the decision was made to explant the device (after a total of 40 hours of support). post removal of the impella 5.5 device, the patient was found to have extensive thrombus burden involving both the peripheral vascular and intracardiac structures. the status of the patient following the event is unknown (unnoted). | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8223 | 1220648-2025-29551 | 22289466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289466&pc=OZD | 05/27/2025 | 08/25/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669069 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported by the physician that the impella implantation was not successful because of kinking above the sheath. implantation with sheath 14f x25cm was not successful. | The complainant reported a patient who was hemodynamically unstable, with high-dose catecholamines, was being implanted with an impella cp device for mechanical circulatory support. the implant was aborted due to the patient's anatomy, and the patient subsequently expired. |  |  |  |  | B5 updated with additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29551. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-29551. h6 medical device problem code 2889 and component code 4742 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29551. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8224 | 1220648-2025-29553 | 22289489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289489&pc=OZD | 06/02/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671964 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Convulsion/Seizure | Ectopic Heartbeat | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella 5.5 for support during a high risk cardiac procedure. the chest was opened and coronary artery bypass graft was performed and an impella 5.5 was placed successfully. during impella 5.5 support, it was reported that the upon arrival to the bedside, the patient was progressing very well with discussions of possibly extubating. however, the patient experienced a seizure lasting approximately 25 seconds. hemodynamics remain unchanged after the seizure. the patient was taken foe a computed tomography to assess and bleeding was noted. additionally ,there was decoupling of the aortic and left ventricular waveforms and ectopy. no suction alarms had occurred and flow was on the lower side of normal for p-7. the patient was intubated and sedated with high dose propofol. the patient was not on systemic heparin due to brain bleed. the patient taken to the operating room and the physician exposed the impella graft and removed sutures. the impella was turned off at this time per the physician's request. the impella required some manipulation to come out of the graft but was ultimately explanted without any issue. of note, while in the operating room, the patient converted to atrial fibrillation with rapid ventricular response and blood pressure (bp) decreased. anesthesia increased epinephrine drip due to drop in bp and gave a push dose neo-synephrine. upon return to intensive care unit, rhythm converted back to normal sinus rhythm with marked improvement in bp, and the patient was being weaned from the epinephrine infusion. also of note, the patient began following commands yesterday. post-operatively, the patient was noted to nod appropriately as well as follow commands. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8225 | 1220648-2025-29555 | 22289503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289503&pc=OZD | 06/02/2025 | 07/31/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female | 39 KG | Non Hispanic | Asian | Malfunction | No | Yes | Malposition of Device | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a shallow position. attempts to reposition the pump were unsuccessful due to the patient's anatomy. the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of positioning issues has been completed. product was not returned for review. the root cause of positioning issue was most likely patient condition related since the patient has small aortic arch and small left ventricle (lv) resulted in difficult advancement the pump to lv.
8226 | 1220648-2025-29560 | 22289660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289660&pc=OZD | 06/02/2025 | 07/31/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653606 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 114 KG | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump could not go higher than p-2 without encountering suction issues. the decision was made to place an impella rp flex pump for additional support, but the patient subsequently passed away following bipella implant. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. low pump flow: data log reviewed: no motor current (mc) spike outside of p-level changes and initiation of impella support observed. suction alarms observed. position unknown alarms seen. placement signal (ps) trending down. the cause of low pump flow most likely was patient condition related due to right ventricle failure. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8227 | 1220648-2025-29562 | 22289684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289684&pc=OZD | 06/02/2025 | 07/30/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583115 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male | 118 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult or Delayed Positioning | Malposition of Device | [5 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was difficult due to the small size of the patient's vessels so a direct aortic approach was used. the guidewire became displaced and the patient experienced ventricular fibrillation requiring resuscitation and defibrillation. later, the impella became positioned toward the mitral complex. hemolysis and possible ectopy was also noted. the impella was explanted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of delivery issues, vt/vf, positioning issues, and hemolysis has been completed. the impella device and data logs were received for investigation. delivery issues: the clinical details reported that "the attempted with 0.18 wire, wire became displaced, device able to be advanced into lv cavity." the guidewire was not returned. due to a lack of clinical details on how the wire became displaced, the root cause of the delivery issues cannot be determined. vt/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. positioning issues: the clinical stated that on the 5th, the device was repositioned. on the 6th, it was stated that overnight the pump was positioned towards the mitral complex. this was the third time the pump was malpositioned and ended up causing hemolysis. the data logs show frequent impella position in aorta and impella position unknown alarms. based on the clinical details, the root cause of the positioning issues is most likely due to use as the pump was malpositioned multiple times during the case. hemolysis: the clinical stated that on the 5th, the device was repositioned. on the 6th, it was stated that overnight the pump was positioned towards the mitral complex. this was the third time the pump was malpositioned and ended up causing hemolysis. the data logs show frequent impella position in aorta and impella position unknown alarms. there were no spikes in motor current or any trends in the placement signal. the returned product had biomaterial on the impeller, but the cause of that is unknown as there was no ingestion signatures in the logs. there were no damages noted with the product. based on the clinical description and data logs, the root cause of the hemolysis is most likely due to improper positioning of the device. thrombus: thrombus was added based on the returned product investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
8228 | 1220648-2025-29563 | 22289686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289686&pc=OZD | 05/25/2025 | 07/15/2025 |  | 06/19/2025 |  |  | 05/25/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026668988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Hemolysis | Sepsis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp support was septic with methicillin-resistant staphylococcus aureus. additionally, labs were coming back hemolyzed. the physician team was not concerned due to the fact that the kidneys were already unrecoverable, and no more repositioning efforts would be made. the p-level was decreased due to hemolysis. it was further reported that the impella cp had a spike in motor current and pump stop occurred.the device was successfully restarted in p2 and ramped up to p7. saturated flows were noticed. the pump was running at p7. there were no anticoagulants infusing related to the patients disease process. the family elected to withdraw care, and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of temporary pump stop, infection/sepsis, hemolysis, and saturated flow has been completed. the impella device was not returned for evaluation. temporary pump stop: data logs were reviewed and a motor current spike with speed deviation occurred just before an alarm 13 "impella stopped - motor current high" alarm was triggered. pump was able to be restarted, purge pressure increased and purge flows decreased after restart. pump flows also became fully saturated and were able to regain some pulsatility as pump continued to run. clinical details noted that nurse observed a motor current spike prior to pump stop occurring. pump was able to be restarted at p-2 and turned up to p-7. clinical details also noted that the patient was not on any anticoagulation infusion at the time of the pump stop. the root cause of the temporary pump stop could not be definitively determined as no product was returned for evaluation to confirm pattern observed in the logs. hemolysis: data logs were reviewed and lt log at time of reported hemolysis showed no positioning issues, suction alarms, or motor current spikes. clinical details noted that patients labs were hemolyzing. patient had chronic kidney disease (ckd) and clinical team was aware that kidneys were unrecoverable. clinical team opted to not reposition pump at that time. the root cause of the hemolysis was most likely patient condition related per clinical details. saturated flows: data logs were reviewed and lt log showed that saturated flows occurred after a motor current spike and pump stop. prior to motor current spike, pump flows are within normal range, after motor current spike and pump stop, pump flows are fully saturated but regain some pulsatility as pump continues to run. clinical details noted that after pump was restarted after pump stop, flows became saturated. the root cause of the saturated flows could not be definitively determined as no product was returned to confirm pattern observed in the data logs. infection/sepsis: without additional clinical details the root cause of the infection cannot be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: entering cardiac output. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
8229 | 1220648-2025-29567 | 22289734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289734&pc=OZD | 05/24/2025 | 07/15/2025 |  | 06/19/2025 |  |  | 05/24/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639621 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient's heart was noted to be dilated on transesophageal echocardiogram so the decision was made for an impella cp. during impella cp support, it was noted that the patient remained critically ill on extracorporeal membrane oxygenation and impella cp support. bival was paused for gastrointestinal bleed concern. however, heparin was restarted. plasma-free hemoglobin/lactate dehydrogenase results were discussed. the doctor was not concerned that it was impella cp related rather total disease state driven. it was also noted that the patient had hemolysis and was on dialysis. position was verified on echocardiogram. the patient was ultimately escalated to an impella 5.5. impella cp removal was successful after placement of the impella 5.5. additional information was reported that no gastrointestinal bleed was noted when doing rounds. coloscopy showed no active ore recently active bleeding sites. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Clinical summary: hemolysis occurred on the 5th day of support. impella cp in perfect position on echo, no waveform abnormalities, alarms, or flow concerns. md does not attribute pfhgb or ldh increase to impella cp. more so probable of ecmo circuit and entire patient disease state. data review: data logs showed pump ran without issue during support. there were positioning alarms triggered during the case but resolved quickly. no suction alarms. no mc spikes outside p level changes. product was not returned for review. the root cause of the hemolysis was most likely patient condition related based on clinical description that the pump was in proper position, the patient was critically ill and on ecmo, md does not attribute issue to impella, and data log review shows no other root cause indications.
8230 | 1220648-2025-29571 | 22289773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289773&pc=OZD | 05/25/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/25/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621972 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support had a history of atrial fibrillation with rapid ventricular response. while on support, the patient went back into atrial fibrillation with rapid ventricular response. the next day, the team was hoping to get the heart rate under control as the patient has been in and out of atrial fibrillation with rapid ventricular response and becomes symptomatic when she loses her atrial kick. the patient was successfully weaned and explanted the following day. | [5 empty] | The impella device was not returned, and the investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the atrial fibrillation was not determined.
8231 | 1220648-2025-29573 | 22289791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289791&pc=OZD | 05/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026671736 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient on impella 5.5 support had occasional ventricular tachycardia (vt). the impella was repositioned in an attempt to help with the vt and the patient remained on support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8232 | 1220648-2025-29575 | 22289792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289792&pc=OZD | 05/25/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/25/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657970 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that following implant, a hematoma developed proximal and distal to the access site and spreading into the scrotum. heparin was halted in response to the condition and manual pressure was held. support continued following the intervention. the pump was successfully weaned and explanted two days later. | [5 empty] | The impella device was discarded by the customer and the investigation has been completed. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8233 | 1220648-2025-29574 | 22289793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289793&pc=OZD | 06/03/2025 | 07/15/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1246899 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had false purge pressure low alarms during prep. it was stated that this issue occurred previously and it was thought to be the purge cassette at the time. the aic was replaced. there was no patient involvement. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on t the data and device analysis the false low purge pressure alarm was not caused by any issues with the aic, as the aic was adjusting to the pump being connected to the purge cassette. no problem found within the aic.
8234 | 1220648-2025-29576 | 22289803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289803&pc=OZD | 05/30/2025 | 07/26/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella cp support developed a slight hematoma at the access site. femostop was placed with light belt pressure. additionally, the impella position appeared to be deep, and it was repositioned. the next day, the patient¿s urine cleared up. | [5 empty] | Access site bleeding - minor: slight hematoma at access site observed, likely due to patient coughing and high blood pressure and femostop placed with light belt pressure. the cause of access site bleeding - minor issue most likely was patient movement related due to patient coughing and high blood pressure. hemolysis: swan was in the wedge position. red urine noticed. echo confirmed impella position appears deep. reposition performed. next day urine cleared up after repositioning. data log reviewed: no mc spike observed. no positioning or placement signal issues occurred. no indication observed. the cause of hemolysis issue most like due improper technique since the red urine cleared after repositioning performed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8235 | 1220648-2025-29580 | 22289849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289849&pc=OZD | 06/04/2025 | 08/01/2025 |  | 06/19/2025 |  |  | 06/04/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658679 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male | 118 KG | Non Hispanic | White | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a red plug leak was observed on an implanted impella 5.5. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of product damage (hole in catherter) has been completed. product damage (hole in catheter): clinical details note that there was coming from the red impella plug shortly after arriving in icu. data logs are not applicable. product was returned and blood was found in the red handle. there was a puncture hole found in the catheter near the 30 cm marking. the cause of the product damage (hole in catheter) was most likely use-related since there was a puncture hole found near the 30 cm catheter marking that led to blood ingress into the red handle. d9: device returned to manufacturer date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8236 | 1220648-2025-29583 | 22289879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289879&pc=OZD | 06/04/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/04/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650224 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 58 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient had no distal pulse in the impella leg following explant of an impella 5.5. thrombi were subsequently found in the contralateral femoral, right radial, and left radial arteries. a fem-fem was placed prior to impella support and was used upon explantation. the patient was successfully weaned from the pump and survived. | [5 empty] | A5 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
8237 | 1220648-2025-29585 | 22289891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22289891&pc=OZD | 05/26/2025 | 07/15/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480323 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient had all labs drawn and were hemolyzed. the patient¿s urine was dark red. the impella was repositioned, and the p-level was reduced. additionally, limb ischemia also noted with no palpitable or dopplerable pedal pulses. the patient was weaned off support and was explanted. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The pump was returned for investigation. no evidence of physical damage was found on the returned product. no biomaterial was observed on the impeller or inflow cage. the pump passed all electrical testing and was able to run in a bucket of water without any abnormalities. the data log indicates a gradual decrease in the placement signal, with several impella position unknown alarms during the time of reported hemolysis. this trend was resolved in 5 hours. there was no conclusive evidence of positioning issues or motor current spike recorded on the data log. additionally, the clinical notes confirm the resolution of hemolysis later during the case run. the cause of the hemolysis issue was most likely patient condition related, based on data log review. the cause of the positioning issue was most likely use related since pump was pulled out in aorta during patient¿s rolling post-bowel movement when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
8238 | 1220648-2025-29512 | 22291684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291684&pc=OZD | 06/03/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665153 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5, the patient experienced arrhythmia which was treated with vasopressors and cardioversion. additionally, bleeding was noted from a tracheotomy. homeostasis was achieved by stitching the bleed. the impella 5.5 continues to support the patient. | [5 empty] | The impella device was not received from the customer as it continues to support the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8239 | 1220648-2025-29509 | 22291768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291768&pc=OZD | 06/03/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659447 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and experienced access site bleeding requiring intervention and device malpositioning leading to explant. the patient presented to the emergency room with shortness of breath and dizziness. the patient was taken to the catheterization lab with st elevation in the anterior leads. angiogram revealed total occlusion of the proximal left anterior descending artery. percutaneous coronary intervention was performed and the patient decompensated. the decision was made to implant an impella cp which was successfully completed, and angle matching was completed in the catheterization lab before transporting the patient to the unit. the groin was dry and intact at this time. approximately 20 minutes after being in the unit, there was oozing. partial thromboplastin time (ptt) was greater than 200 seconds and mean arterial pressure (map) was greater than 90. the physician decided to hold heparin until ptt was therapeutic and started a nitroglycerin drip to get the blood pressure and map down. the grind was redressed with a new angle match in place and surgi-cell was applied to the site as well. the access site was checked the following morning, and the site was again dry and intact and no hematoma. additionally occurring during this impella implant experience, the patient had hematuria. the pump was moved back and was interfering with the anterior mitral valve leaflet. the impella cp device was advanced. however, it appeared to be stuck in the papillary muscle and unable to reposition out of it. subsequently, the patient was taken to the catheterization lab for impella cp explant which was completed without incident. post check it was noted, urine output cleared after adjustments. there was no more hematuria and the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8240 | 1220648-2025-29514 | 22291799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291799&pc=PYX | 06/03/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  | 12/31/2025 |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025448793 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported elective placement, insertion of impella rp flex device for mechanical circulatory support (mcs) post coronary artery bypass graft surgery was unsuccessful. a 7 french balloon wedge catheter floated pre-sternotomy into the right pulmonary artery. coronary artery bypass graft times two, mitral valve replacement, tricuspid valve replacement and left atrial appendage were completed. the impella rp flex wire was inserted through balloon wedge catheter, and a right ventricle bleed was noted. the area of the bleed was repaired, and the impella rp flex implant was aborted. subsequent mcs treatment to the patient, if any, was unnoted. the patient was reported to be in stable condition following the event. | [5 empty] | The impella guidewire and pump were discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
8241 | 1220648-2025-29516 | 22291826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291826&pc=OZD | 06/02/2025 | 07/30/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026654950 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Improper or Incorrect Procedure or Method | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 insertion, no wire was used and the heart was digitally manipulated causing a left ventricle perforation. the left ventricle was patched and the impella 5.5 was placed and turned on. the patient ended up expiring and the device stayed in the patient. | [5 empty] | The impella device was not received from the customer as it was kept inside the deceased patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of ventricle perforation has been completed. no product was returned for evaluation. clinical details noted that during direct access pump insertion without a wire, the lv was perforated as heart was being manipulated by surgeon. the root cause of the ventricle perforation was most likely a use related issue as the heart shifted during insertion procedure. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29516.
8242 | 1220648-2025-29513 | 22291910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291910&pc=OZD | 06/03/2025 | 07/31/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665275 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | Yes | Failure to Advance | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp for mechanical circulatory support but was unsuccessful. upon insertion of the .018 guidewire into the left ventricle, the patient went into ventricular tachycardia. the pump was then inserted into the femoral artery. however, the pump was unable to advance beyond the iliac. as cardiopulmonary resuscitation was being performed, the decision was made to remove the impella cp and place the patient on extracorporeal membrane oxygenation (ecmo) support. upon removal of the impella cp, biomaterial was found between the wire and the catheter pigtail, which prevented further advancement of the impella into the patient. the patient survived. however, the patient was noted to be in critical condition following the event. | [5 empty] | The impella pump, purge cassette and guidewire were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues, vt/vf, and thrombus has been completed. delivery issue - reported feeling resistance in the iliac and was no longer able to advance the impella cp. pump delivery was aborted and patient was put on extracorporeal membrane oxygenation (ecmo). when the pump was removed, there was tissue between the wire and pigtail inhibiting the pump to advance on the guidewire. impella cp, guidewire and purge cassette were returned. guidewire was in the pump cannula when returned. there was dried blood around the guidewire and pigtail but no biomaterial observed. the root cause of the delivery issue was most likely patient condition related since it was reported there was found some tissue between the wire and the pigtail which inhibited the ability for the impact to be inserted any further in to the patient. vt/vf - patient went into vt after the guidewire was inserted. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. thrombus - there was tissue found between the guidewire and pigtail after the pump was explanted. the root cause of the thrombus was unable to be determined due to insufficient clinical details.
8243 | 1220648-2025-29517 | 22291925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291925&pc=OZD | 05/26/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660463 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp and experienced ventricular fibrillation which required defibrillation. two days later, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8244 | 1220648-2025-29519 | 22291946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22291946&pc=OZD | 06/03/2025 | 07/18/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1082391 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was supporting a patient admitted in with cardiomyopathy when after over 2 weeks of support, the purge fluid was changed. the automated impella controller (aic) froze on the screen where you edit the purge fluid information. the team was unable to move forward or exit out. the aic was changed to the back up aic. the patients outcome was stable. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other and aic - kbd/rotary knob issues has been completed. data review: console logs were consistent with was reported in the complaint. device analysis: console was tested but the kbd issue was unable to be reproduced. the consoles kbd was responsive and had no issues advancing through the purge procedure. conclusion: the root cause of this issue was unable to be conclusively determined due to the inability to reproduce. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8245 | 1220648-2025-29522 | 22292080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292080&pc=OZD | 06/03/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617438 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support and was having runs of ventricular tachycardia. the impella was repositioned from 5.7 centimeters to 4.7 centimeters. the patient was successfully weaned. the patient survived the explant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8246 | 1220648-2025-29526 | 22292092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292092&pc=PYX | 05/31/2025 | 07/08/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517798 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, the flows dropped from 3.0 to 1.4. the doctor stated there was a pump stop. the pump was restarted and then the pump was explanted and biomaterial was observed. | [5 empty] | Insufficient product was returned for analysis. the pump was severed at the catheter and only cannula was returned. biomaterial was observed by the impeller. no data logs were returned. reproduction testing could not be performed. low pump flow: the root cause of the low pump low was unable to be determined as no logs were returned for analysis and insufficient product was returned for evaluation. temporary pump stop: the root cause of the temporary pump stop was unable to be determined as no logs were returned for analysis and insufficient product was returned for evaluation. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
8247 | 1220648-2025-29532 | 22292122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292122&pc=PYX | 05/27/2025 | 07/10/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  | 08/31/2025 |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024396637 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 81 KG | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella rp device for mechanical circulatory support. the pump was replaced due to delivery issues and lost optical signal. the impella rp was replaced, and support proceeded. | [5 empty] | Product changed to 0.035gw based on investigation revealing issue with wire in clinical details. failure mode 'difficulty inserting/removing introducer' added. no product or data logs were returned for analysis. delivery issues: the root cause of the delivery issue was most likely patient condition related as evidenced by clinical detail of issues passing wires with resistance in the femoral vein. difficulty inserting/removing introducer: the root cause of the difficulty inserting/removing introducer was most likely patient condition related as evidenced by clinical detail of issues of being unable to pass through vasculature with resistance in the femoral vein. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8248 | 1220648-2025-29534 | 22292172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292172&pc=OZD | 05/31/2025 | 07/18/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653469 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the pump had intermittent suction alarms. packed red blood cells were ordered. additionally, the pump was positioned towards the mitral valve (mv) and there was an inability to reposition, and the pump was explanted and replaced with a new impella 5.5. | [5 empty] | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. low pump flow: clinical stated 5.5 device exchanged due to malposition and inability to reposition at bedside under echo. data log reviewed: no motor current (mc) spike outside of p-level changes observed. no positioning or placement signal issue seen. many suction alarms were observed at the first day of support followed by a drop in flow at p6 . also suctions observed on the last 2hrs before impella exchange. the cause of low pump flow issue most likely was improper positioning due to improper technique, device exchanged due to malposition and inability to reposition at bedside under echocardiogram. positioning issues: clinical stated: impella 5.5 positioned toward the mv and unable to reposition. left 5.5 in over night, could not get appropriate flows without suctioning. unknown why impella malposition. the cause of positioning issues cannot be determined due to insufficient clinical details. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8249 | 1220648-2025-29525 | 22292196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292196&pc=OZD | 05/29/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1817137 |  | 042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that when the vad coordinator went to transfer automated impella controller (aic) to aic there was a purge disk error alarm. a new aic was used. there was no patient harm. | [5 empty] | The investigation for the aic (console) purge disc/flag issues has been completed. data logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) several times. the console was returned for evaluation finding a broken purge disc flag was identified. therefore, the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
8250 | 1220648-2025-29523 | 22292241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292241&pc=OZD | 05/30/2025 | 08/28/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095483 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | Device Damaged Prior to Use | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was shipped to the facility, and it arrived in a wet and damaged box. the controller had a crack on the rear housing near the base. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. data analysis: log analysis was uninformative for this complaint. there was nothing noteworthy in the logs and the console was last booted up two weeks before the complaint date. device analysis: the console was examined for any sign of water damage but none were found. the console was run overnight and performed as expected. there were two cracks in the rear housing as described in the complaint but they did not inhibit console performance. root cause: the cause of the cracks in the rear housing was not determined.
8251 | 1220648-2025-29530 | 22292289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292289&pc=OZD | 05/31/2025 | 07/16/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664133 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was explanted due to flows dropping to zero. a new impella cp was placed, and the patient remained hemodynamically stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. the returned complaint cp pump visually inspected. no biomaterial observed. no cannula kink or broken blade found. impella cp connected to console and no low flow can be reproduced and no other issue observed. data logs reviewed. brief motor current (mc) spike with all parameters resolving on (b)(6). on (b)(6) 2025 suction alarms occurred with sudden drop in flows and motor current at p-7. no suction alarms prior to this event. flows are just above 0. 3 minutes later the p-level is turned down to p-6, then p-2. flows remain just above 0 at p2 even when the pump is turned to p-0 twice with mc spikes as it turns back to p-2. flows remain close to 0 with no resolve from then turning p-level to p-4. ps and purge pressure remain stable. the real time (rt) logs not available at time of issue. low pump flow: the cause of low pump flow issue cannot be determined due to lack of clinical details, no abnormalities on returned product, no rt logs at time of issue, and low flow not reproduced on returned pump.
8252 | 1220648-2025-29535 | 22292300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292300&pc=OZD | 05/27/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665340 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 94 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that upon placement of the impella cp repositioning sheath, the staff was unable to find distal pulses in the lower right extremity. a fem-fem bypass was subsequently placed to restore flow to the limb. support continued as planned with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller. section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
8253 | 1220648-2025-29533 | 22292426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292426&pc=PYX | 05/31/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025547368 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with known cardiac history and end-stage renal disease (esrd) was implanted with an impella rp flex device for mechanical circulatory support. while on support, there was a diagnosis made of heparin induced thrombocytopenia (hit) that prompted heparin removal from systemic iv, the purge was not running with heparin but rather sodium bicarb. the pump supported for 4 days and then the patient expired from care. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
8254 | 1220648-2025-29542 | 22292435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292435&pc=OZD | 05/27/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1708923 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was being prepped for an implant. when the pump was connected to the automated impella controller (aic) a controller error alarm occurred. the staff exchanged the aic for a new one and attempted to prime the impella, but the alarm persisted and the impella pump could not be primed successfully for insertion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8255 | 1220648-2025-29556 | 22292565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292565&pc=OZD | 03/05/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 03/05/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620883 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male | 99 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella cp exhibited access site bleeding. two 8fr angioseals were deployed to obtain hemostasis. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8256 | 1220648-2025-29545 | 22292573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292573&pc=OZD | 05/27/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815958 |  | 0042-0040-US-L |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an impella 5.5 was being prepped for an implant. when the pump was connected to the automated impella controller (aic) a controller error alarm occurred. the staff exchanged the aic for a new one and attempted to prime the impella, but the alarm persisted and the impella pump could not be primed successfully for insertion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8257 | 1220648-2025-29539 | 22292578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292578&pc=OZD | 06/02/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584509 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient presented with complaints of back pain. pyelonephritis, sepsis, and elevated troponin were noted. the patient was taken for percutaneous coronary intervention of the left anterior descending artery. however, the patient became unstable. the decision for an impella cp was made and was placed successfully. the impella cp was started in auto then decreased to p4 due to suction alarms. the patient taken to intensive care unit on support. it was noted that the impella cp was weaned and explanted due to positive blood cultures and continued sepsis. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist, section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8258 | 1220648-2025-29550 | 22292579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292579&pc=OZD | 06/02/2025 | 07/18/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581657 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted when an angiogram showed the patient had aortic coarctation and graft dissection. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella couldn't pass through the vasculature since angio was taken and aortic coarcatation and graft dissection was discovered. | A4 and a5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8259 | 1220648-2025-29540 | 22292581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292581&pc=OZD | 05/31/2025 | 07/21/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644232 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was inadvertently pulled back into the aorta. attempts to recross the valve were unsuccessful and the decision was made to remove the pump. a new impella cp pump was placed for ongoing support. | [5 empty] | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was pulled back into aorta inadvertently. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance." understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8260 | 1220648-2025-29548 | 22292594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292594&pc=OZD | 06/02/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666909 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella 5.5 was inserted and the impella was initiated at p2. the impella was increased to p4 and held due to left atrial thrombus and high after load from vasopressors. the physician decided to leave the impella 5.5 at p4 as to not dislodge the left atrial thrombus. the patient continues on impella 5.5 support. | [5 empty] | The impella device was not received from the customer as the pump is still supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8261 | 1220648-2025-29549 | 22292606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292606&pc=PYX | 06/02/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026629377 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male | 120 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella rp flex developed thrombocytopenia. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8262 | 1220648-2025-29541 | 22292609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292609&pc=OZD | 05/29/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631971 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump supported patient who was escalated to the 5.5 after just 19 hours of support and undergoing coronary artery bypass surgery. the 5.5 pump was placed but the patient did suffer from pulseless electrical activity and arrested on the 5.5 pump. the patient had chest compressions and veno-arterial extracorporeal membrane oxygenation emergently placed. the patient remains on for support. | [5 empty] | The investigation has been completed. no product was returned. to determine the true root cause of the cardiac arrest, the clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the was not determined.
8263 | 1220648-2025-29552 | 22292610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292610&pc=OZD | 05/29/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643553 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported impella cp support ongoing when decision made to remove and escalate to the 5.5 pump. when explanting the cp the access site was closed by perclose x3 but, post there were no pulses and so vascular surgery performed repair to remedy the limb ischemia. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation has been completed. no product was returned. to determine the true root cause of limb ischemia, the clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8264 | 1220648-2025-29554 | 22292692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292692&pc=OZD | 05/27/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/27/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026660254 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp developed gastrointestinal bleeding. the patient was transfused with packed red blood cells and platelets, and anticoagulation was withheld to obtain hemostasis. the patient was successfully weaned from the impella and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
8265 | 1220648-2025-29546 | 22292736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292736&pc=OZD | 05/29/2025 | 07/16/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653650 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 was not able to cross into the aorta. after multiple and unsuccessful attempts, a kink was noted to the impella 5.5 cannula and catheter. another impella was used via direct approach. there was no patient harm reported. | [5 empty] | Delivery issue: the root cause of delivery issue is determined to be patient condition because the pump was unable to pass cross into aorta despite placement of 0.035 wire. product damage (cannula kink): after multiple attempts during the insertion, a kink on the catheter and cannula were noted. the root cause of product damage is not determined because of insufficient clinical details the product was not returned for investigation. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
8266 | 1220648-2025-29558 | 22292970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292970&pc=OZD | 05/23/2025 | 07/07/2025 |  | 06/19/2025 |  |  | 05/23/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660185 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male | 100 KG | Non Hispanic | White | Death | Yes | Yes | Failure to Prime | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The medical safety officer reviewed and determined the following: there is not enough evidence to exclude the impella cp pump 1 used as an associated factor in the patient's outcome. there was an interruption in the procedure as a result of the device malfunction (no purge noted exiting the pump). however, it must be noted care was withdrawal because the patient never quite recovered. | The complainant reported that an impella cp pump failed to properly prime during preparation for implant. the pump was replaced to resolve the issue. no patient harm was reported. |  |  |  |  | The investigation of priming issue has been completed. based on purge pressure and flows being within specification at the beginning of the case in field data logs and returned products being able to prime without issue, the root cause of the priming issue was determined to most likely be no problem found. the medical safety officer reviewed and determined the following: there is not enough evidence to exclude the impella cp pump 1 used as an associated factor in the patient's outcome. there was an interruption in the procedure as a result of the device malfunction (no purge noted exiting the pump). however, it must be noted care was withdrawal because the patient never quite recovered. d9 device returned to manufacturer date added b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29558. b2 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29558. b5 narrative of patient outcome and medical safety officer review was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29558. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29558. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-29558. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-29558. new code entered. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29558. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8267 | 1220648-2025-29557 | 22292976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22292976&pc=PYX | 05/29/2025 | 07/15/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026661912 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella rp flex pump could not be delivered successfully for patient support. the implant was aborted due to the patient's anatomy. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: rp flex with smart assist system with automated impella controller ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient anatomy.
8268 | 1220648-2025-29561 | 22293087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293087&pc=PYX | 05/28/2025 | 07/16/2025 |  | 06/19/2025 |  |  | 05/28/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600735 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male | 91 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, low flows occurred when the patient was getting red blood cells (rbc)s infused. the rbcs were given as well as continuous renal replacement therapy (crrt) throughout the support. after just under seven days the pump was explanted due to the flows. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Insufficient product returned for analysis. the returned pump had the cannula severed off at the catheter. thrombus failure mode added as observed on returned product. placement signal issue: biomaterial was observed in the inflow cage around the impeller. the data logs from the case show placement signal drifting up while flows drifted down with flows at 2.5l/min at p7. there were no motor current spikes outside p-level changes before decreased flows. the cvp was stable and ps and flows continued to drift. the device was explanted due to this issue. the returned pump could not be tested as it was returned severed at the catheter. the root cause of the placement signal issue was most likely sensor drift as evidenced by the data log pattern observed. low pump flow: based on the investigation, no problem was found as the root cause of the low pump flow. the perceived low pump flow was as a result of the sensor drift of the dp sensor which calculated flow is based on. thrombus: thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
8269 | 1220648-2025-29564 | 22293139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293139&pc=OZD | 06/02/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026666036 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump placed along with ecmo to support the patient with known cardiac and systemic comorbidities. the pump was supporting when the limb with both ecmo and cp cannula became ischemic. the medical therapy dosing may have contributed to the ischemia, and the pump was not explanted. after just under 5 days the pump support was stopped/withdrawn due to underlying condition and anoxic brain injury. the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8270 | 1220648-2025-29559 | 22293172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293172&pc=OZD | 05/28/2025 | 08/26/2025 |  | 06/19/2025 |  |  | 05/28/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651373 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implantation of an impella cp, the companion sheath met some slight resistance when passing through the peel away sheath. the implant was successful. upon removal of the companion sheath, some slight damage was noted at the tip of the component. the damage had no impact on the procedure. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of introducer damage ¿ mechanical has been completed. the 7fr low-profile companion sheath was returned for investigation. sheath tip damage was noted on the returned introducer. according to the clinical notes, slight friction was encountered during the placement of the 7fr sheath into the abiomed peel-away sheath, and upon explant, tip damage was discovered on the returned product. the cause of the introducer damage issue was determined to be sheath tip damage; however, the cause of the sheath tip damage could not be concluded without sufficient clinical information during the procedure. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29559.
8271 | 1220648-2025-29566 | 22293224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293224&pc=OZD | 12/23/2024 | 07/09/2025 |  | 06/19/2025 |  |  | 11/27/2024 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546146 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Epistaxis | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received regarding a patient that experienced a massive epistaxis, thrombocytopenia, and a ventricular tachycardia arrest all with a possible relationship to the impella device used. the patient presented in cardiogenic shock. echocardiogram reflected ejection fraction of 10% with moderate to severe mitral regurgitation and right ventricular dysfunction. during an attempt to place a swan the patient went into sustained ventricular tachycardia and the decision was made to place the impella 5.5 device. the patient remained on support for approximately four days. on day three it was noted the patient still very little pulsatility noted with left ventricular assist device (lvad) planned for the following day. the impella 5.5 device was explanted the following day with a survived outcome (it is unknown if the lvad was placed). as reported per the registry, the patient underwent esophagogastroduodenoscopy (esd) without obvious source of bleeding from esophagus. the egd was negative. however, the patient was found to have massive epistaxis, likely oropharyngeal bleeding. rhino rockets and packing were completed, and the patient was given afrin and epinephrine. the packing removed approximately six days later, and patient had some recurrent bleeding. repacking with rhinnorocket was completed. the patient required a total of 10 units of packed red blood cells products. the patient recovered with no sequelae. several days later it was noted the patient's platelets decreased post-procedure. likely consumptive. subcutaneous heparin was stopped. a total of 7 units of platelets were transfused between over a five-day period. it was noted the patient did not recovered/event not resolved. and approximately five days later after the thrombocytopenia event, the patient's amiodarone had decreased, and the patient went into ventricular tachycardia arrest. the patient given epi and shocked with return of spontaneous circulation achieved. the patient was noted following commands post-code; the patient recovered with no sequelae. | [5 empty] | Additional information was received and noted that regarding the thrombocytopenia, it started (b)(6) 2024 post explant of device (provided as date captured to section d6b). consequently, update was made to reflect the implant date as well. as the actual event date still remains unknown, the date of the thrombocytopenia occurred is captured as the provisional event date. note: investigation and its conclusion as previously reported remains unchanged. | Event date and implant/explant dates are conflicting and are being clarified. consequently, these respective dates fields were wittingly left blank. should additional/clarification information be received a supplemental mdr will be submitted accordingly. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. however, an investigation was completed and noted the following: vascular damage peripheral vessel and thrombocytopenia root cause(s) could be determined since the product was not returned, and insufficient clinical details were provided. however, additionally, the cardiac arrythmia (ventricular tachycardia/ventricular fibrillation) can be reported during impella support. however, to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease, timing of the arrythmia event in relation to impella support (during or post-implant), electrocardiogram (ekg) recordings of the arrhythmia episodes, evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances, assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements, presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities, response to any antiarrhythmic treatments or interventions during the event, any documented device-related issues or complications during impella support, evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand, and review of any concomitant medications that may influence cardiac electrophysiology. furthermore, with a returned impella device used, it can be inspected for any burrs or rough edges that may contacted the heart wall. therefore, to determine the true root cause of the ventricular tachycardia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia was not determined.
8272 | 1220648-2025-29568 | 22293313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293313&pc=OZD | 06/02/2025 | 09/06/2025 |  | 06/19/2025 |  |  | 06/02/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026670473 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that the patient progressed to further liver and kidney failure and started on continuous renal replacement therapy (crrt). no other information is available. | The user facility reported that the patient on impella cp support developed hematuria. the cp was repositioned under fluoroscopy. additionally, a large laminar thrombus along lateral lv myocardium was observed. the team is discussing options. it was further reported that there was a loss of aortic placement signal and lv waveform after extubating and the patient coughing. patient survived the procedure. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. neither product nor data logs were available for investigation. since neither product nor data logs were returned for investigation, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 medical device problem code: 2917 device sensing problem were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29568. | The investigation of hemolysis, positioning issues, thrombus, and renal failure has been completed. the impella device was not returned for evaluation. hemolysis: based on the clinical details provided, the pump's position could have been the cause of hemolysis, however, it was not explicitly reported. additionally, the report of low filling volume on 3/jun/2025 even though no blood products were administered is a potential contributing factor to hemolysis. since insufficient clinical details were provided and neither product nor data logs were returned for investigation, the root cause of the hemolysis could not be determined. positioning issues: neither product nor data logs were returned for investigation. since insufficient clinical details were provided and product wasn't returned for investigation, the root cause of the positioning issue could not be determined. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. renal failure: the cause of the renal failure was not determined due to insufficient clinical details. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29568. h6 health effect - clinical code 4499 and health effect - impact code 4641 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29568. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8273 | 1220648-2025-29565 | 22293348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293348&pc=OZD | 05/24/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/24/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630615 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was eventually implanted with an impella 5.5 device for mechanical circulatory support (mcs) after an unsuccessful attempt with the first impella 5.5 device. however, following the implant of the impella 5.5 (pump 2), the patient developed an access site hematoma. a jackson-pratt drain was placed, heparin held, and platelets and fresh frozen plasma were administered. hematoma evacuation was a consideration if needed. support was continued with plans for percutaneous coronary intervention (pci) within the next two days. two days later, the patient experienced non-sustained tachycardia. however, the following day. the patient would undergo pci to the left main coronary artery and experienced a vessel dissection treated with a covered stent. echocardiogram confirmed no effusion. the patient continued on impella mcs, and over the next couple of days, the patient progressed to the chair and walking. on day eight of support, the device was successfully weaned and explanted with a survived outcome. | [5 empty] | This is one of two devices associated with the patient's implant experience. refer to manufacturing report number 1220648-2025-29405 for the report on the first impella 5.5 used/attempted for implant. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8274 | 1220648-2025-29569 | 22293353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293353&pc=OZD | 05/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502104 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that distal pulses were not found following impella cp implant. the staff opted to continue with support while monitoring the condition. the patient passed away the following day during support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8275 | 1220648-2025-29570 | 22293381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293381&pc=OZD | 06/03/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 06/03/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1414254 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during prepping the impella aic had optical sensor failure. the aic was exchanged. there was no patient harm. | [5 empty] | The investigation for the optical signal issue with the aic (console) has been completed. console logs from the reported day of event were reviewed and are consistent with complaint and show that after connecting pump during case start, a ¿placement signal not reliable¿ event occurred, accompanied by event invalid_bad_snr_sample_mask (invalid sample due to snr below minimum). calibration of the optical sensor failed, and case start wizard presented ¿optical sensor system error¿ prompt. pump was disconnected, then reconnected, but ¿optical sensor system detection progress error¿ occurred. after consecutive attempt failed, the operator was prompted to replace the console. the real time (rt) log data corresponding to the case start indicated that optical snr was below 50, and lamp and gain were maxed out at 100% and 2.5 accordingly, indicating low quality and low power of optical signal. the console was returned for evaluation and a piece of debris (most likely fragment of white pump plug) was found lodged in the blue receptacle, which would result in air gap and very low snr and low power of optical signal. optical ¿5s¿ parameters were measured (using two test cavities bypassing the blue receptacle) but results differed, depending on test cavity length. monitoring of analog output of the ccd detector in optical bench revealed deformation of the baseline (presumably defect of ccd) which significantly lowered snr for shorter cavity lengths (higher values of pressure). ultimately, the inability to start support with this console was due to perceived optical sensor failure was mostly caused by debris in pump connector preventing pump plug to be inserted fully, which resulted in an air gap lowering the snr.
8276 | 1220648-2025-29577 | 22293418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293418&pc=OZD | 04/30/2025 | 07/09/2025 |  | 06/19/2025 |  |  | 06/04/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630621 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 84 KG | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump began to experience sensor issues during patient support. the pump was explanted the following day and a thrombus was discovered on the outlet near the motor. no additional action was required. the patient was successfully weaned from the pump and survived. | [5 empty] | The investigation of optical signal issue and thrombus has been completed. optical signal issue: the returned product was inspected and biomaterial was observed by the sensor and outlet cage. the 5s parameters on the returned product were within specification and pump passed the light continuity test. the data logs show stable 5s parameters but confirm placement signal not reliable alarm. there were no motor current spikes outside p-level changes. the root cause of the optical signal issue was most likely biomaterial as evidenced by biomaterial observed on the sensor on returned product and stable 5s parameters thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8277 | 1220648-2025-29578 | 22293503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293503&pc=OZD | 06/04/2025 | 07/31/2025 |  | 06/19/2025 |  |  | 06/04/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026674574 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation | Distributor |  |  |  | 64 YR |  | Female | 78 KG | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp developed red colored urine. the impella was repositioned to resolve the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of hemolysis and positioning issues has been completed. hemolysis: data logs confirmed pump was in improper position corresponded with some impella position in aorta events on 6/3/25. also there were some suction alarms. no other issues were found on the logs. product was not returned for review. based on clinical description and data review, the root cause of hemolysis was most likely due to pump was not in the proper position since the clinical team confirmed poor position of impella caused hemolysis and resolved after repositioning of the pump. positioning issues: the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
8278 | 1220648-2025-29579 | 22293518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293518&pc=OZD | 05/25/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/25/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2026644138 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella cp was placed for a cardiogenic shock patient in addition to the extracorporeal membrane oxygenation circuit. the pump was placed but the limb became ischemic. the pump was explanted and escalated to an impella 5.5. the limb had become cold and pulse was faint. additional information was received. the limb ischemia was due to pressors and not the pump. once the patient arrived to the intenstive care unit and was balanced out, color and pulse came back. the patient was taken for an impella 5.5 upgrade three hours later. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8279 | 1220648-2025-29572 | 22293538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293538&pc=OZD | 05/30/2025 | 10/21/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665413 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella cp implant and during the implant, the impella 5.5 was unable to transverse into the ascending aorta due to anatomy. the pump was explanted and attempted again. the impella was backloaded, placed into artery, but again unable to get through anatomy. the axillary artery was dissected and the implant was aborted. treatment or patient outcome was unknown. | [5 empty] | Delivery issue: the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. vascular damage: the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
8280 | 1220648-2025-29584 | 22293612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293612&pc=OZD | 05/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568428 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp / rp / 5.5 device for mechanical circulatory support. while on support, there were new episodes of arrhythmia. there was both ventricular tachycardia (vt) and atrial fibrillation. the pump remains on for support. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8281 | 1220648-2025-29586 | 22293655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293655&pc=OZD | 05/30/2025 | 07/28/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666895 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Detachment of Device or Device Component | [5 empty] | Tachycardia | Great Vessel Perforation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, the patient had frequent ectopy and tachycardia. support continued. additionally, there were suction alarms. two (2) units of blood products were given to the patient. also, the patient had a pericardial hematoma. the patient was intubated and cannulated for peripheral va-ecmo. at the time of cannulation, impella flows acutely dropped following a suction alarm with dampening motor current, jump in ao placement signal (that no longer correlated with arterial pressure) and extremely low impella flows. patient taken to the operating room for chest washout. chest re-opened, copious amount of old clot removed from the pericardium. no left ventricle thrombus was noted. impella removed via graft with usual amount of tension, but separated at the junction between the outlet cage and cannula. graft clamped at level of aorta, distal portion of the cannula located within the graft and removed. cannula and outlet/impeller inspected, no clot visualized. | [5 empty] | Vt/vf in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. thrombus in order to make a root cause determination, all of the clinical details outlined in es2023-0158 would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessels the patient had bleeding throughout support but the source of the bleeding was unknown. due to limited clinical information, the root cause of the vascular damage - peripheral vessels cannot be determined. low pump flow throughout support volume was administered to the patient, and it was stated that the patient had very little pulsatility but it would improve with volume. on (b)(6), the patient had an rv collapse and was cannulated for va-ecmo, and impella flows decreased. the data logs show suction alarms throughout support, and on the 5th, there is a lack of pulsatility seen in the mc and trending ps, and flows significantly decrease. the returned product did not show any thrombus or kinks in the cannula. the pump was unable to be run as the cannula was detached. based on the clinical details provided and data log analysis, the root cause of the low pump flow is most likely due to the patients condition. product damage - detached inflow/outflow there are not enough details surrounding the removal/explant of the device to make a determination about the angle of explant, or what occurred during this time. this issue has previously been seen on direct cases during removal and can be referenced in the previous escalation pia-0000148. the design testing has stated that the bond is tested to a 30 n specification per dir550-1.1.7.2. additionally, the bond was ripped at the cannula to ph bond and was jagged remaining on the cannula and not the motor housing. on the inside of the cannula there was also residual glue marks where it would be attached to the motor housing. there were no kinks seen in the cannula or the catheter. due to a lack of sufficient evidence during the removal of the 5.5 and evidence of the glue bond, the root cause of the product damage (detached inflow/outflow) cannot be determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8282 | 1220648-2025-29581 | 22293712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293712&pc=OZD | 05/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657937 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and after the implant, the patient had two runs of ventricular tachycardia (vt) and was shocked. support continued. additionally, there were suction alarms, and the patient was given albumin. support continued. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8283 | 1220648-2025-29582 | 22293772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293772&pc=OZD | 05/26/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026663674 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed some bleeding at the groin site. femostop was placed, however, shortly after, perfusion was no longer present in the impella leg. the femostop was noted to have been inflated too much, which led to the poor perfusion of the extremity. the femostop was adjusted and the leg recovered. support continued with the implanted pump. the patients outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8284 | 1220648-2025-29588 | 22293786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293786&pc=OZD | 06/04/2025 | 07/21/2025 |  | 06/19/2025 |  |  | 06/04/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658318 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to a kink in shaft of the catheter. no patient harm was reported. | [5 empty] | The investigation of delivery issues and product damage (cannula kink) has been completed. impella 5.5 was returned and catheter kinks were found at 27cm, 31cm and 36cm. no cannula kinks noted. delivery issue: the root cause of the delivery issue was not determined due to insufficient clinical details. catheter kink: the cause of the catheter kink was not determined due to lack of clinical details. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29588. | A4 is unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
8285 | 1220648-2025-29592 | 22293852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293852&pc=OZD | 05/26/2025 | 07/23/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644241 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, upon insertion of the repositioning sheath during impella cp implant, oozing was noted at the site. the existing perclose slack was tightened and forward tension sutures were applied to the site. once in the intensive care unit, the distal pulses were absent, and no flow was seen in the vessels. the patient was transferred to a high level of care in response to the condition. it was believed that the perclose may have tightened the lateral wall of arteriotomy. that evening, the patient began to bleed from their mouth. anticoagulation was stopped and two units of blood were given to treat the condition. by the following day, a hematoma had developed at the access site and dopplerable pulses were present. support continued with the implanted pump. the patients outcome is critical. | [5 empty] | B1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29592. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29592. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Access site bleed - major: clinical details report that after the introducer sheath was removed and the repositioning unit was advanced, the access site was oozing. forward tension sutures had not been applied at this time. the percloses were tightened and forward tension sutures were applied. the complication resolved and the patient went back to the icu with the site c/d/i. product was not returned for investigation. based on the clinical details provided that forward tension sutures were not in place at the time of the bleed and the issue resolved once they were, the root cause of the access site bleed was determined to most likely be a use issue. vascular damage - peripheral: clinical details reported that once the patient was back in the icu on the first day of support, the patient was bleeding from the mouth. anticoagulation was stopped and blood products were administered. it is unclear from clinical details if the issue resolved. product was not returned for investigation. since product wasn't returned and limited clinical details were provided, the root cause of the vascular damage could not be determined. limb ischemia - reversible: clinical details report that once the patient arrived in the icu, distal pulses were absent. the decision was made to transfer the patient to hloc. when the patient arrived at the new hospital on (b)(6) 2025, the impella extremity was warm, and dopplerable pulses were noted. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. positioning issues: when the pump was initially delivered into the lv, the physician noted that it was under a papillary muscle. no attempts to reposition resolved the complication, so the pump was explanted and reinserted. once this was done, the pump was able to be positioned satisfactorily. no clinical details were provided regarding the patient's anatomy. data logs were reviewed and there were 4 position in aorta alarms roughly 15 minutes into the case, likely during the time that the physician was attempting to reposition the pump. there were no position in ventricle alarms to confirm that the pump was inserted too deep. product was not returned for investigation. since insufficient clinical details were provided and product wasn't returned for investigation, the root cause of the positioning issue could not be determined. | Access site bleed - major: clinical details report that after the introducer sheath was removed and the repositioning unit was advanced, the access site was oozing. forward tension sutures had not been applied at this time. the percloses were tightened and forward tension sutures were applied. the complication resolved and the patient went back to the icu with the site c/d/i. product was not returned for investigation. based on the clinical details provided that forward tension sutures were not in place at the time of the bleed and the issue resolved once they were, the root cause of the access site bleed was determined to most likely be a use issue. vascular damage - peripheral: clinical details reported that once the patient was back in the icu on the first day of support, the patient was bleeding from the mouth. anticoagulation was stopped and blood products were administered. it is unclear from clinical details if the issue resolved. product was not returned for investigation. since product wasn't returned and limited clinical details were provided, the root cause of the vascular damage could not be determined. limb ischemia - reversible: clinical details report that once the patient arrived in the icu, distal pulses were absent. the decision was made to transfer the patient to hloc. when the patient arrived at the (b)(6) on (b)(6) 2025, the impella extremity was warm, and dopplerable pulses were noted. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. positioning issues: when the pump was initially delivered into the lv, the physician noted that it was under a papillary muscle. no attempts to reposition resolved the complication, so the pump was explanted and reinserted. once this was done, the pump was able to be positioned satisfactorily. no clinical details were provided regarding the patient's anatomy. data logs were reviewed and there were 4 position in aorta alarms roughly 15 minutes into the case, likely during the time that the physician was attempting to reposition the pump. there were no position in ventricle alarms to confirm that the pump was inserted too deep. product was not returned for investigation. since insufficient clinical details were provided and product wasn't returned for investigation, the root cause of the positioning issue could not be determined.
8286 | 1220648-2025-29587 | 22293856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293856&pc=OZD | 05/26/2025 | 07/21/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661088 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient went into supraventricular tachycardia (svt). on evho noted impella was mispositioned. post reposition svt was managed and then dark red urine noted later that day that stayed persistent. hemolysis was resolved with impella removal. | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the physician had to stick the right common femoral artery twice. micro-puncture was used both times. first stick was too low and dangerously close to the bifurcation. the second stick seemed appropriate. bleeding occurred. manual compression, dressing changes, and additional hemostatic stitch were done. the groin site is stable. additionally, the impella cp was delivered into the left ventricle without issue, the position optimized and started in auto. the patient went into supraventricular tachycardia and the impella was noted to be malpositioned. the position was optimized again with good results. |  |  |  |  | The investigation of access site bleeding, vf/vt/af/at, hemolysis, and positioning issues has been completed. hemolysis: impella cp was returned and biomaterial was found around the impeller. no other damages or abnormalities were found. data logs show a motor current spike on (b)(6) 2025 at around 6am followed by suction and a drop in flows. the root cause of the hemolysis was most likely biomaterial since there was biomaterial found around the impeller and an ingestion pattern in the data logs. access site bleeding: the root cause of the access site bleeding was most likely use related since the femoral artery was stuck multiple times and the bleeding was resolved with an additional hemostatic stitch. positioning issues: the root cause of the positioning issue was not determined due to insufficient clinical details. vt/vf: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. instructions for use: impella cp with smartassist system section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29587. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29587. h6 health effect - clinical code 1886, health effect - impact code 46247, and medical device problem code 3009 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29587. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-29587. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8287 | 1220648-2025-29591 | 22293887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293887&pc=OZD | 05/30/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533053 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Hemolysis | Tachycardia | Gastrointestinal Hemorrhage | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, the patient had atrial fibrillation and ventricular tachycardia. support continued. additionally, the patient had hemolysis. the pump was found at 7cm, so the doctor repositioned the pump and the patient was given one unit of packed red blood cells. the patient had a gastrointestinal bleed. heparin was withheld. furthermore, care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8288 | 1220648-2025-29590 | 22293889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293889&pc=OZD | 05/26/2025 | 07/21/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564752 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the pump became dislodged in cardiac catheterization lab on initial implant. impella cp was explanted successfully and the site rewired and a new peel away was inserted. a new impella cp was inserted and the patient's outcome is stable. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. data logs show case length of 17 minutes and impella position in aorta alarms with suction. the pump and introducer were returned and inspected. the introducer was partially torn along the peel away score lines. no cannula kink was observed. the cause of the positioning issues was most likely use related due to the dislodgment occurring during transition from peel away to repo sheath.
8289 | 1220648-2025-29589 | 22293897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22293897&pc=OZD | 05/30/2025 | 07/17/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026675685 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella 5.5 delivery through the axillary artery was unsuccessful. a direct approach was then completed successfully for patient support. the patient survived the procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The root cause of delivery issue was most likely patient condition since pump was unable to pass through artery.
8290 | 1220648-2025-29593 | 22294080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22294080&pc=OZD | 03/11/2024 | 07/18/2025 |  | 06/19/2025 |  |  | 05/30/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The impact study reported that impella 5.5 procedure was attempted on (b)(6) 2024, but the 5.5 was not successfully placed. the patient had aortic calcification and aortic root needed to be expanded and replaced. the impella procedure was aborted. it was further reported that on (b)(6) 2024, the patient had an ischemic cerebrovascular accident. the patient recovered on (b)(6) 2024. cec relatedness is ¿possible¿ to impella device and impella procedure. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1- a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | Delivery issue: the root cause of delivery issue is determined to patient condition because the implantation was cancelled due to vessel stenosis. stroke/cva: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿
8291 | 1220648-2025-29594 | 22294089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22294089&pc=OZD | 05/26/2025 | 06/19/2025 |  | 06/19/2025 |  |  | 05/26/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1281596 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Display or Visual Feedback Problem | Decrease in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an air in purge alarm appeared on the automated impella controller (aic) during patient support. while attempting to de-air the system, the screen on the aic froze. a purge cassette exchange was performed, but the issue persisted. the aic was subsequently swapped to resolve the failures. support continued following the exchange. however, after 11 days of support, the patient eventually expired. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8292 | 1220648-2025-29596 | 22294172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22294172&pc=OZD | 05/29/2025 | 07/23/2025 |  | 06/19/2025 |  |  | 05/29/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656660 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. during anastomosis of graft, arterial tissue was injured requiring surgical repair. once the 23 french sheath was placed, pigtails over table wire were advanced without any complications. .035 wire was exchanged for .018 wire. the impella was advanced over the .018 wire while advancing the impella into the innominate artery the impella was hard to advanced. the impella was pulled back and an angiogram was performed showing dissection, and no distal flow. femoral access was obtained and a covered stent was placed, surgical was stapled to the graft and the case was aborted. | [5 empty] | No product was returned for analysis vascular damage - great vessel: clinical mention that graft arterial tissue was injured. no other clinical information related to how failure occurred was provided. the root cause of the vascular damage - great vessel was not determined as limited clinical information related to failure was provided. delivery issues: clinical mention while advancing impella into the innominate artery the impella was hard to advance. poor vasculature was mentioned in patient history. the root cause of the delivery issue was most likely patient condition related as evidenced by the difficulty advancing pump due to poor patient vasculature vascular damage - peripheral vessel: clinical mention dissection observed when pump was advanced. no other clinical information was provided as to how dissection occurred. the root cause of the vascular damage - peripheral vessel was not determined as insufficient clinical information was provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | H6 medical device problem code 2682 was inadvertently omitted on the initial manufacturer device report number.
8293 | 1220648-2025-29595 | 22294244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22294244&pc=OZD | 05/31/2025 | 08/26/2025 |  | 06/19/2025 |  |  | 05/31/2025 | 06/19/2025 |  |  |  |  | 06/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570742 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp and ecmo support ongoing for the patient admitted in with acute myocardial infarction. the pump was supporting when all access sites were bleeding. the hemoglobin lab values dropped and so the patient was transfused with many units of blood and products. the pump remains on despite the bleeding. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related since bleeding on every accesses and massive coagulation disorder observed. | The impella device has not been received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8294 | 1220648-2025-29413 | 22282174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282174&pc=OZD | 05/22/2025 | 07/15/2025 |  | 06/18/2025 |  |  | 05/22/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584669 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was taken to cardiovascular operating room for a mitral valve replacement (mvr) and impella 5.5 axillary placement. successful mvr was noted, however, due to tortuous anatomy there were unable to get thw impella 5.5 positioned well. the decision was made to explant the impella 5.5 to avoid interference with the new mitral valve. additional information was received by the complaints team via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient had endured recurring thrombocytopenia with a "possible" relationship to the impella device used: ¿the patient developed recurring thrombocytopenia. the patient has had several procedures and experiences multifactorial acute blood loss anemia. admission platelets 166 k/ul, nadir 23 k/ul ((b)(6) 2025). the patient received 12 units of platelets.¿ the patient's outcome is unknown. the complaints team also received via abiomed¿s loqi impella registry that the patient endured recurring procedural mediastinal hemorrhage with a "definite" relationship to the impella device used: ¿the patient underwent a procedure on (b)(6) 2025 for mitral valve repair, decannulation of va ecmo to vv ecmo, cardiopulmonary bypass ((b)(6) 2025 1643 - 2106) impella cp removal, and impella 5.5 implant. the (b)(6) procedure began at 1559 and ended on (b)(6) 2025 at 0037. the patient received 3 units of packed red blood and 3 units of fresh frozen plasma during the procedure. the surgeon mentions a pericardial effusion and bilateral pleural effusions discovered during the procedure. admission hemoglobin 12.5 mg/dl; hgb prior to first procedure 8.4 mg/dl, hgb post-first-procedure 9.9 mg/dl." the patient had mediastinal and bilateral pleural chest tubes after the procedure.¿ the patient's outcome is unknown. | [5 empty] | The investigation of positioning issues, thrombocytopenia, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. positioning: the cause of the positioning issue was most likely patient condition related due to patient's tortuous anatomy. vascular damage peripheral: the cause of the vascular damage was not determined due to lack of clinical details. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8295 | 1220648-2025-29414 | 22282215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282215&pc=OZD | 05/23/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/23/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594823 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had been bleeding from the axillary site. one unit of red blood cells were given to the patient and heparin was reduced. support continued for seven more days and the patient's outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8296 | 1220648-2025-29415 | 22282436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282436&pc=OZD | 05/19/2025 | 07/14/2025 |  | 06/18/2025 |  |  | 05/19/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655573 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | IMPELLA RP FLEX | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a cardiac history and had come in for the ross (aortic root replacement) procedure and when condition deteriorated an emergent coronary artery bypass graft. the patient was unable to come off bypass after reported six attempts. extracorporeal membrane oxygenation was considered prior to placing intra -aortic balloon pump (iabp). the patient failed iabp/inotropic support and the impella team was notified of need. during impella support, the patient had extensive surgical/chest bleeding and protamine was administered due to surgical oozing. additionally, it was reported that the impella 5.5 was unable to flow optimally due to low volume. the team attempted to correct bleeding surgically before chest was closed. the patient continued to bleed post closure. replacement clotting factors were administered at 1.5 liters cell saver along with other blood products. the team decided that surgical bleed could not be corrected and the decision was made to transfer the patient to the unit. the patient remained on support to allow the family time to decide on whether to withdraw care or escalate care through the addition of an impella rp flex. the patient was escalated during care. as the patient continued support, hemodynamics were declining despite multiple pressors and bicarbonate pushes. ultimately, the patient coded and was not able to achieve return of spontaneous circulation. | [5 empty] | The investigation of low pump flow and access site bleeding has been completed. the pump was not returned for investigation. the data log shows several suction alarms throughout the case run, with a trending placement signal. there was no motor current spike or positioning-related alarm noted. the cause of the low pump flow issue was most likely patient condition related, based on data log review and given clinical details. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | Medical safety review of the event noted there is not enough evidence to exclude the impella 5.5 device used as an associated factor to the patient's outcome given related events of bleeding and low pump flow. this is one of two related medical device reports associated with this patients care. the associated report will reflect the impella rp flex. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8297 | 1220648-2025-29416 | 22282497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282497&pc=OZD | 05/24/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102523 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the console failed to recognize two purge cassettes during change. the automated impella controller was swapped to resolve the issue. there was no patient harm reported. | [5 empty] | The console was received from the customer and an investigation regarding the returned device has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined and a broken purge disc flag was identified. purge unit driver connector lock was broken. the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc.
8298 | 1220648-2025-29417 | 22282601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282601&pc=OZD | 01/04/2020 | 06/18/2025 |  | 06/18/2025 |  |  | 01/04/2020 | 06/18/2025 |  |  | 04/30/2021 |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1415469 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient endured hemolysis during impella cp support. following implant, the patient began to develop hematuria. the support level was reduced to p-2 and the device was repositioned to address the condition. the following day, the decision was made to explant the pump as the hemolysis had not yet resolved. the patient passed away the next day. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8299 | 1220648-2025-29418 | 22282626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282626&pc=OZD | 05/24/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666898 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, micro clots were discovered in both ventricles, and heparin was started with the patient on bipella support. it was further reported that the patient had embolic stemi post 5.5 removal. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller.
8300 | 1220648-2025-29419 | 22282695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282695&pc=OZD | 05/24/2025 | 08/04/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627399 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient spiked a fever. blood cultures were drawn, and it was positive for gram rods. the team suspects dialysis catheter as the culprit, antibiotics regimen started. it was further reported that high purge pressure and purge system blocked alarm occurred. tissue plasminogen activator (tpa) was administered and the purge pressures normalized and the pump remains implanted for support. | [5 empty] | The investigation of high purge pressure and infection/sepsis has been completed. high purge pressure: throughout the case, purge pressure and purge flows were steady. there was one increase on 5/14 triggering high purge pressure but there was no spike in motor current. the rise time was 3 minutes and it remained high for 3 minutes and then rapidly declined. this was not during a bag change. the root cause of the first high purge pressure event is most likely due to a kinked purge line but location of kink is unknown. the purge pressure remains steady until 5/24 when a small increase is seen in the purge pressure to about 800 and then it goes back down, no significant changes in motor current (mc). about 2 hours later, the purge pressure increases over the span of an hour. during this time the motor current increases and flows become slightly saturated. it remains high for about 17 hours until it starts to gradually decrease. the clinical stated tissue plasminogen activator (tpa) was started and it resolved. based on the data logs and clinical details, the root cause of the high purge pressure event was most likely due to biomaterial buildup. infection/sepsis : the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. thrombus: the failure mode thrombus was added based on the investigation findings. the root cause of the thrombus was unable to be determined due to insufficient clinical details. d6b explant date added. | The impella device remains implanted and was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8301 | 1220648-2025-29420 | 22282696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282696&pc=OZD | 05/24/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026669900 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the patient was hypotensive and was developing a hematoma. the hematoma continued to grow. seven units of red blood cells, fresh frozen plasma, six units of albumin and additional chemical support were given to the patient. the pump was explanted and replaced. once removed, the surgeon noted that there were multiple tears in the vessel. it looked like upon access with the needle, the needles went through and through. posterior femoral tear from access needle was noted. the patients outcome is stable. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. upon removal, multiple vessel tears that looked to be from the access needle were noted. impella cp was returned cut at the catheter. no other damages or abnormalities were found. the root cause of the vascular damage - peripheral vessel was most likely use related since the vessel tears looked to be from the access needle. h6 health effect - clinical code was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29420. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
8302 | 1220648-2025-29421 | 22282749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282749&pc=OZD | 05/24/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660174 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after an impella cp was inserted, the patient¿s distal pulses diminished and eventually became non-dopplerable. the impella cp was explanted and flow afterwards down the left leg was excellent. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8303 | 1220648-2025-29422 | 22282751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282751&pc=PYX | 05/24/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025507443 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, micro clots were discovered in both ventricles and heparin was started to mitigate the condition. the rp flex was removed bedside with no complications. the patient remained on impella 5.5 support which is slowly being weaned down. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8304 | 1220648-2025-29423 | 22282811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282811&pc=OZD | 05/25/2025 | 08/21/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026659901 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was an access site bleed treated by suture, medical management, and blood product delivery. the pump remains on and was successfully weaned and explanted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. impella cp was returned and no damages or abnormalities were found. leak test was performed and the repositioning unit met specifications. the root cause of the access site bleeding was not determined due to insufficient clinical information.
8305 | 1220648-2025-29424 | 22282887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282887&pc=OZD | 05/23/2025 | 10/30/2025 |  | 06/18/2025 |  |  | 05/23/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607599 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male | 122 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support. while on support there was continuous suction on automated impella controller. motor current was fast, impella flow ranges were saturated within 0.1 of each other. thought impella was initially tangled in papillary, after further repositioning under echocardiogram and fluoroscopy, the impella was mid ventricular. it was also reported that the impella was also being removed due to a clot and a new device would be inserted. after the impella cp was explanted significant repair needed to repair right femoral artery and 2 units of packed red blood cells given for blood loss. an impella 5.5 was inserted and the patient was noted in stable condition. | [5 empty] | The investigation of positioning issues, saturated flow, thrombus, vascular damage - peripheral vessel, low pump flow, and access site bleeding has been completed. the pump was returned for investigation. data logs were returned for analysis. positioning issues: data log review shows no positioning alarms. the root cause of the positioning issues was not determined since insufficient clinical details were provided. saturated flow: data log review showed a large motor current spike on the night of (b)(6) 2025. earlier on (b)(6) 2025, there was disturbances in purge pressure and a shift up in motor current (mc) and pump flow. product evaluation revealed biomaterial around the impeller blades. the root cause of the saturated flow was most likely biomaterial ingestion based on clinical details, data log analysis, and product evaluation. thrombus: data log review shows a large motor current spike. the root cause of the thrombus was unable to be determined due to insufficient clinical details. access site bleeding - major: the root cause of the access site bleeding was not determined due to insufficient clinical details and inconclusive evidence from product evaluation. vascular damage - peripheral vessel: the root cause of the vascular damage - peripheral vessel was not determined due to insufficient clinical details and inconclusive evidence from product evaluation. low pump flow: data log review shows that the suction alarms began when the pump was turned down from p-7. there had been a mc spike during the suction alarms. the pump was turned down from p-7 to p-4 after the mc spike and there was low flow at p-4. product evaluation revealed biomaterial around the impeller blades. the root cause of the low pump flow was most likely biomaterial since a large motor current spike was seen in the data logs and a clot was observed in the patient and on the outflow of the pump upon explant. d9 device returned to manufacturer date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8306 | 1220648-2025-29425 | 22282892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282892&pc=OZD | 06/06/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/06/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information was obtained from a literature report. it was noted that a patient underwent percutaneous coronary intervention under impella support for st-elevation acute myocardial infarction. on the sixth day, the impella device was removed, and the impella cp, which had been inserted via the right femoral artery, was surgically extracted in the intensive care unit. the patient had a history of stent placement due to bilateral external iliac artery stenosis, and a computed tomography (ct) scan upon arrival revealed calcification and severe stenosis in the right common iliac artery. resistance was noted during the removal of the impella, and immediately after its extraction, there was no detectable pulse in the femoral artery, raising suspicion of acute arterial occlusion due to arterial dissection. a contrast-enhanced ct scan revealed occlusion from a thrombus in the common iliac artery, leading to the diagnosis of acute arterial occlusion due to thrombotic occlusion. surgical thrombectomy and stenting of the right common iliac artery was performed in the hybrid operating room. postoperatively, the dorsalis pedis artery was palpated well, and the patient continued without lower limb ischemia. | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8307 | 1220648-2025-29426 | 22282961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22282961&pc=OZD | 05/27/2025 | 08/22/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1115051 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller had a broken purge disc bracket. this issue was discovered after removal from patient use and there was no patient harm. | [5 empty] | A.1, a.2, a.3, a.4 and a.5 are unknown. e.1 name prefix/title, first name and last name are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for aic - purge disc/flag issues has been completed. console logs did not contain any data relevant to this complaint. the console was returned for evaluation. the reported broken purge disc retainer was confirmed. purge retainer was completely broken off. the root cause of this issue was a broken purge disc retainer.
8308 | 1220648-2025-29428 | 22283076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283076&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026651707 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had blood loss related to development of gi bleed. reduction in total systemic volume leading to suspected decreased flow past distal extremities. one unit of packed red blood cells (prbc) was given. doppler pulses were intermittent bilaterally but only right extremity painful. computed tomography (ct) run off shows significant chronic peripheral arterial disease (pad). the impella continued to be weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ and cardiac or vascular injury (including ventricular perforation.¿
8309 | 1220648-2025-29427 | 22283114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283114&pc=OZD | 06/06/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 06/06/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 54 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported from information obtained from a literature review that a patient had extensive myocardial ischemia accompanied by st elevation. on the fifth day of hospitalization, the intra-aortic balloon pump (iabp) was changed to an impella cp in order to improve pulmonary hypertension by reducing the left ventricular load. on the eighth day of hospitalization, the impella cp was removed due to worsening aortic regurgitation, and peripheral veno-arterial extracorporeal membrane oxygenation (va ecmo) was changed to central va ecmo to ensure antegrade arterial blood flow. however, despite intensive care, left ventricular wall motion did not recover, and mechanical circulatory support with central va ecmo was continued. on the 23rd day of hospitalization, the patient developed a hemorrhagic cerebral infarction in the right cerebral hemisphere, and died on the 26th day of hospitalization. the relationship between the impella and cause of death is unknown | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter, impella catheter inlet to the aorta: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
8310 | 1220648-2025-29429 | 22283167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283167&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632576 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had ventricular tachycardia (vt) and torsades. the patient received amnio boluses, mag and lidocaine. the patient was shocked once and support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8311 | 1220648-2025-29430 | 22283245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283245&pc=OZD | 05/27/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608904 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the implant, there were suction alarms. an echocardiogram (echo) preformed and position near mitral valve was noted. unable to reposition and resolve suction. 1.5 liters of fluid was administered. the echo demonstrated impella deep and obstructed under the papillary muscle. the team attempted to reposition several times with echo guidance. anatomically, the patient noted to have sigmoid septum obstructing cannula position into the left ventricle. the decision was made to explant the pump. the patient remained hemodynamically stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms, ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position, ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
8312 | 1220648-2025-29431 | 22283311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283311&pc=OZD | 05/25/2025 | 07/10/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026657506 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was aborted due to inability to pass through the patient's vasculature. a new device was prepped for implantation. no patient harm was reported. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to resistance at the femoral artery.
8313 | 1220648-2025-29432 | 22283580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283580&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562286 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male | 90 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while implanted with an impella 5.5 the patient experienced ventricular fibrillation. the patient was defibrillated over ten times and amiodarone and lidocaine were administered. due to lack of brain function the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8314 | 1220648-2025-29433 | 22283680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283680&pc=OZD | 05/24/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/24/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615044 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the access site was sore with mild tenderness and support continued. additionally, there was serosanguinous fluid in the anticontamination sleeve and the sleeve had a rip in it. the doctor removed the fluid and sent off the culture. everything was sterilized and covered in tegaderm. the patient was put on a broad spectrum antibiotics prophylactically. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components: "the infusion filter prevents bacterial contamination and air from entering the purge lumen." section: warnings and cautions: warnings: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
8315 | 1220648-2025-29434 | 22283732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283732&pc=OZD | 05/25/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635537 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an implanted impella cp was removed due to failure by the clinicians to prime the pump appropriately. the catheter was inserted before the priming or connecting the yellow luers. the clinicians replaced the pump with another impella cp to continue support. | [5 empty] | The investigation of priming issue has been completed. priming issue: clinical details state that the cp was inserted into body without priming/connecting yellow luers. decision made to remove the pump and replace with a new cp. no data logs are available. the pump was returned and evaluated. the pump was primed and ran without issue in the lab. there was a brief period of high purge pressure that is unrelated to the reported issue. the root cause of the priming issue was use-related as clinical details state that the pump was inserted prior to priming/connecting yellow luers. d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29434. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-29434. h10 narrative erroneously stated the device was not received on the initial manufacturer device report number 1220648-2025-29434. | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8316 | 1220648-2025-29435 | 22283768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283768&pc=OZD | 05/28/2025 | 07/21/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660464 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed bleeding at the impella cp access site. pressure was held at the site, heparin was stopped, the dressing was changed, and quick clot was applied to manage the condition. two units of blood were given to the patient in response to the bleed. it was additionally reported that the patient's labs had hemolyzed. the support level was reduced, and two units of blood were given to treat the condition. later support was withdrawn and the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and access site bleeding has been completed. data review: there is no data log from 05/27 - 5/28 returned for review. product analysis: pump was returned for analysis. visual inspection found no issues on the pump. pump was conducted hemolysis test to observe if any defect on the pump. pump passed hemolysis test. conclusion: the root cause of hemolysis was unable to be determined as limited clinical details were provided and no issues were found on the pump. the root cause of access site bleeding was most likely use-related since angle matching resolved the bleed. b2 intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29435.
8317 | 1220648-2025-29436 | 22283805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283805&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, IN.C |  |  | US |  | 2026665561 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an implanted impella cp was explanted due to apical thrombus noted on an echocardiogram. the patient was successfully weaned from the pump and survived. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8318 | 1220648-2025-29437 | 22283950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283950&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672691 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the implanted impella cp was unable to advance above p-1 without encountering suction during patient support. imaging revealed that the cannula of the implanted pump was kinked. the suction events persisted throughout support at various levels of support. the patient was successfully weaned from the pump and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8319 | 1220648-2025-29438 | 22283956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283956&pc=OZD | 06/01/2025 | 08/25/2025 |  | 06/18/2025 |  |  | 06/01/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026659768 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, a high purge pressure alarm was encountered. no spikes in motor current were observed and no abnormalities or kinks were noted. the purge cassette was replaced; however, the issue remained. medical information for tpa-lyse was provided to the physician and support continued uninterrupted. the following day low pump flow was experienced and no volume was given. the impella 5.5 device was explanted and thrombus was discovered on the inlet. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation has been completed. the pump was returned for investigation. the data log was only available until 29-may-2050. the returned pump had a cut catheter; therefore, no investigation was performed on the low pump flow. thrombus was noted on the inlet cage. the motor side of the cut pump was tested for purge pressure using controlled manual syringe pressure. the purge was able to exit from the purge gap at normal purge pressure limits, justifying the absence of biomaterial in the purge gap. according to the clinical details, lysis therapy successfully resolved the high purge pressure issue. the cause of the high purge pressure issue was not determined without data log review. the cause of the low pump flow issue was not determined without data log review. in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the thrombus was unable to be determined due to insufficient clinical details. h6 medica device problem code was inadvertently omitted on the initial manufacturer device report number. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella 5.5 with smart assist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8320 | 1220648-2025-29439 | 22283965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22283965&pc=OZD | 05/28/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026656646 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during repositioning of the impella 5.5, the pump was pulled across the aortic valve (av) and into the aorta. the patient was taken to the operating room and under transesophageal echocardiogram (tee) still unable to readvance pump across the atrioventricular (av). the 5.5 was exchanged for a new one. | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced suction events. suction was resolved with volume and albumin. there was no patient harm. |  |  |  |  | B5 additional details from the clinical site regarding positioning failure. h6 updated to reflect the additional failure mode. | The investigation of positioning issues and low pump flow has been completed. product was not returned for review. data logs confirm suction alarms on reported date. positioning alarms on day of reported positioning issue. based on clinical description, the root cause of positioning issue was most likely use related since the pump was inadvertently pulled out of the left ventricle (lv) while repositioning of the pump. the root cause of low flow was most likely patient condition related since suction alarms were resolved by fluid compensation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8321 | 1220648-2025-29440 | 22284225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284225&pc=OZD | 05/27/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664167 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had disseminated intravascular coagulation requiring platelets to be administered and the impella was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
8322 | 1220648-2025-29441 | 22284279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284279&pc=OZD | 05/22/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 05/22/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026665353 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Atrial Fibrillation | Tachycardia | Ventricular Fibrillation | Ectopic Heartbeat | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 supported ongoing for approximately two weeks a surgical/coronary artery bypass graft patient. a day after implant, purge system blocked alarm occurred. one bag of tissue plasminogen activator was given via the purge. the current purge flow was abover 11 milliliters per hour. the plan was to change back to bicarbonate purge at the completion of the bag. the next day, there was concerns with the patients heart rate being tachycardic. testing was ordered to rule out atrial fibrillation. an episode of ventricular fibrillation occurred the following morning which prompted defibrillation and amiodarone bolus. the patient was on bilevel positive airway pressure post vf and the patient is now in atrial fibrillation with rapid ventricular response. the patient continued to have premature ventricular contractions and the plan was to turn of lidocaine. days later, amiodarone bolus was given for ventricular tachycardia. the following day, ectopy was noted to be decreasing and patient continued on amiodarone. the impella was planned to be weaned and explanted. the patient survived the explant. | [5 empty] | The investigation of vf/vt/af/at and high purge pressure has been completed. the pump was not returned for investigation. the data log shows a motor current spike followed by a 5-minute gradual rise in purge pressure and high purge pressure alarm. the purge pressure trend returned to normal after a few hours of the case run. the cause of the high purge pressure issue was most likely biomaterial ingestion, based on data log review. thrombus failure mode was added as an investigated failure mode based on the high purge pressure investigation. the root cause of the thrombus was unable to be determined due to insufficient clinical details. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8323 | 1220648-2025-29442 | 22284288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284288&pc=OZD | 05/25/2025 | 07/17/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616758 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male | 82 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp the patient exhibited oozing at the access site. anticoagulation was administered and forward tension sutures were applied. the patient survived. | [5 empty] | Access site bleeding: the root cause of access site bleeding is determined to be use issue because the bleeding was controlled by dressing change and angle matching of the repo sheath. access site bleeding: there was oozing from the surgical sites. the root cause of the access site bleeding was determined to be patient condition because of the bleeding from multiple surgical sites. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8324 | 1220648-2025-29443 | 22284380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284380&pc=OZD | 05/27/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026655877 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient endured an episode of ventricular tachycardia during impella cp support. no intervention was noted and support continued uninterrupted with the implanted pump. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8325 | 1220648-2025-29444 | 22284453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284453&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629265 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump placed for the patient being escalated from an intra-aortic balloon pump for surgery/coronary artery bypass graft. the pump was placed however there was an access site ooze. the pump remained on but later there was concerns of a retroperitoneal bleed and heparin was held systemically. the pump had sodium bicarb in the purge. the patient experienced vt and was shocked. blood panel was positive for heparin-induced thrombocytopenia. the pump still supports the patient despite the harms/injuries. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8326 | 1220648-2025-29445 | 22284475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284475&pc=OZD | 05/25/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657691 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a high purge pressure alarm occurred on the impella 5.5. the purge cassette was replaced and tissue plasminogen activator was added to the purge, but it was unsuccessful. additionally, a thrombus was identified. the patient was taken back to surgery for implantation of a new impella 5.5. the patient's outcome is stable. | [5 empty] | High purge pressure: clinical details state that purge flow had a sudden drop from 12 ml/hr to 1.5 then to less than 1.0ml/hr and low purge flow/high purge pressure alarms triggered. the issue was not resolved by replacing purge cassette. no purge line kinks were observed in the field. it was noted in the days leading up to the hpp event that the patient had been going on daily walks. data log review confirmed the drop in purge flow and a 40-second rise in pp following by stalled purge ticks, purge system blocked alarms, and purge pressure high alarms. the motor current and pump flow also slightly increased at the time of the pp rise. a de-airing procedure was completed 15 minutes after the pp rise and the purge cassette was replaced an hour after the pp rise, but the hpp did not resolve. the pump was explanted after 3 hours of hpp. the investigation of high purge pressure (hpp) and thrombus has been completed. the pump and a purge cassette were returned and evaluated. there was some blood found in the purge gap that was still present after purging the pump. no kinks were found on the returned pump or purge cassette. the returned purge cassette was not used for the reproduction testing since there was a tight knot present in the tubing. the pump was purged, it reproduced hpp, and the hpp resolved when cutting the catheter near the motor housing. no blood ingress was found in the purge line. it is likely that the high purge pressure began with a kinked purge line that then led to the blood in the purge gap seen in the lab. the location of the kink was unable to be identified and could have been a transit kink that resulted in blood ingress despite late efforts to replace the purge cassette. the root cause of the high purge pressure was most likely a kinked purge line since the data logs were characteristic of a kink (at an unknown location) with a sudden pp rise and purge system blocked alarms with no motor current (mc) spikes that led to a slight increase in mc and pump flow due to biomaterial or blood ingress. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. product damage (purge tubing kink): the fm product damage (purge tubing kink) was added since investigation found that the root cause of the high purge pressure was most likely a kinked purge line. no kinks were found on the returned pump or purge cassette; however, a tight knot was tied in the tubing of the purge cassette. the root cause of the product damage (purge tubing kink) was most likely patient movement as clinical details stated that the patient had been ambulating daily. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
8327 | 1220648-2025-29446 | 22284536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284536&pc=OZD | 05/27/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026667422 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient endured a couple episodes of ventricular tachycardia during impella 5.5 support. the condition resolved without intervention and support continued uninterrupted. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8328 | 1220648-2025-29447 | 22284747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284747&pc=OZD | 05/27/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657554 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support the patient had small runs of ventricular tachycardia and amio was started. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8329 | 1220648-2025-29448 | 22284882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284882&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487277 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and experienced oozing from the groin site and a hematoma. partial thromboplastin time (ptt) was still significantly elevated. argatroban was on hold until ptt in normal range. femostop was in place with good distal pulses. the patient was extremely labile with blood pressure. the next day, bumex drip was stopped with plan to give albumin. groin was stable and femostop was being removed slowly. the following day, femstop was back on. the patient became hypertensive overnight. a nitroglycerin drip was started for blood pressure. vitamin k was replaced and two units of blood were given. femostop was removed the next day and a new one was placed to low pressure for tracking oozing. argatroban was turned off. the patient was placed on a low dose of epipinephrine to maintain mean arterial pressure. after consulting the surgeon, the decision was made to remove the impella, place a dryseal sheath and then take the patient to the operating room for surgical cut down and closure of the femoral artery. possible impella 5.5 for post cardiotomy. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8330 | 1220648-2025-29449 | 22284947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22284947&pc=OZD | 05/25/2025 | 08/12/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 -EU |  |  | ABIOMED, INC. |  |  | US |  | 2026656009 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella 5.5 support and during support, there was high purge pressure with no flow. tissue plasma activator was added to the purge. the pump was explanted and replaced. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was returned for investigation. data logs were not provided for this case or could not be retrieved from the remote server. according to the clinical notes, high purge pressure was observed with no pump flow, and the lysis therapy did not work. no physical abnormalities were observed on the returned product. biomaterial was observed on the impeller and outflow cage. the pump was primed, and high purge pressure did not reproduce during testing. the cause of the high purge pressure issue was not determined without data log review.
8331 | 1220648-2025-29450 | 22285009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285009&pc=OZD | 05/27/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026662847 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the medical doctor inserted impella cp to the right groin per protocol and performed coronary angiogram. the medical doctor removed the impella sheath and secured impella per protocol. the patient moved mouth and got more sedation. the patient became nonpulsatile, with no pulse and blood pressure in the low 40s. cardiopulmonary resuscitation was initiated and there was a return of spontaneous circulation (cpr) achieved after 11 minutes of cpr. the patient was successfully transitioned to an impella 5.5 within 24 hours. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8332 | 1220648-2025-29451 | 22285085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285085&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632073 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient had an impella cp implant and after the implant, the pump started with flows 3.5 on auto. after peel away sheath was removed and repo sheath advanced, flows dropped to 0.5 on p4, position was confirmed with echo and was in ideal position. the pump was explanted and exchanged. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8333 | 1220648-2025-29452 | 22285114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285114&pc=OZD | 05/28/2025 | 07/16/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610822 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and during the implant, the anesthesiologist noticed a large clot formation around the coronary cusp. the doctor aborted the case because of the clot and sent the patient back to the icu on ecmo. the clinical rep confirmed the thrombus was determined to be preexisting to the impella implant and the pump was not related to the clot. the patient's condition limited ability for surgeon to place the pump. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | Delivery issue: the root cause of delivery issue is determined to be patient condition because of large clot around the coronary cusp and insertion was aborted. thrombus: thrombus failure mode was added because of the large clot formation around the coronary cusp. "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined".
8334 | 1220648-2025-29453 | 22285130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285130&pc=OZD | 05/25/2025 | 10/02/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026658304 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while weaning a patient after successfully impella cp support, the patient used a table knife to cut thru the purge side arm. purge bypass was performed by perfusionist but was not successful. no pressure could be built up from the automated impella controller. weaning was successful and the pump was explanted. the patient survived. | [5 empty] | The investigation of purge leak ¿ pump has been completed. impella cp was returned for evaluation. upon visual inspection, no purge sidearm was returned for evaluation and purge line was cut at red handle. no other abnormalities were noted. clinical details noted that the patient tried to cut through the purge sidearm with a table knife. a purge bypass was attempted but was not successful and pump was continued to be weaned then explanted. the root cause of the purge leak - pump was most likely due to a damaged sidearm based on clinical details however the exact location of the leak could not be determined as the purge sidearm was not returned with the pump. | The impella pump was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
8335 | 1220648-2025-29454 | 22285272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285272&pc=OZD | 05/28/2025 | 07/22/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026663329 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, after the pump was in the left ventricle, the pump was started. the motor current was high, and the pump stopped and was unable to restart. the implant was aborted and there was no patient harm. | [5 empty] | The investigation of pump did not start has been completed. impella 5.5 was returned for evaluation. pump did not start: upon visual inspection, no visual abnormalities were observed. no biomaterial or pump damage was observed. pump passed electrical testing. pump was connected to console and was able to run without issue for approximately 20 minutes. data logs were reviewed and pump was connected to console for approximately 1 minute before disconnecting then reconnecting to console. when the pump was re-connected and inserted into the patient, the pump was attempted to be started approximately 4 times with an immediate motor current spike and "impella stopped - motor current high" alarm was triggered each time. clinical details noted that once the pump was placed in a good position in the left ventricle (lv), a motor current high alarm was noted and the pump was unable to start. the root cause of the pump not starting could not be definitively determined as limited clinical details were provided and pump stop was unable to be reproduced with returned product. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
8336 | 1220648-2025-29455 | 22285285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285285&pc=OZD | 05/26/2025 | 07/21/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1180866 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the pump was being prepped for support and the automated impella controller (aic) would not prime the pump. the aic was exchanged and the pump was primed. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of priming issue has been completed. during the data and device analysis the root cause for the reported priming issue was not able to determine due to inability to reproduce.
8337 | 1220648-2025-29456 | 22285325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285325&pc=OZD | 05/28/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026672990 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a patient admitted in for the coronary percutaneous coronary intervention with impella cp support. it was reported that during impella cp support the distal leg became ischemic. the 14fr peel away sheath left in place from ccl. the peel away sheath was removed and repositioning sheath placed with resolution of flow to the lower extremity. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
8338 | 1220648-2025-29458 | 22285328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285328&pc=OZD | 05/24/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613200 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella cp support and during support, the patient developed limb ischemia in their leg. the pump was removed and the ischemia improved after the impella removal. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8339 | 1220648-2025-29457 | 22285330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285330&pc=OZD | 05/28/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664650 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. hemolysis has been noted based off urine as labs have not been drawn. the impella p-level was lowered. the next day, the patient improved slightly overnight. urine cleared to pale yellow and the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
8340 | 1220648-2025-29459 | 22285344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285344&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614956 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/25/2025 |  |  |  |  | A complainant reported that the impella cp was placed on a patient with acute myocardial infarction/cardiogenic shock. it was reported that during implantation, the purge cassette fell to the floor and became unsterile. a new purge cassette was used. additionally, there was a call from the physician that reported the had a suction alarm and pump was repositioned and pulled back a little. in addition to placement signal alarm on the automated impella controller (aic) and cardiopulmonary resuscitation after electrical storm. however, after optimizing the position, the aic values were in range. it was noted that care was withdrawn due to total heart failure and multi organ failure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8341 | 1220648-2025-29460 | 22285433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285433&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538948 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a cp pump support placed for the patient admitted in with acute myocardial infarction. after 11 days the motor current rose and pump failed/stopped. the pump was removed and not replaced. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ a.1- a.5 are unknown.
8342 | 1220648-2025-29461 | 22285461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285461&pc=OZD | 05/28/2025 | 07/22/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399951 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Malfunction | No | Yes | Battery Problem | Electrical Shorting | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a cp pump support ongoing for the patient and was transferred to the tertiary center by medflight. the automated impella controller (aic) was then seen to be malfunctioning with power cable having sparks and inability to charge with wires exposed. the aic was removed due to the noted issue. there was no reported harm or injury. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage during therapy and aic - battery charging/other issues has been completed. data analysis: the case start on 28th may 2025, the ac power was connected, as ac power alarms was cleared. however, after two hours of pump operation, the "ac power disconnected - alarm" (event 109) was triggered and could not be cleared until the console was shut down. device analysis: the console was returned for further investigation. upon inspection of the aic, it was discovered that the ac cord plug had detached from the ac cord, resulting in all three wires being broken. further examination of the plug revealed a blackened spot inside, which was caused by sparking at the live wire terminal. it appears that the wire covering or outer sheath was improperly cut, as per the vendor wiring instructions. also, the clamps were not tight enough, and light clamping marks were visible on the outer sheath, which likely contributed to the wire becoming loose and detaching from the plug and causing a spark. additionally, the batteries were checked by replacing the ac cord with a known good one, and no charging issues were found. the console was also tested with a known good pump and purge cassette, and no issues were observed. root cause: the reported complaint about the power cable malfunction with sparks and inability to charge was due to the loose wires inside the ac cord plug.
8343 | 1220648-2025-29462 | 22285497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285497&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026665349 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and developed limb ischemia concerns. a fasciotomy was performed and the pump was replaced with a new impella cp. later, it was determined that the flow issue was caused by the patient's blood pressure and not the impella pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
8344 | 1220648-2025-29463 | 22285503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285503&pc=OZD | 05/28/2025 | 07/11/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026666949 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be placed during attempted delivery due to the patient¿s anatomy and the implant was aborted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient anatomy.
8345 | 1220648-2025-29464 | 22285578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285578&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026661893 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had impella cp support and after the explant the patient had no flow to their foot. vascular surgery performed a clot removal which was successful. it was noted at that time that the pump dissected the superficial femoral artery and the artery was repaired. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
8346 | 1220648-2025-29465 | 22285666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285666&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623700 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. percutaneous coronary intervention (pci) started with impella cp support. the physician weaned the impella midway through pci due to the patient having oozing around the peel-away sheath. the physician was concerned about the oozing when the patient coughed. the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8347 | 1220648-2025-29466 | 22285698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285698&pc=OZD | 05/28/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655564 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, after 19 days there was a purge leak. the team replaced the leaking purge cassette and then found the sidearm was also a site of purge leak. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir."
8348 | 1220648-2025-29467 | 22285699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285699&pc=OZD | 08/09/2024 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513130 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used: ¿on (b)(6) 2024, hgb 12.5 on admission. on (b)(6) 2024, hgb 11.4 before cardiac surgery. on (b)(6) 2024, hgb dropped to 7.5.¿ it was further reported that care was withdrawn and the patient expired. there is not enough information to exclude the impella used as an associated factor to the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller. displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart. failure.¿
8349 | 1220648-2025-29468 | 22285708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285708&pc=PYX | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026655512 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, the limb had a fasciotomy performed for concerns of limb ischemia which may have been due to bp issues rather than true pump caused ischemia. the family made the choice to withdraw care after four (4) days of support and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8350 | 1220648-2025-29469 | 22285749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285749&pc=OZD | 05/25/2025 | 07/09/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026661707 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella 5.5 was aborted due to the patient¿s anatomy. an impella cp was implanted instead. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to resistance in axillary artery and patient's anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
8351 | 1220648-2025-29470 | 22285956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285956&pc=OZD | 05/28/2025 | 07/22/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026678546 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella cp was placed, however was started with poor flows. it was noted that the pump was kinked at prox cannula. the impella was pulled back in to descending aorta and kink resolved. several attempts to place pump w/o kinking were made w/o success. the doctor declined attempts to place a new pump stating that the aortic root was too small for the cannula to fit w/o kinking. pump was removed and there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of positioning issues and product damage (cannula kink) has been completed. the clinical details state that "a kink in the cannula was noted and pump pulled into descending aorta resolving kink. several attempts to position pump led physician to believe that the aortic root was too small, kinking cannula." the pump was returned and there was a significant kink seen in the cannula near the cage. based on the clinical details, the root cause of the positioning issue is most likely due to the patients condition as they had a small aortic root leading to a kink in the cannula. the root cause of the product damage (cannula kink) is most likely due to a cannula kink from due to crossing a small aortic root. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-29470.
8352 | 1220648-2025-29471 | 22285997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22285997&pc=OZD | 05/28/2025 | 07/28/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026664237 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that the impella cp was inserted and started but the cp cannula kinked, causing low flow and an impella stopped alarm. the cp was removed and a new impella cp was placed. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Returned cp pump visually inspected. cannula kink near outlet observed. no broken blade or biomaterial observed. low pump flow: clinical details do not specify how cannula kink occurred. no patient vessel abnormalities were mentioned. cannula kink confirmed on returned product. the root cause of the low pump flow issue was cannula kink as evidenced by kink seen on returned product, it is unknown how cannula kink occurred. product damage (cannula kink): cannula kink confirmed on returned product near pump outlet. the cause of product damage (cannula kink) was determined to be a cannula kink as evidenced by kink seen on returned product, it is unknown how cannula kink occurred. pump stop: cannula kink confirmed on returned product near pump outlet. the logs show an alarm #18 - impella stopped - reverse flow which per impella cp ifu, alarm #18 is retrograde flow meaning pump not running. the alarm occurrence here is likely as a result of a user input turning the pump down to p0. the returned pump passed electrical testing and there were no other abnormalities beside cannula kink. no problem was found as the root cause of the pump stop as returned product functioned as expected and no data log abnormalities were observed. | Returned cp pump visually inspected. cannula kink near outlet observed. no broken blade or biomaterial observed. low pump flow: clinical details do not specify how cannula kink occurred. no patient vessel abnormalities were mentioned. cannula kink confirmed on returned product. the root cause of the low pump flow issue was cannula kink as evidenced by kink seen on returned product, it is unknown how cannula kink occurred. product damage (cannula kink): cannula kink confirmed on returned product near pump outlet. the cause of product damage (cannula kink) was determined to be a cannula kink as evidenced by kink seen on returned product, it is unknown how cannula kink occurred. pump stop: cannula kink confirmed on returned product near pump outlet. the logs show an alarm #18 - impella stopped - reverse flow which per impella cp ifu, alarm #18 is retrograde flow meaning pump not running. the alarm occurrence here is likely as a result of a user input turning the pump down to p0. the returned pump passed electrical testing and there were no other abnormalities beside cannula kink. no problem was found as the root cause of the pump stop as returned product functioned as expected and no data log abnormalities were observed. | H6 medical device problem code 1503 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29471.
8353 | 1220648-2025-29472 | 22286005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286005&pc=OZD | 05/25/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026668071 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ectopic Heartbeat | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient endured occasional runs of ventricular tachycardia and ectopy during impella 5.5 support due to patient positioning which self resolved. support continued with the implanted pump. the patient was bridged to transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
8354 | 1220648-2025-29473 | 22286015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286015&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638759 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. a femstop was placed. the clinical rep confirmed that heparin was not withheld nor were blood products used. additionally, the patient had ventricular tachycardia (vt). the patient was shocked twice and support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8355 | 1220648-2025-29474 | 22286154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286154&pc=OZD | 05/21/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/21/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had multiple episodes for ventricular tachycardia (vt) prior and post impella cp insertion. once the cp was inserted mean arterial pressures were stable during persistent vt. anti arrhythmic were administered to control vt and subsided once drugs were administered and circulated. it was further reported that when the patient was transferred to another facility, the patient arrived bleeding from the right groin. femstop was placed and heparin was held. it was further reported that during explant of the impella cp that the patient had no pedal pulses in the right leg. a right thrombectomy and patch performed successfully. ultimately, the patient care was withdrawn and the patient expired. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8356 | 1220648-2025-29475 | 22286169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286169&pc=OZD | 05/28/2025 | 07/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636388 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp pump's flows were less than expected and its positioning was believed to be behind the papillary muscle. as the patient was stabilizing, the staff attempted to reposition the pump. when unsuccessful, the decision was made to remove the device and rewire the left ventricle for better positioning. a steelcore .018 wire was utilized to reimplant the pump, however, upon attempted removal, the wire became stuck in the impella. the patient began to rapidly decompensate, and the decision was made to replace the pump. despite the new pump implant, the patient arrested multiple times with a temporary rosc achieved after several rounds of cpr. when the patient arrested again shortly after, they could not be stabilized and the patient passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of positioning issues and low pump flow has been completed. the pump was not returned for investigation. the data log shows that the pump was started in auto mode p9 but was unable to achieve flow more than 3 l/min. there was one suction alarm at the end of the case. no positioning issues or motor current spikes were reported during the case run. according to the clinical notes, the doctor believed the pump was positioned behind the papillary muscle, and a failed attempt was made to reposition it. the doctor ultimately removed the pump and rewired it again with a steelcore 0.018 wire. the cause of the low pump flow issue was not determined since product was not returned and data log was not conclusive to derive root cause. the cause of the positioning issue was not determined since sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
8357 | 1220648-2025-29476 | 22286196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286196&pc=OZD | 05/25/2025 | 09/18/2025 |  | 06/18/2025 |  |  | 05/25/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1885740 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) displayed a yellow controller error alarm during patient support. a prompt appeared instructing the staff to switch to a backup controller. the aic was swapped to resolve the failure. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. based on the device and data analysis the root cause of the controller error could not be determined due to the inability to reproduce the issue.
8358 | 1220648-2025-29477 | 22286197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286197&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026681745 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient endured recurrent episodes of ventricular tachycardia (vt). antiarrhythmic medications were initially increased, and the patient was sedated in response to the condition. when the vt returned, the patient was shocked 4-5 times and received two rounds of cpr until stable. the patient's condition continued to worsen, and their family changed their status to do not resuscitate (dnr). the patient passed away after two days of support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8359 | 1220648-2025-29478 | 22286249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286249&pc=OZD | 05/28/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026679907 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp for support and during support, the patient had a computed tomography (ct) that showed a diffuse anoxic injury and multiple foci in intraparenchymal hemorrhage. neurosurgery consulted and recommended comfort measures. the patient expired on support. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
8360 | 1220648-2025-29479 | 22286296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286296&pc=OZD | 05/28/2025 | 07/22/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026669001 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient was sitting up in bed had a severe coughing episode and the impella slipped back into the aorta and the pump was explanted. there was no patient harm. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the positioning issue was most likely patient movement since the pump slipped back into the aorta during a severe coughing episode. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8361 | 1220648-2025-29480 | 22286326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286326&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026660972 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into ventricular tachycardia (vt) and was shocked two times, was given four minutes of cpr, and the pump remained on for support. pump was still supporting while planning for transfer to tertiary center for care escalation. after just over four days the patient expired when care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8362 | 1220648-2025-29481 | 22286397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286397&pc=OZD | 05/28/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026639105 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the pump stopped unexpectedly. the pump could not be restarted, and the decision was made to remove the pump. an impella cp pump was subsequently placed for ongoing support. the patient was noted as stable. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8363 | 1220648-2025-29482 | 22286404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286404&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026657609 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on day two of support, the patient went into ventricular tachycardia (vt) that only resolved with p level reduction. on the 4th day the pump had a purge leak when the pc tubing disconnected. the csc helped to trouble shoot and the pump remains on for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8364 | 1220648-2025-29484 | 22286428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286428&pc=OZD | 05/23/2025 | 07/21/2025 |  | 06/18/2025 |  |  | 05/23/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563938 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and had episodes of ventricular tachycardia requiring shocks and once was converted to atrial fibrillation. additionally, the patient had gastrointestinal bleed and heparin was on hold. the patient remains on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel and vf/vt has been completed. there were no abnormalities were observed with returned product and data logs showed no hard failures vascular damage - peripheral vessel: clinical details noted that the patient was experiencing a gi bleed while on support and required and ex-lap where a perforation was found in the right upper quadrant. the root cause of the vascular damage could not be definitively determined as limited clinical details were provided. vf/vt: the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. d6b added d9 device returned to manufacturer date added.
8365 | 1220648-2025-29483 | 22286429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286429&pc=OZD | 05/28/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/28/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607568 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, to prevent retrograde flow, amyloid precursor protein was increased p level and nipride, decreased milrinone. the patient was extubated and one unit of packed red blood cells was given. additionally, the patient had several episodes of ventricular tachycardia (vt) where he was shocked, the 5.5 was repositioned, however the (b)(6) believed that the pump was not causing the vt. regardless, the vt continued throughout support and again the pump was repositioned and on one episode it was noted of the patient being shocked 16 times. also, albumin was given and suction alarms resolved. the patient aspirated when he threw up stool, noted fever and elevated white blood cells and they started antibiotics. the patient continued on support. echocardiogram turndown looked and plan to possibly remove early next week. impella in aorta alarms when suctioning patient or coughing fits, however no more arrhythmias. the patient remains on support. | [5 empty] | The impella device remains implanted and was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
8366 | 1220648-2025-29485 | 22286514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286514&pc=OZD | 05/26/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650083 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  |  | Death | Yes | No | Malposition of Device | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp implanted in a patient was pulled 1 cm out of the body and the diastolic placement signal was 45 instead of 27. no placement signal alarmed. the patient expired on support though the medical staff stated the impella did not cause the death of the patient. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
8367 | 1220648-2025-29486 | 22286555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286555&pc=OZD | 05/29/2025 | 06/18/2025 |  | 06/18/2025 |  |  | 05/29/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1138237 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) abruptly powered off and turned back on after 10 seconds. the aic was replaced in response to the failure and there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
8368 | 1220648-2025-29487 | 22286559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286559&pc=OZD | 05/26/2025 | 06/26/2025 |  | 06/18/2025 |  |  | 05/26/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627685 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed a hematoma following transfer to another facility. one unit of packed red blood cells were transfused to obtain hemostasis. the patient was successfully weaned from the pump and survived. | On (b)(6) 2025 a computed tomography scan of head showed multifocal strokes, most notable right middle cerebral artery (mca). possibly due to hypoperfusion or clot related to have atrial fibrillation (afib) and not on anticoagulation. magnetic resonance imaging (mri) scheduled for further evaluation. the impella was weaned the following day with recovered cardiac function. |  |  |  |  | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely patient movement since hematoma was noted after transferring of patient to different facility. the investigation of thrombus, stroke, and af have not been completed. should new information or investigation results become available a supplemental report will be filed. instructions for use: impella cp with smart assist system for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b4 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-29487. h6 health effect - clinical codes 1770, 1729, 4440 and health effect - impact code 4639 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29487. h6 type of investigation, investigation findings, and investigation conclusions codes were erroneously entered on the initial manufacturer device report number 1220648-2025-29487. h10 investigation results and instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-29487.
8369 | 1220648-2025-29489 | 22286688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286688&pc=OZD | 05/29/2025 | 07/23/2025 |  | 06/18/2025 |  |  | 05/29/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026658339 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the physician pulled the impella 5.5 across the valve in attempt to optimize positioning but was not able to recross the atrioventricular (av). a decision was made to remove the impella and place a new one. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was most likely use related since the physician pulled the pump across the valve.
8370 | 1220648-2025-29490 | 22286711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=22286711&pc=OZD | 05/26/2025 | 07/22/2025 |  | 06/18/2025 |  |  | 05/27/2025 | 06/18/2025 |  |  |  |  | 06/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604012 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female | 47 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | Malposition of Device |  |  |  |  | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp support needing mechanical circulatory support. during support, foley revealing coca-cola colored urine. transthoracic echocardiogram performed at bedside, cp measuring 4.4 to 4.9 cm. dr. reviewed images. he is adamant that he does not believe repositioning at bedside is the best option. he explains due to lack of arrhythmias, lack of suction alarms (on p7), and pt having stable flows, he feels positioning is sufficient. discussion was had regarding possible impella outlet location on the valve, which could be corrected by repositioning by 1 cm retraction. it was suggested her hematuria and elevated lactate levels would benefit from repositioning. he re-affirmed his decision that her heart is too small and vasculature is so tiny that he believes bedside repositioning will rock the boat and cause bp/arrhythmia/flow issues. in addition the patient experienced acute decompensation and suction events on pump. 2 packed red blood cells, 1 fresh frozen plasma, 1 cryoprecipitate given. a hematoma development. no rp bleed seen, pressure dressing applied, along with products given. heparin paused during acute bleed, plan to restart when hemodynamics and coags stabilize. later it was reported that the impella stopped. motor current high alarm. this pump was restarted at p2, and then eventually removed about 2 hours later. the patients outcome is unknown. | [5 empty] | The investigation of hemolysis, low pump flow, vascular damage - peripheral vessel, positioning issues, and pump stop has been completed. temporary pump stop: upon visual inspection, a large purge gap was observed on returned pump. data logs were reviewed and pump stop alarm 13 was triggered at time of pump stop. pump was able to be restarted at p-2 with a variable motor current and increased purge pressures and decreased purge flows for 2 hours before explant. clinical details noted that a pump stop occurred on day 11 of support occurred and was able to be restarted and run for 2 more hours before explant. the root cause of the temporary pump stop was due to bearing wear beyond use indications per design trace matrix. hemolysis: returned pump had no damage to cage or epotek and no biomaterial present. pump was unable to be hemolysis tested as pump had bearing wear present. data logs were reviewed and no motor current spikes present at time of reported hemolysis. suction alarms occurring intermittently during reported hemolysis and placement signal is spiking throughout support. no positioning alarms noted at the time of reported hemolysis. clinical details noted that patient's pfhg increased on day 2 of impella support. patient was given multiple units of additional volume and pfhg was able to be decreased. it was noted that pump was position was slightly deep at 4.4 to 4.9cm but physician did not want to reposition with possible decreased flows. the root cause of the hemolysis was most likely patient condition related as data logs showed no ingestion events and clinical details noted that additional volume resolved the issue. vascular damage - peripheral vessel: clinical details noted that patient was bleeding from femoral a-line access site. dressings were applied and blood products given. the root cause of the vascular damage - peripheral vessel could not be definitively determined as limited clinical details were provided. positioning issues: clinical details noted that on second day of support, pump was measuring deep in lv and physician did not want to reposition. it was noted that patient had a small heart and repositioning pump would likely not help. data logs showed no positioning alarms at time of reported issue. the root cause of the positioning issue could not be definitively determined as limited clinical details were provided as to how the pump came out of position, it was not clear if it was related to patient's small heart anatomy, and data logs show no positioning alarms at time of issue. low pump flow: returned pump was unable to be tested for low flow as bearing wear was present on returned pump. a cannula kink was present on returned pump likely due to packaging of the returned pump and likely did not occur while in the patient. clinical details noted that patient experienced acute decompensation and a drop in blood pressures. suction events were occurring on the pump and multiple units of blood products were given. data logs do confirm suction alarms with low flows and some spiking in placement signal. the root cause of the low pump flow was most likely patient condition related as additional volume was administered to resolve the low flows per clinical details. h6 health effect - clinical code 4431 was erroneously entered on the initial manufacturer device report number 1220648-2025-29490. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿