24836 | 1220648-2024-12915 | 19678907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19678907&pc=OZD | 06/09/2024 | 08/01/2024 |  | 07/05/2024 |  |  | 06/09/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486012 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male patient of unknown ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis as evidence by hemolyzed labs. the pump was repositioned and plan to remove the following day. during the night, the patient was unable to resuscitate and expired per abiomed's medical safety officer review, there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. no product nor logs were returned to aid the investigation. based on the clinical details provided regarding the bulge in the septum and echocardiogram confirmation that the pump was in an acceptable position based on instructions for use (ifu) guidance, the root cause of the hemolysis issue was determined to be most likely patient condition. the suggested positioning of the pump in the ifu cannot account for unique patient anatomy.
24837 | 1220648-2024-12916 | 19679014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679014&pc=OZD | 06/10/2024 | 07/18/2024 |  | 07/05/2024 |  |  | 06/10/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474759 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 80-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient lost pulses in their right lower extremity. the profunda artery was blocked by the impella sheath position. no intervention was reported. the patient expired while on support. per abiomed's medical safety officer review, there is not enough information to associate the use of device with the patient's outcome. | [5 empty] | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
24838 | 1220648-2024-12917 | 19679058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679058&pc=OZD | 06/12/2024 | 08/05/2024 |  | 07/05/2024 |  |  | 06/13/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454971 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemodynamic instability | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 81-year-old male patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the pump had low pump flows. the pump was unable to flow above p2 without suction, and was unable to wean down bypass flows without significant hypotension. the pump was attempted to be repositioned and noted multiple cinches in catheter that prevented 5.5 from being advanced further. the staff was unsure if this was an issue with the manufacturing or if it was an issue with not deploying the blue bottom completely prior to attempting repositioning. the patient expired while on support. per abiomed's medical safety officer review, there is not enough information to associate the use of device with the patient's outcome. | [5 empty] | The impella pump was returned. visually inspected the pump and catheter was kinked close to the 50 cm marking. there were no issues found with the cath anchor. no biomaterial or cannula kink was seen. no other abnormalities were found. data logs were reviewed, suction alarms were found throughout the case and pump was unable to run greater than p-2 without suction. also, placement signal trending down. the cause of the positioning issue was use related. the cause of the low pump flow issue was most likely patient condition related. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
24839 | 1220648-2024-12905 | 19679202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679202&pc=OZD | 05/22/2024 | 09/09/2024 |  | 07/05/2024 |  |  | 06/14/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1339213 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon plugging in an impella 5.5 pump, the console displayed a controller error. the controller error noted that the pump could be started but that the placement signal would not be available. the aic (automated impella controller) was swapped in an attempt to troubleshoot the failure, but the issue persisted. the pump was then replaced, and the failure resolved. there was no patient harm resulting from the failure. | [5 empty] | The investigation for the console (aic) "controller error "yellow alarm has been completed. data logs were reviewed and a controller failure was not confirmed. instead, it was found that two controller errors were triggered due to the console identifying the optical sensor lamp as defective (event set #1039). preceding each controller error, the console attempted to reset the optical bench to resolve the issue, and this process also revealed that the bad lamp caused the 5s parameters to be out of spec. the aic was returned for evaluation. during testing, the 5s parameters were tested and confirmed to be out of spec, and matched the pattern of a defective lamp. the console was opened up, and the lamp was confirmed to not be emitting any light while the console was powered on. the lamp was confirmed to be defective. the reported complaint was determined to most likely have been caused by a defective lamp within the optical bench. | The impella device was received from the customer on (b)(6)2024 and therefore,¿an evaluation of the device is under way. upon investigation closure, a supplemental mdr will be filed.
24840 | 1220648-2024-12906 | 19679203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679203&pc=OZD | 06/28/2024 | 08/21/2024 |  | 07/05/2024 |  |  | 06/28/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516041 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 74-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that after medical staff exchanged for a repositioning sheath, the patient experienced uncontrollable bleeding around the sheath. a femoral angiogram revealed no extravasation around the access site. the medical staff put the impella cp at p2 and monitored the patient¿s vitals. the medical staff then decided to explant the impella cp due to persistent bleeding at the access site. the patient remained stable and survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, perclose x1 & 8fr. angioseal to achieve hemostasis as per the clinical description. added- h6 component code 925.
24841 | 1220648-2024-12918 | 19679444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679444&pc=OZD | 06/06/2024 | 07/30/2024 |  | 07/05/2024 |  |  | 06/06/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support as escalation of care from an intra-aortic balloon pump. during support, the patient had a large, laminated thrombus totally occluding the left atrium shown on transesophageal echocardiography. the family decided to withdraw care and the patient expired. per abiomed medical safety office, abiomed does not have enough information to associate the use of the device with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in section: entering cardiac output states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
24842 | 1220648-2024-12920 | 19679634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679634&pc=OZD | 06/11/2024 | 03/19/2025 |  | 07/05/2024 |  |  | 06/11/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438724 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during routine cassette and purge bag change the aic (automated impella controller) was unable to complete priming. multiple cassettes were tried and ongoing yellow alarm for not sensing purge fluid in tubing occurred. aic attempted to reconnect during the priming of the cassette to the yellow luer lock, this prompted another alarm to disconnect yellow luer, then the final prime finished. a new bag of purge was attempted again, and the exact same alarm continued to come up. the aic was exchanged. (patient passed - not enough information to associate use of the device with the outcome). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue - other has been completed. after reviewing the data logs, the reported priming issue was confirmed. there were numerous ¿flush fluid not detected¿ prompts found in the logs. the reported priming issue was unable to reproduced in the lab. the console was able to prime normally after multiple attempts. no problem was found with this aic. the device functioned/operated to specification. d.9 date device returned/information was added. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12920. d.2 revised common device name since the initial manufacturer device report 1220648-2024-12920 was submitted in accordance with updated procedures. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12920. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-12920 was submitted in accordance with updated procedures.
24843 | 1220648-2024-12921 | 19679673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679673&pc=OZD | 06/08/2024 | 07/30/2024 |  | 07/05/2024 |  |  | 06/08/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1280682 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a yellow alarm stating error aic. the aic was exchanged to the back-up aic and the issue was resolved. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. the automated impella controller (aic) logs showed power battery manager (pbm) 1 communication errors. the console was tested after arrival and the failure mode was reproduced so when the console was turned on, it immediately output a system self check fail. the super cap on the pbm was confirmed to have corroded the pbm communication trace next to it. the cause of the aic issue was contamination of a communication line on the pbm. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12921 in accordance with updated procedures. e.4 an incorrect selection was made on the initial manufacturer device report number 1220648-2024-12921. it is unknown if the initial reporter also sent the report to the fda. d.6a and d.6b dates were reported on manufacturer device report number 1220648-2024-12921 and should not have been. g.1 revised manufacturing site address section as previously entered incorrectly on initial manufacturer device report number 1220648-2024-12921. h.6 code 3233 results pending completion of investigation was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12921 and has been added to this report. the investigation has been completed.
24844 | 1220648-2024-12922 | 19679704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679704&pc=OZD | 06/10/2024 | 07/23/2024 |  | 07/05/2024 |  |  | 06/10/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient developed bleeding at their access site during impella cp support. the patient was given two units of packed red blood cells to treat the condition and femostop was placed to achieve hemostasis. support continued with the implanted pump following the intervention. (patient passed, not enough information to associate use of the device with the outcome). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the patient's access site bleeding has been completed. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
24845 | 1220648-2024-12923 | 19679956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679956&pc=OZD | 04/08/2024 | 07/09/2024 |  | 07/05/2024 |  |  | 04/08/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454722 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 34 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 34-year-old female patient with cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported leaking from the purge side arm. medical staff exchanged the purge cassette and noted the luer connections were adequately sealed. the purge flow was 112.7 and the purge pressure was 431. medical staff noted a crack on the purge side arm. medical staff decided not to remove the 5.5. excessive force had been applied. the patient remained stable on impella support until weaning and eventual explant. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. based on an analysis of pictures from the field staff, the cause of the pump purge leak was most likely the damaged sidearm, but the exact location of the leak remains unknown since the pump is not available for investigation. should any new information be received, a supplemental mdr will be filed. | The instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock; in section: cleaning states the following: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ added- h6 component code 481, h6 impact code 4627, h6 type of investigation 4112. g1-corrected mfr site name and address.
24846 | 1220648-2024-12924 | 19679976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19679976&pc=OZD | 05/08/2024 | 07/05/2024 |  | 07/05/2024 |  |  | 06/06/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146845 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported the automated impella controller (aic) had an issue where the open-door button was sticking. the purge door would not close. this issue was discovered in a demonstration in service, not with a patient. it was reported that during the investigation, the logs were reviewed and found that a file was deleted. there was a file with a system self-check failed issue and after the console rebooted, the system self-check failed file previous stored was deleted. no patient involvement. | [5 empty] | The investigation has been completed. the automated impella controller (aic) was received and tested and after the console rebooted, the log indicated that the same ¿sh.core¿ file had been deleted. this is an unreported failure mode. unable to reproduce the unexpected reboot or system self-check failed error forcefully by passing the ¿slay -s sigabrt sh¿ command via telnet. the root cause of the intermittent console reboot was not determined due to the inability to reproduce.
24847 | 1220648-2024-12925 | 19680228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680228&pc=OZD | 06/28/2024 | 08/13/2024 |  | 07/05/2024 |  |  | 06/28/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495280 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 47-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed ischemia in their leg during impella cp support. a loss of pulse was noted in the impella leg and a femorofemoral bypass was performed to treat the condition. support was maintained with the implanted pump. | [5 empty] | The impella device was discarded by the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- h6 component code 925.
24848 | 1220648-2024-12926 | 19680358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680358&pc=OZD | 06/11/2024 | 08/20/2024 |  | 07/05/2024 |  |  | 06/12/2024 | 07/05/2024 |  |  | 11/30/2025 |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025436538 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after removing the impella cp, the staff wanted to use a manta as a closure device, but it failed. then, the staff tried to closure with fem-stop, but it also failed. the patient was rushed to the operating room where the patient expired. cause of death is unknown however, the physician reports that the cause death could have been because of the massive bleeding from the access site or a stent occlusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-12926.
24849 | 1220648-2024-12927 | 19680363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680363&pc=OZD | 06/20/2024 | 09/06/2024 |  | 07/05/2024 |  |  | 06/20/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508072 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male patient implanted with an impella cp for mechanical circulatory support for high risk surgery. it was reported that the patient was in delirium and self-explanted the pump. there was no vessel/groin damage reported. the patient had some hemodynamic instability and ecmo was implant while patient was in the icu. patient remained stable on ecmo support. | [5 empty] | The investigation for the positioning issue has been completed. the impella device was not returned for evaluation. clinical information mentions that he pump was explanted by the patient in the state of confusion. therefore, the root cause of the positioning issue was use related to improper technique. g1 corrected mfr site name and address. | The device was not returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
24850 | 1220648-2024-12928 | 19680413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680413&pc=OZD | 06/24/2024 | 08/21/2024 |  | 07/05/2024 |  |  | 06/24/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506054 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the 76-year-old male patient's left foot pulse was weak and he extremity was cold. the pump was explanted due to poor leg perfusion. the patient was noted to be stable and on low dose catecholamines, although both legs were cold. | [5 empty] | The impella was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use for the impella cp with smartassist system states: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
24851 | 1220648-2024-12929 | 19680465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680465&pc=OZD | 04/11/2024 | 07/09/2024 |  | 07/05/2024 |  |  | 06/06/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025459822 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 36-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the pump support is ongoing when purge flow issues were noted, and the pump purge cassette was replaced to troubleshoot. no patient harm was reported. | [5 empty] | The instructions for use for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter states the following: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 type of investigation 4121, h6 impact code 4629, h6 investigation conclusion code 50. | Clinical reported no decrease in purge fluid upon visual inspection. insufficient product was returned for evaluation. the returned pump cassette had a damaged purge disk which leaked purge fluid. this could have happened during lab testing. high purge pressure alarms confirmed in the data logs. the data logs show a gradual build up in purge pressure and a concurrent decrease in purge flow that continued even after purge set was replaced. the root cause of the high purge pressure was most likely biomaterial related due to simultaneous rise in purge pressure and decrease in pump flow.
24852 | 1220648-2024-12930 | 19680497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680497&pc=OZD | 04/25/2024 | 07/05/2024 |  | 07/05/2024 |  |  | 06/07/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456422 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an 80-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that while weaning the patient off support at the end of the procedure a hematoma was observed, it was discovered a vessel had been damaged as the source. perclose unsuccessful. a 10x40mm balloon tamponade performed from right contralateral access for 20 minutes with combination of manual pressure also unsuccessful. a covera covered stent deployed and hemostasis achieved successfully. the patient was reported as stable. | [5 empty] | The investigation into the vascular damage-peripheral is complete. the device was not received for investigation, but the clinical information provided, the root cause of the vascular damage issue was most likely use related, tenting of artery due to sheath angle of entry confirmed as the only likely source as per the clinical description provided. instructions for use for the related event are as follows: section: warnings and cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿
24853 | 1220648-2024-12931 | 19680578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680578&pc=OZD | 06/25/2024 | 08/02/2024 |  | 07/05/2024 |  |  | 06/25/2024 | 07/05/2024 |  |  | 05/31/2025 |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356562 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that while implanting the impella cp on a 46 year old female patient, the patient went into the ventricular tachycardia (vt). the patient was shocked once and the impella was placed successfully. | [5 empty] | The investigation of vt/vf has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12931: e4 should have been left blank. g1 reporting contact facility name missing. g1 reporting contact fax number missing. | The impella device has not been returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
24854 | 1220648-2024-12932 | 19680584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680584&pc=OZD | 06/26/2024 | 12/03/2025 |  | 07/05/2024 |  |  | 06/26/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1384825 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that once the impella cp was connected to the aic (automated impella controller), it would not recognize that there was a catheter plugged in. another console was brought in, and it loaded the catheter as usual. the case was then completed. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g.1. mdr reporting contact email. d.2. device name has been updated. d.4. model number updated. d.9. return date has been added. updated h.3. to device evaluation anticipated. h.6. type of investigation code added.
24855 | 1220648-2024-12933 | 19680647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680647&pc=OZD | 06/27/2024 | 08/20/2024 |  | 07/05/2024 |  |  | 06/27/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025463363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 76-year-old female patient with acute myocardial infarction / cardiogenic shock had an attempted impella cp implant. medical staff reported extreme difficulty when attempting to place the sheath. an aortoiliac system showed a small and diseased common femoral artery and iliac arteries. the implant was aborted, and the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the difficulty inserting the introducer has been completed. the product was not returned for investigation. as per clinical aortoiliac system visualized which confirmed small and diseased common femoral artery and iliacs. the cause of the difficulty inserting introducer issue was most likely patient condition related due to small/diseased vessels.
24856 | 1220648-2024-12934 | 19680658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19680658&pc=OZD | 06/28/2024 | 08/21/2024 |  | 07/05/2024 |  |  | 06/28/2024 | 07/05/2024 |  |  | 02/28/2025 |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310417 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported multiple complications were encountered during support. the placement signal initially was inconsistent with that of the previously implanted pump. support continued because the flow raters were normal. the next day, medical staff withheld heparin from the purge line because the patient experienced persistent bleeding from their mouth and nose. support continued with sodium bicarbonate added into the purge line. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the access site bleed has been completed. the impella device was not returned for evaluation. clinical details state that during the second day of 5.5 support, heparin was stopped due to bleeding from the patient's mouth and access ports. the heparin remained off through that night, but no other clinical details were given about the bleeding. the root cause of the access site bleeding was most likely patient condition. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ added- d6b b1-changed product problem to no. h6 med dev prod problem 2917 changed to 2993.
24857 | 1220648-2024-12947 | 19681171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681171&pc=PYX | 07/01/2018 | 03/10/2025 |  | 07/05/2024 |  |  | 04/23/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A literature review was performed by abiomed, and a publication was found inclusive of an impella rp review of cases from 2007 to 2018 published/authored by the team at new york presbyterian columbia. the patient was noted to have had to be returned to the catheterization lab for impella rp malpositioning issues. manufacturer device report 1220648-2024-10643 will address another patient from this literature article. | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. literature citation: coromilas, e. j., takeda, k., ando, m., cevasco, m., green, p., karmpaliotis, d., kirtane, a., topkara, v. k., yuzefpolskaya, m., takayama, h., naka, y., burkhoff, d., colombo, p. c., & garan, a. r. (2019). comparison of percutaneous and surgical right ventricular assist device support after durable left ventricular assist device insertion. journal of cardiac failure, 25(2), 105¿113. https://doi.org/10.1016/j.cardfail.2018.12.005. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
24858 | 1220648-2024-12948 | 19681351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681351&pc=OZD | 06/26/2024 | 10/17/2025 |  | 07/05/2024 |  |  | 06/26/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504235 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support due to having a left main percutaneous coronary intervention (pci). the team noticed the placement signal disappeared. all other metrics still worked, and patient was hemodynamically stable. the pci was completed and the impella was successfully weaned and explanted due to the placement signal issue. there was no patient harm reported. | [5 empty] | The impella device has not been returned evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product or data logs were returned for evaluation. clinical details noted that a lithotripsy balloon was being used when the placement signal disappeared. all other pump metrics including flow and motor current remained stable. the distance between the pump and the lithotripsy device was not noted. the root cause of the optical signal issue was most likely due to a damaged sensor as a result of interaction with a lithotripsy balloon as ifu 0048-9007 rr notes that the use of shockwave intravascular lithotripsy can interfere or damage the impella optical sensor. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
24859 | 1220648-2024-12949 | 19681419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681419&pc=OZD | 06/27/2024 | 08/16/2024 |  | 07/05/2024 |  |  | 06/27/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490909 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Material Separation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella 5.5 support for heart failure/cardiogenic shock was getting from the bed to the chair, the impella cable ripped. the white impella cable was either stepped on or the staff thought they had more purchase on cable. the grey piece and red impella plug were torn apart, connected by only what appears to be one wire. hard pump stop followed. the patient was urgently taken back to the operating room for a new impella 5.5 placement. the patient is stable. | [5 empty] | The investigation of pump stop has been completed. visually inspected the pump and observed the catheter kink protector pulled off with broken motor cables at the red plug. the catheter was intact to the kink protector and no bond issues were observed. evidence of necking and damages to all three motor cables without any bond issue. the motor cables were damaged with likely excessive force during support. data logs were reviewed, and several 115 pump stop alarms with motor current spike of greater than 30000 ma found. pump stop was observed on 3rd day after 66 hours of support. the cause of the pump stop issue was an open electrical line due to excessive force during support. d9 date device returned to manufacturer added g1 reporting contact facility name missing on manufacturer device report 1220648-2024-12949. | The impella device has not been returned for evaluation at this time. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
24860 | 1220648-2024-12950 | 19681541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681541&pc=OZD | 06/28/2024 | 09/01/2024 |  | 07/05/2024 |  |  | 06/28/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502477 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 70-year-old male patient with cardiomyopathy had an attempted impella 5.5 implant. due to the patient's difficult/tortuous anatomy, the implant was unsuccessful. medical staff performed multiple attempts, and the catheter had significant kinking 15-20 cm from the impella motor. a new impella pump was inserted successfully. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was determined to be patient condition because of tortuous anatomy of the patient. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2024-12950. a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12950 in accordance with updated procedures.
24861 | 1220648-2024-12951 | 19681599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681599&pc=OZD | 06/28/2024 | 09/03/2024 |  | 07/05/2024 |  |  | 06/28/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450435 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during support with an impella cp on a 65 year old female in cardiac arrest, the impella appeared kinked upon checking the placement of the pump while performing cardio pulmonary resuscitation on fluro. physician opted to remove the pump and replaced it with a new impella while patient continued in cardiac arrest. the patient's outcome is unknown at this time. | [5 empty] | The investigation for the low pump flow has been completed. the product and data logs were not returned for review. based on clinical description, the root cause of low flow was most likely cannula kink from cardiopulmonary resuscitation. g.1 revised reporting contact facility name as it was inadvertently omitted from the initial manufacturer device number 1220648-2024-12951. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device number 1220648-2024-12951 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device has not been returned for evaluation at this time. upon completion of the investigation, a supplemental report will be submitted. the instructions for use for the impella cp with smartassist system state the following: section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
24862 | 1220648-2024-12952 | 19681883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681883&pc=OZD | 06/28/2024 | 08/19/2024 |  | 07/05/2024 |  |  | 06/28/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520696 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | User facility reported a 64 year old male patient had an attempted impella cp implant and it was unsuccessful. due to vascular anatomy, they were unable to insert impella into the left ventricle and as a result had to medically manage the patient. there was no harm reported. | [5 empty] | The impella device has not been returned at this time. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use for the impella cp with smartassist system states the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12952. g1 reporting contact facility name missing. h.6 code 4627 was reported incorrectly. new code was added to health effect - impact code.
24863 | 1220648-2024-12953 | 19683103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19683103&pc=OZD | 06/29/2024 | 07/30/2024 |  | 07/05/2024 |  |  | 06/29/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142853 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had a purge disc not detected alarm. upon inspection, the purge disc retainer was cracked. the aic was replaced. there was no patient harm reported. | [5 empty] | The aic has not been returned for evaluation at this time. however, upon completion of the investigation, a supplemental report will be submitted. | The console was received for investigation. after reviewing the data logs reported issue that purge disc couldn't be detected was confirmed. the purge disc retainer was found to be broken (crack on retainer). the root cause of the issue was a broken purge disk retainer. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12953: g1 reporting contact facility name missing h.6 code 4629 was reported incorrectly. new code was added to health effect - impact code.
24864 | 1220648-2024-12955 | 19683306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19683306&pc=OZD | 07/01/2024 | 02/06/2025 |  | 07/05/2024 |  |  | 07/01/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164687 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella aic (automated impella controller) had a controller failure error. the error occurred multiple times but resolved itself. it was further reported seeing high purge pressure error as well. the aic was exchanged. | [5 empty] | The investigation for the console (aic) controller failure-red alarm has been completed. data logs were reviewed which confirm that purge pressure sensor defective alarm was issued. the console was returned for evaluation. the purge sensor defective alarm controller error was reproduced on initial boot up of the console in the lab. inspection of the power of the purge pressure sensor revealed one of the voltages to be much lower than expected. when a known a good purge pressure sensor was swapped in. the failure resolved. the console's imc logs as well as reproducibility confirm that purge pressure sensor defective alarm was issued as a failure mode. the root cause of the alarm was the defective purge pressure sensor. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24865 | 1220648-2024-12954 | 19683390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19683390&pc=OZD | 06/25/2024 | 10/29/2024 |  | 07/05/2024 |  |  | 07/01/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484110 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 71 year old male patient on impella 5.5 support for heart failure/cardiogenic shock. it was reported that the sterile sleeve tore and was unsecure during repositioning. tegaderm was used to secure sleeve back into place. there was no patient harm reported. | [5 empty] | The impella has not been returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | B.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12954. e.1 added phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12954. | The investigation of product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added g1 reporting contact facility name missing on manufacturer device report 1220648-2024-12954.
24866 | 1220648-2024-12956 | 19683428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19683428&pc=OZD | 07/01/2024 | 10/24/2024 |  | 07/05/2024 |  |  | 07/01/2024 | 07/05/2024 |  |  |  |  | 07/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1667361 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella aic (automated impella controller) screen was black. the green lights to soft keys and power were still on and the blue light in the back was blinking six times. the patient pressures dropped and went up on levo temporarily. the aic was exchanged, and the patient went into vtach temporarily, but resolved on its own. the patient is stable. | [5 empty] | The investigation for the console (aic) power issues has been completed. data logs were reviewed which show the monitoring suddenly end on the complaint date. there are no indications for why the logs end abruptly. the console was returned for investigation and the symptoms described in the complaint were reproduced. upon booting up the console, the screen was black, serial communication to the console was non-functional and the soft touch buttons were unresponsive. while powered on, the impellatronic board, optical bench , and purge assembly all had power but the kbd was not being sent voltage from the 4-in-1. additionally there was not voltage to the i/o port which would leave the rlm unpowered. voltage from the pbm to the 4-in-1 was confirmed with and without ac power. a known-good 4-in-1 was swapped in and the console was able to boot up as expected with all the subsystems powered which is when the logs were able to be downloaded. the original 4-in-1 was reinserted with the symptoms returning. the cause of the aic shutting down mid-case and losing impella connect visibility was a defective 4 in 1 board. h6 clinical code 1914, h6 impact code 4644. | The device was not received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
24867 | 1220648-2024-12891 | 19664738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664738&pc=OZD | 06/26/2024 | 08/19/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446054 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient for a stemi (st elevation myocardial infarction) patient. after coding in the ccl (chronic lymphocytic leukemia) the team placed impella and consulted for cabg (coronary artery bypass graft). the pump on day 2 had saturated flows but, remained on for support along with ecmo (extracorporeal membrane oxygenation). the saturated flows additionally prompted the team to use tpa (tissue plasminogen activator) in the purge. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of saturated flow has been completed. data log review showed that the only abnormality on 6/26 was low pump flows rather than saturated. returned product was inspected and no abnormalities were found, specifically cannula kinks or biomaterial. the pump run in the lab successfully, and neither low nor saturated flows reproduced at various p levels. the root cause of the low flow was determined to be most likely patient condition. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report.
24868 | 1220648-2024-12892 | 19664742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664742&pc=OZD | 06/27/2024 | 08/21/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025491240 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and the introducer sheath had leaking from the peel away site. to control the bleeding, the a new introducer kit was used successfully. | [5 empty] | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12892: d4 model number. d6b device explant date. e4 should have been left blank.
24869 | 1220648-2024-12893 | 19664754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664754&pc=OZD | 06/27/2024 | 08/15/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella cp placement the patient tortuous vessel was noted. a decision was made to proceed with the short sheath. upon attempt to advance, the cp could not advance due to extremely tortuous vessel and impella kinked. impella was explanted and integrity of catheter was in question due to the kinking. a new impella was used and a long sheath due to the patient¿s vessel condition. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the patient had tortuous vessel. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added- h6 med dev prob code 2682, 2889 d4-corrected model number h6 impact code 4627 changed to 4629
24870 | 1220648-2024-12894 | 19664756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664756&pc=OZD | 06/27/2024 | 08/16/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485974 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an 85-year-old male noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was unable to pass through the vasculature. the cannula kinked during insertion via subclavian, noted under fluoroscopy. the impella was removed, and another cp was placed successfully using alternative access. there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. e4 should have been left blank on manufacturer device report.
24871 | 1220648-2024-12895 | 19664764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664764&pc=OZD | 06/27/2024 | 08/21/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 22-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump placed to support a patient admitted with multiple cardiac and systemic comorbidities along with ecmo (extracorporeal membrane oxygenation). on day 2 of the support air bubbles/emboli were observed in the vasculature and lv (left ventricle). the etiology was unknown, but the p level was adjusted. hemolysis was not confirmed to be a cause of bubbles. on same day the pump was explanted and escalated to the 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of embolism has been completed. data logs were reviewed, and during the first 24-hour period there were no abnormalities related to the purge pressure or flow noted. the only thing of note was ao placement signal spikes during that period, but that is unrelated to the failure mode. product was returned and both purge cassettes were de-aired and connected to the pump to run in lab for several minutes each. purge pressure and flow remained stable for both cassettes. there were no visual abnormalities with the product. the root cause of the embolism could not be determined as not enough clinical information was provided and there was no product damage noted. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12895. d4 model number. e4 should have been left blank.
24872 | 1220648-2024-12896 | 19664770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664770&pc=OZD | 06/27/2024 | 08/21/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484463 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 47-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported, that when transferring the patient from the cardiac catheterization lab table to the bed, the nurse inadvertently ripped off the purge tubing from the infusion filter site. medical staff then exchanged the cp. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. the root cause of the purge leak was most likely damage to the reservoir to filter joint. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12896. e4 should have been left blank. g1 reporting contact fax number missing.
24873 | 1220648-2024-12897 | 19664776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664776&pc=OZD | 06/27/2024 | 08/07/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1752120 |  | 0042-0040-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller had an issue detecting the purge disc. medical staff attempted multiple purge cassette replacements, but all with the same alarm. no patient involvement and/or harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. after reviewing the logs reported issue that purge disc could not be detected was confirmed. the purge disc flag was found to be broken (missing at the blue purge disc retainer). the root cause of the issue was a broken purge disc flag. a.1 and a.3 revised as information was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12897. d.2 common device name revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12897. d.5 operator of device revised as information was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12897. d.9 revised was the device available for evaluation as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12897. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12897 as this information is unknown. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12897 was submitted.
24874 | 1220648-2024-12898 | 19664788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664788&pc=OZD | 06/19/2024 | 08/24/2024 |  | 07/03/2024 |  |  | 06/28/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481019 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported purge fluid leaked from the luer connector connection. the patient remains on support despite leak and no harm reported. | Days after complaint initially was reported, the team determined that the leak was to be from the pump and not the purge connector (pc). the pump still remained on for support despite the leak. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D.9 updated as the pump and purge cassette were returned for investigation. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning : ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ b.5 and e.1 facility name were revised as this new information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12898. e.1 us state american samoa was entered incorrectly on the initial manufacturer device report number 1220648-2024-12898 and should of been left blank as this event took place in japan. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12898 as it is unknown the initial reporter also sent the report to the fda. g.3 date was entered incorrectly on the initial manufacturer device report number 1220648-2024-12898 and should of been 19-jun-2024. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12898.
24875 | 1220648-2024-12899 | 19664793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664793&pc=OZD | 06/19/2024 | 09/05/2024 |  | 07/03/2024 |  |  | 06/28/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456305 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male patient of unknown ethnicity and race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during the implantation of the peelaway sheath major bleeding occurred, and the patient¿s hemoglobin dropped to 5mg/dl. the bleeding was resolved surgically, a retroperitoneal hematoma was removed, and the groin was instrumented with a redon drain. an unknown amount of rbcs was administered. the patient remains on support and status is unknown. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- d6b, h6 component code 925 b2-selected death and added date of death. b5-added mso review. d4-corrected udi number. g1-corrected mfr site name and address. h1-corrected type of reportable event changed to death. h6 impact code 1802.
24876 | 1220648-2024-12900 | 19664811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19664811&pc=OZD | 06/28/2024 | 08/07/2024 |  | 07/03/2024 |  |  | 06/28/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1377267 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller was being prepped for patient use and there was a system self-check error. medical staff replaced the console. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) self check issue has been completed. the console data log showed a communication issue with the pbm (power battery manager) during the initial bootup, with error messages "sau_ser10 cmd seq.101 timeout. reply err and sem_post" and "post: path1/dev/name/local/sau_powermod_1 failed". due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the console was returned for evaluation and the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination. the root cause of the "system self check failure" was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. added- d9 device return date, h6 impact code 4629. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
24877 | 1220648-2024-12901 | 19665071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19665071&pc=OZD | 06/07/2024 | 07/25/2024 |  | 07/03/2024 |  |  | 06/07/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when repositioning the impella cp the sterile sleeve ripped and was taped by the physician. there was no adverse effect to the patient as a result of this event. additional information noted the patient became critically ill and remained on impella support. the decision was made to withdraw care, and the patient expired while on support. medical safety review of the event noted there is not enough information to associate use of impella device with the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage (sterile sleeve damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
24878 | 1220648-2024-12902 | 19665145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19665145&pc=OZD | 09/20/2023 | 08/05/2024 |  | 07/03/2024 |  |  | 06/11/2024 | 07/03/2024 |  |  | 11/30/2024 |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273495 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [5 empty] | INTRA-AORTIC BALLOON PUMP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complaints team forwarded article from interventional cardiology journal entitled: " first use of the impella 5.5 in a patient with cardiogenic shock to bridge to heart transplantation in poland." the article described the patient was admitted with cardiogenic shock as a result of st-segment elevation myocardial infarction treated by percutaneous coronary intervention. the impella cp pump was post-operatively placed, and the patient developed hemolysis and thrombocytopenia. the team made the decision to upgrade the device. the impella cp was explanted, and an impella 5.5 was placed. two days later, signs of hemolysis and thrombocytopenia ceased, and the patient¿s condition continued to improve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. publication reference: roman przybylski, et al. (2023). first use of impella 5.5 in a patient with cardiogenic shock to bridge to heart transplantation in poland. cardiology journal 2024, vol. 31 no. 2, 355-356. | The investigation for the thrombocytopenia and the hemolysis has been completed. product and data logs were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of thrombocytopenia was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿ g1- corrected mfr site name and address. h6 component code 4755 changed to 925.
24879 | 1220648-2024-12903 | 19665149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19665149&pc=OZD | 06/07/2024 | 07/26/2024 |  | 07/03/2024 |  |  | 06/07/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487637 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted for acute congestive heart failure exacerbation and was in process of being discharged home when he experienced cardiac arrest. an impella cp device was placed intra-operatively for mechanical circulatory support during percutaneous coronary intervention of left anterior descending, diagonal and right coronary arteries, and during support, the patient experienced hemolysis. the patient had peripheral artery disease, and the physician made the decision to implant the impella cp despite the risk for ischemia. at the end of the procedure femoral angiogram showed no visible flow beyond the impella cp insertion site. the physician placed antegrade perfusion sheath and a contralateral retrograde sheath to perform an external fem-fem bypass which resolved the leg ischemia. however, the patient did not survive cardiac event. further information noted the patient was too unstable for transfer, the family decided to make the patient a dnr-do not resuscitate, and the patient expired on support overnight. medical safety review of the event noted there is not enough information to associate use of the device with the patient's demise. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
24880 | 1220648-2024-12904 | 19665244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19665244&pc=PYX | 06/13/2024 | 08/13/2024 |  | 07/03/2024 |  |  | 06/13/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1620355 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient post cardiothoracic surgery was implanted with an impella rp device for bipella mechanical circulatory support. while on support, the automated impella controller (aic) had continuous low-pressure alarms after starting the impella rp support. the staff checked for leaks, and it was noted that the blue notch for the purge disc was partially detached. the console was replaced, and support continued. additional information noted the patient's outcome as expired. post cardiac surgery, the had a "very poor prognosis," and expired approximately after seven hours on support. medical safety review of the event noted there is not enough information to associate the use of the device with the patient's demise. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
24881 | 1220648-2024-12907 | 19667006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19667006&pc=OZD | 06/28/2024 | 03/23/2025 |  | 07/03/2024 |  |  | 06/28/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1646194 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 52-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the power cord on the associated automated impella controller (aic) broke. the medical staff swapped the aic due to the damage. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. the console logs confirmed instances of ac power disconnected alarms. the console was returned for further investigation. upon arrival the visual inspection revealed power cord damage. internal cables within the plug were found to be disconnected. the root cause reported complaint was most likely due to an insufficient length of the wire outer sheath or covering d2 common device name revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h5 selection was inadvertently omitted on the initial manufacturer device report number. h8 selection was inadvertently omitted on the initial manufacturer device report number.
24882 | 1220648-2024-12908 | 19667292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19667292&pc=OZD | 06/28/2024 | 08/21/2024 |  | 07/03/2024 |  |  | 06/28/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512821 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 82-year-old female patient with high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. it was reported that during a percutaneous coronary intervention (pci) case, the 14 french sheath (25 centimeters) was noted to be leaking around the valve. the patient was already on impella support when this occurred. the medical staff noted leaking at the orange and white colored connectors outside of the hemostatic valve. after the pci procedure, the impella cp and peel away sheath were removed. the sheath was then flushed by the side port, and no leaking was noted. leaking was however noted when flushing the 14 french sheath base towards the hemostatic valve. both 14f french peel away sheaths were flushed post-pci, and the 25 centimeters sheath was the one that leaked. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the introducer damage has been completed. the impella was not returned for evaluation. however, clinical details were not sufficient to determine the root cause. the root cause of the introducer damage was not determined since no product/photos of the damage were returned. added- h6 impact code 4641.
24883 | 1220648-2024-12909 | 19667900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19667900&pc=OZD | 06/15/2024 | 08/30/2024 |  | 07/03/2024 |  |  | 06/15/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1435042 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 heart pump for mechanical circulatory support during and after a high-risk coronary intervention. during impella support, the physician reported that the purge system was leaking, and it was discovered that the two wings holding the purge disc were broken off. therefore, the physician successfully switched to the backup console with assistance, the irrigation system was also replaced, and the problem was resolved. the patient was noted to be in stable condition. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. note: products-specific information is unknown, and therefore section d4 fields are blank. upon receipt of product-specific or any relevant additional information, a supplemental report will be submitted. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. data logs are irrelevant to this complaint. upon examining aic, it was observed that the purge disc retainer was broken off. the root cause of the issue was a broken purge retainer.
24884 | 1220648-2024-12910 | 19668828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19668828&pc=OZD | 06/18/2024 | 03/05/2025 |  | 07/03/2024 |  |  | 06/18/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1838368 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported the automated impella controller was received into service depot and during operation check, horizontal noise was observed on the display screen. after isolating the problem, the display was replaced, and the problem was resolved without recurrence. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) ¿ display issue has been completed. the console logs were not available and were not necessary to this investigation. the console was visually inspected and was observed to have significant lcd screen flicker. the root cause for the aic display issue was most likely a defective lcd screen. d2 common device name revised on manufacturer device report 1220648-2024-12910 in accordance with updated procedures. e2 and e3 were revised on manufacturer device report 1220648-2024-12910 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-12910 in accordance with updated procedures. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. note: initial reporter information is unknown and therefore section e 1 section is blank (excluding country). upon receipt of initial reporter or any relevant information, a supplemental report will be submitted.
24885 | 1220648-2024-12911 | 19668850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19668850&pc=OZD | 06/25/2024 | 08/30/2024 |  | 07/03/2024 |  |  | 06/25/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509335 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed a high purge pressure alarm. there were no kinks observed on the line. the purge cassette was replaced but the issue did not resolve. tissue plasminogen activator (tpa)protocol was initiated by adding tpa to purge line and the issue resolved. there was no reported patient harm. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The investigation for the high purge pressure has been completed. according to clinical, there was high purge pressure reported on day 6 of support. tissue plasminogen activator (tpa) was administered and the issue resolved. a partial pump with a severed catheter was returned for investigation. biomaterial was found in the inflow cage and around the impeller. no data logs were available for review. the root cause of the high purge pressure was determined to be biomaterial. added- d6b, d9 device return date, h6 impact code 4644 g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.
24886 | 1220648-2024-12912 | 19668892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19668892&pc=OZD | 06/25/2024 | 09/06/2024 |  | 07/03/2024 |  |  | 06/25/2024 | 07/03/2024 |  |  | 10/31/2025 |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412904 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp smart assist support and during the implant the patient had ventricular tachycardia (vt). the patient also experienced arrhythmia issues after their rotor procedure. the impella was repositioned, and the patient stabilized. it is unknown if the vt was related to the impella and it is unknown if the patient had prior arrhythmia issues. additionally, the impella was removed early due to an esophageal injury during intubation unrelated to the impella. reportable due to patient's vt and the early removal of the impella for an esophageal injury. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the ventricular tachycardia and vascular damage has been completed. per clinical details, patient experienced arrhythmia issues during implant and impella was removed early due to an esophageal injury during intubation unrelated to the impella. no product was returned for investigation. without additional clinical details, the root cause of the vt and vascular damage could not be determined.
24887 | 1220648-2024-12914 | 19669284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19669284&pc=OZD | 06/27/2024 | 08/20/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510310 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support was rolled on the side and impella insertion site started to bleed. hemoglobin dropped. heparin on hold and two units prbc (packed red blood cells) were transfused. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement. the patient on impella cp support was rolled on the side and impella insertion site started to bleed as per the clinical description. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.6 code 4755 was reported incorrectly on the initial manufacturer device report number. h.6 codes 925 and 4641 were added as they were inadvertently omitted from the initial manufacturer device report number. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
24888 | 1220648-2024-12913 | 19669325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19669325&pc=OZD | 06/27/2024 | 08/21/2024 |  | 07/03/2024 |  |  | 06/27/2024 | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489304 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient¿s anatomy would not allow for insertion of impella cp due to high bifurcation of the femoral artery on both sides. the physician had dilated up the right profunda and attempted to place cp, but it was unsuccessful. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to high bifurcation of the femoral artery on both sides. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
24889 | MW5157023 | 19681447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19681447&pc=OZD | 06/15/2024 | 06/28/2024 |  | 07/03/2024 |  |  |  | 07/03/2024 |  |  |  |  | 07/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CONTROLLER |  |  | OZD |  |  | 0042-000-US |  |  | ABIOMED, INC. | [8 empty] | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 83 YR |  | Male | 76 KG | Non Hispanic | Asian | Death | No | Yes | [7 empty] | Cardiogenic Shock | [12 empty] | Death | [5 empty] | IMPELLA CP SET WITH SMART ASSIST: (B)(6) 2024 | [22 empty] | Initial submission | [9 empty] | On (b)(6) 2024 at 23:30 the motor current was lost on an impella device and it began to alarm. no alarms were noted to have gone off on the device prior to this incident. all connections were inspected and the impella help desk was contacted to help troubleshoot. attempts were made to restart the motor but were unsuccessful. the patient expired an hour and 15 minutes after loss of current. the patient was found to have evidence of acute myocardial infarction and underwent code stemi for which he underwent cardiac catheterization. he developed worsening dyspnea and required emergent intubation. he required cardiac catheterization demonstrating critical multivessel coronary artery disease with critical lesions of the left main, left anterior descending coronary artery. he had a dominant left coronary circulation with small nondominant right coronary with severe disease. ejection fraction was depressed at 30 to 35% with mild mitral valve regurgitation. an impella device was placed due to the patient's cardiogenic shock. patient initially did not require pressors but then required increasing doses of levophed. troponin was noted to be at 17,000. patient is currently intubated. cvt was asked to evaluate the patient for consideration of possible coronary artery bypass surgery. past medical history of copd (chronic obstructive pulmonary disease).
24890 | 1220648-2024-12857 | 19653743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19653743&pc=OZD | 06/13/2024 | 07/22/2024 |  | 07/02/2024 |  |  | 06/13/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1268753 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that when hooking up the purge, it was noted that where the purge disc snaps in was completely missing from the automated impella controller (aic). the purge disc was unable to be placed at all into the aic. a new console was used for the case. there was no harm to the patient. | [5 empty] | A.1 patient identifier and a.5 ethnicity and race are unknown. the console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller purge disc issues has been completed. data review: no data review is needed for this complaint because it was only reported that the retainer was broken. device analysis: the purge disc retainer was found to be broken (retainer completely broken off). h6 medical device problem code corrected from 2994 no pressure to 1260 fracture.
24891 | 1220648-2024-12858 | 19653787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19653787&pc=OZD | 06/13/2024 | 08/21/2025 |  | 07/02/2024 |  |  | 06/13/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494702 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, there was high purge pressure and purge blocked alarms. the staff checked for kinks and no kinks were found. the purge system was de-aired and there was good flow for the purge fluid to yellow luer. once the yellow to yellow lures were connected, the alarm continued. the staff was aware that the pump could stop due to the pressures. the pump was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. the pump was not returned for investigation. data logs shows 12-minute gradual rise in purge pressure without any abnormalities related to motor current spikes or placement signal trends. according to the clinical report, the doctor was unsuccessful in troubleshooting high purge pressure by de-airing the purge system and checking for kinks in the purge tubing. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was returned for investigation. data logs show a 12-minute gradual rise in purge pressure without any abnormalities related to motor current spikes or placement signal trends. according to the clinical report, the doctor was unsuccessful in troubleshooting the high purge pressure by de-airing the purge system and checking for kinks in the purge tubing. the returned pump reproduced the high purge pressure during the priming test. no biomaterial or physical abnormalities were noted on the returned device. the pump was cut near the motor housing to verify the blockage after priming using manual controlled syringe pressure, and the purge was unable to exit from the motor. upon teardown, biomaterial was observed on the bearing. the cause of the high purge pressure issue was biomaterial, based on data log review and returned product investigation. d9 device returned to manufacturer date added g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-12858.
24892 | 1220648-2024-12859 | 19653879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19653879&pc=OZD | 06/14/2024 | 08/06/2024 |  | 07/02/2024 |  |  | 06/14/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500979 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, there were purge pressure high alarms. there was a spike in motor current and the pump stopped. the pump was successfully restarted. tissue plasminogen activator was added to the purge and then saturated flows were noticed. the pump stopped again and was unable to be restarted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the high purge pressure, saturated flow, and pump stop has been completed. clinical details states that during support, multiple purge pressure high alarms occurred. there were no purge pressure high alarms the day of the pump stops. the root cause of the high purge pressure was not determined since product was not returned and insufficient clinical detail was provided. when the pump stopped occurred, it was successfully restarted after the second attempt. the pump stopped again after a spike in motor current (mc) and lv signal shortly after. after tissue plasminogen activator (tpa) was initiated in the purge, saturated flows with lv signal rapidly climbing above ps was noted. shortly after, the pump stopped again and could not be restarted. the root cause of the pump stops and saturated flows was not determined since product was not returned and insufficient clinical detail was provided.
24893 | 1220648-2024-12860 | 19653950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19653950&pc=OZD | 06/19/2024 | 08/13/2024 |  | 07/02/2024 |  |  | 06/19/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the during routine check, the automated impella controller (aic) sounded an alarm for battery failure. the power switch under the aic switched off, then on without resolution of the alarm condition. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - battery failure has been completed. after reviewing the logs, the reported battery failure alarm issue was confirmed. there were numerous instances of transport connection blown/missing alarms found in the logs from the reported occurrence date. the console was tested after arriving in field service. the reported battery failure alarm issue was unable to be reproduced. results of running batttest showed that the health of the batteries were normal. reported battery failure issue was determined to be use related. it is likely that the console was booted up while the console was connected to ac power and the battery switch was disengaged since the issue occurred upon first boot of the console after routine functional check was performed on it.
24894 | 1220648-2024-12861 | 19654086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19654086&pc=OZD | 06/19/2024 | 08/20/2024 |  | 07/02/2024 |  |  | 06/19/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506261 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump had high purge pressure due to the purge system being blocked. the pump and purge cassette were replaced, and the procedure was successfully completed. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the high purge pressure was most likely biomaterial ingestion.
24895 | 1220648-2024-12862 | 19654139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19654139&pc=PYX | 06/19/2024 | 08/12/2024 |  | 07/02/2024 |  |  | 06/19/2024 | 07/02/2024 |  |  | 06/30/2025 |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024359045 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella rp device. on day two of support, a suction alarm occurred, and purge flow dropped. the issue was unable to be resolved, and the patient was noted to be unstable and without right ventricle support. consequently, the impella rp was explanted, and venoarterial extracorporeal membrane oxygenation support was emergently placed. following the event, the patient was noted to be stable. | [5 empty] | The pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella was returned for evaluation. during inspection, the inflow of the pump was filled with biomaterial and heavy biomaterial was found wrapped throughout the impeller blades. data logs were returned. the rise in purge pressure was gradual and there were motor current spikes corresponding to the start of the purge pressure rise. the purge flow and the pump flow was also noticed to be decreasing corresponding to the event. these all are the characteristics of a high purge pressure due to biomaterial. therefore, the root cause of the high purge pressure issue was ingested biomaterial.
24896 | 1220648-2024-12863 | 19654297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19654297&pc=OZD | 06/26/2024 | 08/21/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481098 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and they had a hematoma. a femostop was placed and support continued. during support, the repositioning sheath became occlusive distal to the access site. attempts were made to restore flow to no avail. the patient had a computed tomography performed revealing an occlusion by the impella and multiple thrombi in lower extremity. additionally, the patient also had red urine with plasma free 320. the decision was made to remove the impella. vascular surgeon performed a cutdown and arterial repair with fasciotomy and thrombectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of limb ischemia, hemolysis, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added.
24897 | 1220648-2024-12864 | 19654795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19654795&pc=OZD | 06/19/2024 | 08/21/2024 |  | 07/02/2024 |  |  | 06/19/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501963 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient experienced malperfusion on both legs (perfusion right < left) as the patient lost pulses at the limb. there was no perfusion distal to the right knee. the staff was advised several times to escalate care to impella 5.5 to re-establish perfusion on right leg. however, the patient had a distal perfusion cannula placed on both sides and the event was resolved successfully. the impella was successfully explanted six days later, and the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for the reported limb ischemia has been completed. the device was returned for evaluation. upon inspection, leak was found under leak test from piston-cylinder (unrelated to limb ischemia). additionally, clinical information was insufficient to determine the root cause. therefore, the root cause of the limb ischemia could not be determined. added-d9 device return date g1-corrected mfr site name and address. h6 component code 4755 changed to 925.
24898 | 1220648-2024-12865 | 19654876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19654876&pc=OZD | 06/20/2024 | 08/15/2024 |  | 07/02/2024 |  |  | 06/20/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439179 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the patient lost pulses on both limbs. the patient expired after a day of support. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. g1-corrected mfr site name and address h6 component code 4755 changed to 925 | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
24899 | 1220648-2024-12866 | 19654935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19654935&pc=OZD | 06/21/2024 | 09/09/2024 |  | 07/02/2024 |  |  | 06/21/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025471401 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old female patient was implanted with an impella 5.5 device for mechanical circulatory support for an unknown indication. it was reported there was flow saturation (p6 - max/min/mean all 5.2 lpm) leading to pump explant. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of saturated flow has been completed. the pump was returned for investigation. data was not provided for investigation. a biomaterial was observed on impeller. no other physical abnormalities were observed. the cause of the saturated pump flow issue was biomaterial. d9 date device returned to manufacturer added.
24900 | 1220648-2024-12867 | 19655060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655060&pc=OZD | 06/12/2024 | 08/16/2024 |  | 07/02/2024 |  |  | 06/12/2024 | 07/02/2024 |  |  | 08/31/2025 |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024389861 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the procedure was aborted due to the patient's aortic anatomy. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because case was aborted due to aortic anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12867 a5 race was inadvertently left blank, e4 should have been left blank, g1 reporting contact fax number missing.
24901 | 1220648-2024-12868 | 19655076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655076&pc=OZD | 06/14/2024 | 04/15/2025 |  | 07/02/2024 |  |  | 06/14/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATEC IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119120 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported, upon plugging the new device into the automated imeplla controller (aic), it displayed an "impella 2.5" device with instruction to connect the gray cables. however, the device being connected was an impella cp and progress to the next step was not possible in the moment. eventually after unplugging and re-plugging the aic several times, it allowed progress as an impella cp to the next insertion step. there was no report of harm to the patient. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - pump issue has been completed. the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-12868 in accordance with updated procedures. d6b should have been left blank on initial manufacturer device report 1220648-2024-12868. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-12868 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-12868 in accordance with updated procedures.
24902 | 1220648-2024-12869 | 19655090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655090&pc=OZD | 06/18/2024 | 08/21/2024 |  | 07/02/2024 |  |  | 06/18/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025481561 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during initial implant for 72-year-old male patient, the impella cp had suction and a spike with high motor current, which was displayed in the impella connect. the physician then pulled a bit too hard and the impella cp slipped out of the left ventricle. the impella was attempted to be repositioned and position was good, however, when the impella was started, the optical sensor failed and there was a max of 0.9l. there were no kinks on the cannula, no thrombus or biomaterial observed. therefore, the physician decided to explant the pump and complete the procedure without impella support. there was no reported patient harm. | [5 empty] | The investigation for the positioning issue and low pump flow has been completed. product and data were not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of low flow was unable to be determined as limited clinical details were provided and no product, logs were returned for review. the cause of the optical signal issue was not determined since product and data logs were not returned. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24903 | 1220648-2024-12870 | 19655100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655100&pc=OZD | 06/21/2024 | 08/14/2024 |  | 07/02/2024 |  |  | 06/21/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492500 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the staff had difficulty with the peel-away sheath due to the patient's calcification. the patient developed a hematoma and vascular surgery was consulted. the patient went to the operating room. it is unknown what was performed to resolve the hematoma at the time of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the insertion issues has been completed. the product was not returned. the cause of the vascular damage issue was most likely patient condition related due to calcified vessels. h6 med dev prob code 2682 changed to 1316.
24904 | 1220648-2024-12871 | 19655120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655120&pc=OZD | 06/25/2024 | 09/09/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025482102 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp had high purge pressure and purge flow decreased after cp was started. the purge cassette was replaced, but there was no improvement. the impella was removed. the removed impella cp was checked, but no thrombus was found attached to the outflow cage. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of high-pressure purge (hpp) has been completed. data logs show that upon starting the pump there is a 6-minute rise in purge pressure and drop in purge flow after a mc spike. hpp and purge system blocked alarms are triggered. the pump is removed 30 minutes later with no hpp resolve. biomaterial was found in the purge gap. when purging pump, hpp was reproduced. running the pump resolved the hpp and biomaterial was expelled from the purge gap. the cause of the high purge pressure was biomaterial in the purge gap.
24905 | 1220648-2024-12872 | 19655182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655182&pc=OZD | 06/24/2024 | 09/09/2024 |  | 07/02/2024 |  |  | 06/24/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025479165 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the patient experienced hemolysis. the pump was noted to have been replaced due to increased motor current and the possibility of thrombus. no blood products were given. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the hemolysis and thrombus has been completed. clinical details state that an impella cp was implanted, and at some point in the next 3 days, hemolysis occurred. the pump was then explanted because of high motor current and the probability of a thrombus. the product was returned and evaluated. there was thrombus found wrapped around the impeller. no other abnormalities were found. the logs showed an increase in purge pressure with a small spike in motor current on the second day of support. there were several suction alarms shortly after. the motor current was also higher than expected for the set p-levels after this. therefore, the root cause of the hemolysis was most likely biomaterial. h5-changed to yes. h6 impact code 4629 changed to 4627.
24906 | 1220648-2024-12873 | 19655185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655185&pc=OZD | 06/25/2024 | 09/09/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025455937 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was bleeding reported at the access site. manual pressure was held. a cat scan revealed that the artery had been damaged when the impella sheath was sutured in place, the bleeding was not at the impella insertion site. the issue resolved with manual pressure. the patient remains on support and is reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage - peripheral vessel issue was use related to improper technique. d6b added h6 code 1884 and 4755 were reported incorrectly on manufacturer device report 1220648-2024-12873. new code was added to component code.
24907 | 1220648-2024-12874 | 19655191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655191&pc=OZD | 06/24/2024 | 08/13/2024 |  | 07/02/2024 |  |  | 06/24/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025481858 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 39-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the patient developed an access site hematoma and was taken back to operating room and had a cutdown and wash out of the small bleed found on the graft anastomosis. the patient is stable and remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
24908 | 1220648-2024-12875 | 19655215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655215&pc=OZD | 06/25/2024 | 07/16/2025 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483305 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 78-year-old male patient with post-cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that medical staff suspected the patient of having hemolysis during support. plasma-free hemoglobin (pfhgb) was found to be elevated at 2500, and a second lab draw still showed elevation. medical staff did not administer heparin to the patient until after the fourth day of impella support. medical staff then removed the 5.5, and the patient survived and remained stable. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed pump ran without issue during support. there were suction alarms were triggered in early of the case but resolved. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The investigation of hemolysis has been completed. data review: data logs showed pump ran without issue during support. there were suction alarms were triggered in early of the case but resolved. no motor current spikes. ps began to trend down on 6/24 through 6/25. pump was returned for analysis. visual inspection found no issues on the pump. pump passed hemolysis test. the root cause of hemolysis was most likely patient condition related since the pump passed hemolysis testing and good positioning was confirmed. d9 device returned to manufacturer date added h6 investigation finding revised based on the returned device investigation results. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-12875 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-12875 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12875.
24909 | 1220648-2024-12877 | 19655236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655236&pc=OZD | 06/26/2024 | 09/30/2025 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1335325 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 64-year-old female patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) had issues with the pump screen - at the start, the pump screen appeared normal, but when starting the pump, the aic did not show the pump flows or the placement left ventricle waveforms. medical staff then replaced the aic. no patient harm has been reported. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) issue has been completed. console logs from the reported day of the event are consistent with the complaint and displayed "algs suspended opm signal invalid" and "invalid sample due to snr below minimum" error messages (low snr was likely due to air gap at the pump connector), and the case start wizard prompted operator to disconnect the pump due to "optical sensor system error" (rtlog data confirms very low snr and at times invalid placement signal). pump was reconnected within 1 second its eeprom hadn¿t been read, but g0 calibration succeeded again with a new valid g0 value. after this moment, although optical snr was high, and raw placement signal was valid, the placement signal on screen (and calculated placement signal in lt and rt logs) had persistent value of 0, which resulted in placement signal low alarm (#439). the aic was returned for evaluation. the failure was not reproducible upon initial boot. persistent debris was observed along the core of the front optical connector. a known good pump and purge system was run with the console with no observable anomaly. the root cause of the optical signal issue was not determined since the issue could not be reproduced. added-d9-date of device return corrections to the initial mfr report: b1-added adverse event, b2-selected death and added date of death, b5-mso review added, d2-corrected common device name, g1 mdr reporting contact email, g4-pma number added, h1-changed to death, h6 impact code changed to 1802. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. device analysis: the console was inspected in the lab. the failure was not reproducible upon initial boot. persistent debris was observed along the core of the front optical connector. section 23 of the preventative maintenance was performed and the console found to be within specification. a known good pump and purge system was run with the console with no observable anomaly. note: the analysis of console logs indicated that the observed behavior was fully consistent with three prior complaints: (b)(4), were an investigation into a possible aic sw defect (gid-dft-3734730) concluded that the software operated as per design and correctly responded to all triggering criteria of pump disconnection and connection. conclusion: the root cause of the optical signal issue was not determined since the issue could not be reproduced. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b,h,h6: adding death as part of a retrospective review of optical signal issues in accordance with fda recommendation.
24910 | 1220648-2024-12878 | 19655241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655241&pc=OZD | 06/25/2024 | 08/20/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1762205 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a flickering screen. the aic was send back for investigation. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. the console was received for investigation and the root cause for the flickering screen was not determined due to the inability to reproduce. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12878. d9 device returned to manufacturer inadvertently was marked no. e4 should have been left blank . g1 reporting contact fax number missing.
24911 | 1220648-2024-12879 | 19655253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655253&pc=OZD | 06/24/2024 | 09/18/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025464960 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the automated impella controller (aic) displayed "impella in the aorta". the transthoracic echocardiogram (tte) showed the impella was in contact with structure and repositioned. the issue resolved and the alarm disappeared. there was no reported patient harm, and the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. clinical details stated that the impella cp was successfully implanted through the right femoral artery of a patient. during the first day of support, the automated impella controller displayed an "impella in the aorta" alarm. in the subsequent transthoracic echocardiogram (tte), the impella was probably correctly in the ventricle. however, the tte showed that the impella is in contact with a structure, and its position was improved by rotating it. the alarm resolved after this repositioning. the root cause of the optical signal issue was not determined since no product/logs were returned. d.4 revised unique identifier (udi) # was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12879. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12879 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site address line 1, address line 2, city and zip code as they were entered incorrectly on the initial manufacturer device report number 1220648-2024-12879. h.6 code 3009 was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12879 and a new code was added.
24912 | 1220648-2024-12880 | 19655267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655267&pc=OZD | 06/25/2024 | 08/26/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479473 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and there was a hematoma at the access site. one unit of packed red blood cells was provided to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D.6b explant date was added. the investigation for the reported access site bleed has been completed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12880. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12880 in accordance with updated procedures. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12880 as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12880 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
24913 | 1220648-2024-12876 | 19655268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655268&pc=OZD | 06/25/2024 | 08/22/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  | 10/31/2024 |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268247 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during implant procedure for 68-year-old male patient, the impella cp was unable to be delivered through the aorta due to high amount of calcification and tortuosity, and moderate stenosis. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12876: d4 model number. e4 should have been left blank.
24914 | 1220648-2024-12881 | 19655288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655288&pc=OZD | 06/25/2024 | 08/20/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that towards the end of the impella cp procedure the patient experienced oozing around the access site. the physician placed a femstop and two forward tension sutures and angle of entry dressing. the oozing was resolved. | Additional information provided: as the team was coming to do the heart transplant, the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12881: b7 relevant history missing. d4 model number. d6b explant date. e4 should have been left blank. g1 manufacturing site fax number.
24915 | 1220648-2024-12882 | 19655292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655292&pc=OZD | 06/25/2024 | 08/16/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025486377 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male patient of unknown race with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that air bubbles were present in the patient¿s ventricle, which was revealed on echocardiogram. staff concluded that the bubble was a possible cavitation. additionally, the patient had hemolysis. the pump was repositioned, and the p level was reduced to p-4 and the patient tolerated well. it is unknown if the hemolysis was resolved by the repositioning and p level reduction. the impella 5.5 was exchanged for left ventricular assist device as part of the plan of care, unrelated to the reported events and is reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of embolism and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-12882.
24916 | 1220648-2024-12883 | 19655308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655308&pc=OZD | 05/26/2024 | 09/18/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443547 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a console alarm/failure occur after removal. the console displayed a controller failure. there was no noted harm or injury. the event is a reportable malfunction. later in review of the automated impella controller logs at abiomed the engineering team discovered high purge pressure alarms noted on the console with the pump support. | [5 empty] | The investigation for the high purge pressure has been completed. data logs were reviewed which showed high purge pressure with purge pressure steadily increased and purge flow decreased in the last 8 days of support until it triggered high purge pressure alarms. there were several suction alarms during the days leading up to the rise in purge pressure and no significant motor current spikes during the rise in purge pressure. according to the data logs, there was heparin in the purge for only the first 4 days of support. it is unknown if heparin was administered otherwise. the gradual increase in purge pressure most likely indicates a gradual buildup of biomaterial in the purge gap. the root cause of the high purge pressure was most likely biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24917 | 1220648-2024-12884 | 19655320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655320&pc=OZD | 06/26/2024 | 08/21/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025486369 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 had a high purge pressure purge system blocked alarm overnight and was still active in the morning. tissue plasminogen activator purge solution was started. the patient remains on support and is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. according clinical data, there were high purge pressure alarms, so they started tissue plasminogen activator (tpa) purge solution. the high purge pressure alarms resolved about 48 hours later. tpa bag finished and the purge was switched back to bicarb. after 43 days of support, the patient was bridged to lvad. the root cause of the high purge pressure was most likely biomaterial because of a gradual increase in purge pressure that was resolved after adding tpa. added- d6b, h6 impact code 4644.
24918 | 1220648-2024-12885 | 19655335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655335&pc=OZD | 06/25/2024 | 08/17/2024 |  | 07/02/2024 |  |  | 06/25/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025495362 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an 80-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient experienced irregular rhythms, an echocardiogram was done, and pump was turned down to p-3. it was also reported that upon removing the cp under fluoroscopy there was thrombus in the left iliac artery and left common femoral artery. multiple attempts were made with pneumbra to remove clot, but it was unsuccessful. the patient was sent to surgery for surgical cutdown. the patient remained hemodynamically stable throughout the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of ventricular tachycardia/ventricular fibrillation (vt/vf) and thrombus has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf and thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12885. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12885 as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12885 in accordance with updated procedures.
24919 | 1220648-2024-12886 | 19655358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655358&pc=OZD | 06/26/2024 | 08/16/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1303848 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during prep for emergency impella implant procedure, it was noticed that following points 3a and 3b on the automated impella controller (aic), further steps were automatically skipped. the screen then displayed the placement screen, and the pump was set to p-0. the purge system was de-aired, and a new purge cassette was placed, and the pump was successfully implanted. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) display issue has been completed. the console was returned for evaluation. data logs were reviewed finding that the console performed as expected during all of this. upon functional inspection of the console, the uninitiated pump was able to go through all of the case wizard steps versus an initiated pump skipped the last steps of the case wizard reproducing what was done in the field. there was no malfunctions with the console. there was no issue found during the case. the device functioned/operated to specification as an initiated pump is supposed to exit the wizard. added- d9 device return date d4-corrected model number g1-corrected mfr site name and address h8 changed to reuse
24920 | 1220648-2024-12887 | 19655360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655360&pc=PYX | 06/26/2024 | 10/21/2025 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025502642 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported an 81-year-old female patient postoperative cardio-thoracic surgery was implanted with an impella rp for mechanical circulatory support. the right ventricle was noted to be thrombosed and had a washout and decompensated so the team placed a rp flex for support. the pump was placed but the flows were low and so sodium bicarb was used. unable to resolve with repositioning or change in p level. the decision was made to withdraw care due to poor prognosis and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started, motor current (mc) spiked, with placement signal (ps) trending upwards followed by low flow. this event remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial. | This report is being filed as part of a retrospective review of historical records. the investigation of low pump flow has been completed. the impella device was not returned for evaluation. data stick was returned, not the pump. data review: data logs confirmed the failure mode and clinical narrative. logs showed after pump started, motor current (mc) spiked, with placement signal (ps) stepping upwards with low flow. this event remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12887. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-12887 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-12887 in accordance with updated procedures. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2024-12887.
24921 | 1220648-2024-12888 | 19655371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655371&pc=OZD | 06/26/2024 | 08/30/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025478971 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during support for 69-year-old male patient, the impella cp migrated to the proximal ascending aorta (paa). the patient and his hemodynamics, despite being septic, were improving. however, the pump was explanted as a result. the patient is stable. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. | The impella device was not received from the customer and therefore an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24922 | 1220648-2024-12889 | 19655393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19655393&pc=OZD | 06/26/2024 | 08/14/2024 |  | 07/02/2024 |  |  | 06/26/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025479306 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during repositioning of impella 5.5 at the bedside by the surgeon, the sterile sleeve broke away. the exposed catheter was disinfected and a tegaderm was placed. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-12889.
24923 | 1220648-2024-12890 | 19656225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19656225&pc=PYX | 04/23/2024 | 03/20/2025 |  | 07/02/2024 |  |  | 04/23/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received an italian publication titled "timing, positioning, and ancillary therapies: the triad for successful impella rp". the publication outlined three patient cases. this report represents the third patient which had pump malpositioning and kinking. the patient expired. manufacturer report number 1220648-2024-10844 reflects patient number (b)(6). | [5 empty] | A.1, a.4 and a.5 are unknown. b.2 date of death is unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. literature citation: bertolin, s., maj, g., secco, g. g., pullara, a., gherli, r., cardinale, a., d¿ettore, n., cavozza, c., audo, a., & pappalardo, f. (2023). timing, positioning and ancillary therapies: the triad for successful impella rp. journal of cardiovascular medicine. https://doi.org/10.2459/jcm.0000000000001563. | The investigation of positioning issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined g1 reporting contact email revised on manufacturer device report 1220648-2024-12890 in accordance with updated procedures.
24924 | 1220648-2024-12813 | 19657124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19657124&pc=OZD | 06/21/2024 | 09/03/2024 |  | 07/02/2024 |  |  | 06/21/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245452 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | No patient involvement: us complainant reported the automated impella controller (aic) was being serviced and there was a system self check error. | [5 empty] | The impella device was received from the customer on (b)(6) 2024 and therefore, an evaluation of the device is under way. upon investigation closure, a supplemental mdr will be filed. | The investigation for the automated impella controller self check error has been completed. the data logs from the console confirm that the system self check failure alarm was issued. the console was returned for evaluation. upon functional inspection, the console booted up in system-self check failure. inspection of the a power battery manager (pbm) under a digital microscope confirmed that the corrosion of copper traces in vicinity of super capacitors c69 and c70. the root cause of the system self check failure was contamination of pbm through supercapacitor electrolyte leakage leading to pbm corrosion. a.3 revised as incorrect selection was made on the initial manufacturer device report number 1220648-2024-12813. e.1 revised name prefix/title, first name and last name were entered incorrectly on the initial manufacturer device report number 1220648-2024-12813. e.2 and e.3 revised since initial submission of manufacturer device report number 1220648-2024-12813 to reflect the proper information that belongs with the person being reported in e.1. h.3 device not returned to manufacturer for evaluation should of reflected that the device was returned on the initial manufacturer device report number 1220648-2024-12813. if the device was not evaluated selection on the manufacturer device report number 1220648-2024-12813 should of reflected device evaluation anticipated, but not yet begun.
24925 | 1220648-2024-12812 | 19657129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19657129&pc=OZD | 06/21/2024 | 09/09/2024 |  | 07/02/2024 |  |  | 06/21/2024 | 07/02/2024 |  |  |  |  | 07/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025491413 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 84-year-old male of unknown ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the automated impella controller (aic) displayed purge line had a leak. there was no fluid noted outside the line or the yellow luer lock detected. there was no reported patient harm. the pump was not explanted as a result. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the pump was returned for investigation. when the pump was primed, there was a leak observed from the holes of the purge filter. when the purge filter is exposed to isopropyl alcohol (ipa) solution, similar leaks have been observed in past. therefore, as per the product investigation, the root cause of the purge leak was sidearm filter damage due to ipa interaction. d9 date device returned to manufacturer added g1 manufacturing site name and address was incorrect on manufacturer device report 1220648-2024-12812.
24926 | 1220648-2024-12808 | 19643192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643192&pc=OZD | 06/18/2024 | 08/12/2024 |  | 07/01/2024 |  |  | 06/18/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025436739 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and while supporting the patient the impella 5.5 had pump saturation followed by pump stoppage. the pump was restarted at a lower p-level. the patient was taken to the catheterization lab for exchange of the impella 5.5 to an impella cp, which was successfully implanted. there was no report of adverse effects to the patient who was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump stop and the saturated flow has been completed. pump was returned and during investigation, heavy biomaterial was found in the cannula, outflow cage and all around the impeller blades. data logs were returned and there was a motor current spike noticed corresponding to the rise in the purge pressure. the pressure was stayed stable after that event for about 3 hours and then the pressure started gradually increasing to the point of pump stop. following the first motor current spike saturated flows were observed. this events recorded in the logs confirm that biomaterial ingestion was the root cause of the pump stop failure. therefore, per the product evaluation and data logs review, the root cause of the pump stop issue was biomaterial. added- d9 device return date
24927 | 1220648-2024-12809 | 19643225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643225&pc=OZD | 06/18/2024 | 08/05/2024 |  | 07/01/2024 |  |  | 06/18/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465510 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experience an access site complication. following the percutaneous coronary intervention procedure, bleeding was noted around peel away sheath. an angiogram revealed extravasation at the arteriotomy. the decision was made to gain access to the left side and move impella cp to the left groin. balloon tamponade attempted of the right groin via right radial approach, however, it was unsuccessful. vascular surgeon performed a cutdown to close arteriotomy. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12809. new code was added to component code
24928 | 1220648-2024-12810 | 19643265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643265&pc=OZD | 06/18/2024 | 08/09/2024 |  | 07/01/2024 |  |  | 06/18/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488161 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | VA EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. however, after insertion of the impella cp device, the patient was diagnosed with lower limb ischemia. the sheath was inserted in reverse and blood was sent via extra-corporeal membrane oxygenation support to relieve ischemia. the patient was noted to be in serious condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12810.
24929 | 1220648-2024-12811 | 19643303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643303&pc=OZD | 06/20/2024 | 08/01/2024 |  | 07/01/2024 |  |  | 06/20/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025491653 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the patient had a pocket bleed. the pocket was washed out, and the artery was repaired. two units of packed red blood cells were given. the patient was stable post procedure and remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added explant date h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12811. new code was added to component code.
24930 | 1220648-2024-12814 | 19643469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643469&pc=OZD | 06/23/2024 | 08/21/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  | 06/30/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380702 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old male patient with acute myocardial infarction / cardiogenic shock had an attempted implant with an impella cp for mechanical circulatory support. the complainant reported delivery failed due patient anatomy - there was a stenotic aortic valve. medical staff aborted the delivery, and no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g2 report source: foreign inadvertently was not selected for manufacturer device report.
24931 | 1220648-2024-12815 | 19643526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643526&pc=OZD | 06/23/2024 | 08/30/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025479165 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the female patient experienced hemolysis. act (activated clotting time) was noted to not be in the target range a few hours following implant procedure (150s / antix 0.2). pump position was confirmed to be good via imaging. decision was made to explant the pump after a day of support. however, care was withdrawn. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient data logs was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. b2-selected death and added date of patient death. g1-corrected mfr site name and address. h1-type of reportable event changed to death. h6 impact code 1802. h6 component code 4755 changed to 925.
24932 | 1220648-2024-12816 | 19643591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643591&pc=OZD | 06/24/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  | 05/31/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024354080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed dark brown urine that transitioned to merlot colored urine. labs hemolyzed. the pump was repositioned, and the patient still had hemolysis. the pump was explanted and escalated to an impella 5.5. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed finding that suction alarms occurred. no positioning issue observed. no motor current spike observed. no indications to confirm the cause of the reported issue. the cause of the hemolysis issue cannot be determined since complaint device not returned for investigation and insufficient clinical data provided. h6 component code 4755 changed to 925.
24933 | 1220648-2024-12817 | 19643681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643681&pc=OZD | 06/24/2024 | 08/01/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1443894 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility biomedical department reported the automated impella controller (aic) having a controller failure. the aic was removed from use and the instructions for use (ifu) were followed and a backup console will be used for a patient support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: for the given complaint, it is most likely referring to the unexpected controller shutdown alarm. the previous bootup, the console was undergoing field testing for impella connect and at the conclusion the console was unexpectedly shutdown. most likely, the transport switch was flipped so, since the a/c power was removed, the console immediately shut off. this would not cause an attempted reboot and would trigger a 603 alarm on the next boot up. device analysis: the console performed as expected. the series of events performed in the field could be performed resulting in the same aic logs as seen from the complaint date. root cause: the cause of the unexpected shutdown alarm was most likely a user-initiated shutdown. d9 date device returned to manufacturer added.
24934 | 1220648-2024-12818 | 19643749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643749&pc=OZD | 06/24/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025487701 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 25-year-old female patient of unknown race with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during placement of the impella 5.5 sustained flows were unable to be achieve. fluid boluses were given, and correct positioning was confirmed by transesophageal echocardiogram (tee). flows reached zero several times. the 5.5 was removed and a new 5.5 was placed successfully. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device nor the data logs were returned for evaluation. therefore, without sufficient clinical details, the root cause of the low pump flow could not be determined.
24935 | 1220648-2024-12819 | 19643781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643781&pc=OZD | 06/24/2024 | 09/01/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support post coronary artery bypass graft and aortic valve replacement. implantation of the impella 5.5 was unsuccessful. the impella was unable to advance completely through the artery into left ventricle. the implant was aborted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump was aborted despite multiple attempts. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number. g.1 revised reporting contact fax number from the initial submission of manufacturer device report number in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
24936 | 1220648-2024-12820 | 19643828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643828&pc=PYX | 06/24/2024 | 10/15/2025 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025484867 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 61-year-old female patient post-operative cardio thoracic surgery was implanted with an impella rp device for mechanical circulatory support. it was reported that while on support, low pump flows were observed, and the patient had hemolysis. flows progressively dropped over several days until sustained flows were less than 1.5 lpm on p5 (on p5 due to suction events and hemolysis). blood products were given. the patient was ultimately taken to the operating room for a wash out of clot from the chest and rp removal. the patient was reported to be in critical condition. | The investigation uncovered optical signal issues. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow and hemolysis has been completed. the rp flex pump was returned for evaluation and visually inspected. no cannula kink observed. some biomaterial around purge gap observed. no other abnormalities found. hemolysis test conducted. no hemolysis result reproduced and mih=3.16 field data logs were reviewed finding suction alarms occurred. no positioning issue observed. no motor current spike observed. drop in flow and raise in placement signal observed in the last 4-5 hours of impella support at p-6. same issue with flow drop observed while the hemolysis test conducting. after conducting the hemolysis test, lab data logs reviewed: drop in flow and raise in placement signal observed again while the pump was running for testing. this issue led us to sensor drift failure. optical sensor visually inspected. no biomaterial observed. 5s parameters check and it was within specification limits. electric resistance checked and it was within specification limits. the cause of hemolysis issue most likely was patient condition related as she had a small ventricle and suspect tamponade limited pump flows. the cause of placement signal issue most likely was sensor drift due to component reliability. added- d9 device return date, h6 impact code 4624 g1 corrected mfr site name and address. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
24937 | 1220648-2024-12821 | 19643889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19643889&pc=OZD | 06/24/2024 | 08/14/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  | 08/31/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390826 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Hemorrhagic Stroke | Renal Impairment | [9 empty] | Required Intervention | Other | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 77-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported during support, the patient had a subdural hematoma and a gastrointestinal bleed. medical staff then withheld anticoagulants and later added heparin to the purge when the bleeding subsided from both locations. it is unknown if the patient had the hematoma before impella support, but a clinical representative confirmed medical staff discovered it after the patient had already started impella support. the patient was also on continuous renal replacement therapy (crrt) while on support. crrt was started in the setting of acute kidney injury, with decreased urine output and rising creatinine levels. the patient also pulled on the impella pump during support, which caused a pump stop, and medical staff could not restart the pump. medical staff then explanted the pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop, vascular damage, and stroke has been completed. data logs were reviewed, and several #115 pump stop alarms, #119 alarm were observed. motor current spike to > 30000 ma was found with 115 pump stop alarm. pump was unable to be restarted. as per clinical patient pulled on impella and pump stopped. the cause of the pump stop issue was most likely an open electrical line per log analysis and clinical information. the cause of the vascular damage issue was not determined with limited clinical information and due to unknown relation to impella. the cause of the renal failure issue was not determined since product was not returned and due to insufficient information. the cause of the stroke issue was not determined since product was not returned and due to insufficient information.
24938 | 1220648-2024-12822 | 19644230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644230&pc=OZD | 06/21/2024 | 08/08/2024 |  | 07/01/2024 |  |  | 06/21/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025487265 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 47-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient continued going into ventricular tachycardia. the impella was explanted on day one of support. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the impella cp with smartassist system staes the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
24939 | 1220648-2024-12823 | 19644246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644246&pc=OZD | 06/21/2024 | 08/13/2024 |  | 07/01/2024 |  |  | 06/21/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482017 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that due to the patient¿s vascular anatomy the impella 5.5 was not able to be placed. the impella was not able to advance into the ascending aorta from the right axillary artery. the 5.5 case was aborted and impella cp was placed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issue has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because impella 5.5 was unable to pass through the vasculature and cp was successfully implanted. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12823: d4 model number e4 should have been left blank
24940 | 1220648-2024-12824 | 19644261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644261&pc=OZD | 06/21/2024 | 09/03/2024 |  | 07/01/2024 |  |  | 06/21/2024 | 07/01/2024 |  |  | 11/30/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434318 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that towards the end of the cabg (coronary artery bypass grafting¿) procedure and while attempting to close the chest the patient decompensated. the chest was reopened for possible 5.5 placement. multiple attempts were made to insert the impella however, the turn from graft into ascending aorta was not able to be made. the graft was untunneled and 5.5. was attempted to pass with a straight drop into the aorta. the impella with difficulty passing through the anastomoses, digital manipulation and able to pass pump motor housing. the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | D.9 updated as the pump was returned for investigation. the investigation for the reported delivery issues has been completed. the root cause of the delivery issue was most likely patient condition related. multiple attempts were made to pass the impella. however they could not make the turn from graft into ascending aorta as per the clinical description provided. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12824. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12824 in accordance with updated procedures. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12824 as it is unknown the initial reporter also sent the report to the fda. h.6 code 2993 was entered incorrectly on the initial manufacturer device report number 1220648-2024-12824. a new code has been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12824 in accordance with updated procedures. | H.3 was the device evaluated by the manufacturer? and evaluation summary attached were revised as they should have been checked as yes on the previously submitted manufacturer device report number .
24941 | 1220648-2024-12825 | 19644273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644273&pc=OZD | 06/22/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/22/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503851 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed ischemia in the right lower extremity, which was cold. medical staff decided the patient no longer needed impella support, so they explanted the device. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12825: e4 should have been left blank. g1 reporting contact fax number missing.
24942 | 1220648-2024-12827 | 19644373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644373&pc=OZD | 06/21/2024 | 08/27/2024 |  | 07/01/2024 |  |  | 06/22/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025488622 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 had a failure. the impella was exchanged. the patient is stable. | [5 empty] | The investigation for the pump stop has been completed. no product was returned for evaluation. clinical details noted that the pump was plugged into the console and case start was not prompted. data logs showed that the pump was recognized by the console and able to go through case start without issue. the pump had real time log with motor current and ps of 0 for one hour before pump was unplugged and reconnected back to the automated impella controller (aic). a different pump was then disconnected from the aic again and data logs showed a new pump and purge cassette was connected. the root cause of the pump did not start was most like due to a use issue as the pump was able to complete case start without issue. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.2 selection was made incorrectly on the initial manufacturer device report number 1220648-2024-12827. there was no adverse event. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12827. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12827 in accordance with updated procedures. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12827 as it is unknown the initial reporter also sent the report to the fda. h.6 code 2993 was entered incorrectly on the initial manufacturer device report number 1220648-2024-12827. a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12827 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
24943 | 1220648-2024-12826 | 19644376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644376&pc=OZD | 06/24/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469355 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support pre percutaneous coronary intervention. despite antegrade's bilaterally, and thrombectomy to the right superficial femoral artery (sfa) and inflow through profunda, the doppler signal is faint. the vascular team feel the only option is decannulation, ischemic damage is irreversible at this point. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.). | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12826. new code was added to component code.
24944 | 1220648-2024-12828 | 19644490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644490&pc=PYX | 06/24/2024 | 08/20/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  | 10/31/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024423965 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 55-year-old male with pulmonary embolism was to be implanted with an impella rp device for mechanical circulatory support. it was reported that the impella rp was unable to advance in the pulmonary artery. the rp kept getting kinked in the inferior vena cava (ivc). the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings and cautions: do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition because the device was getting kinked during insertion and unable to advance into the pulmonary artery and the doctor has decided to insert the impella rp flex. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12828 in accordance with updated procedures.
24945 | 1220648-2024-12829 | 19644552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644552&pc=OZD | 06/22/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/22/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025515024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an 83-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a concern for hemolysis and one unit of packed red blood cells was given. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-12829.
24946 | 1220648-2024-12830 | 19644569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644569&pc=OZD | 06/11/2024 | 07/25/2024 |  | 07/01/2024 |  |  | 06/11/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490040 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 50-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed a hematoma at the impella¿s axillary site. the medical staff administered one unit of packed red blood cells and pressure dressing was applied. a thrombus in the right arm, with swelling, was also reported. the medical staff then performed a thrombectomy and explanted the impella cp. the patient¿s outcome remains unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for thrombus and hematoma has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added.
24947 | 1220648-2024-12831 | 19644630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644630&pc=OZD | 06/23/2024 | 08/16/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025481758 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 41-year-old male patient with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported attempted impella 5.5 implant through the left axillary and it was unsuccessful. the endograft suture ring from previous procedure prevented passage of impella, the implant was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. e4 should have been left blank on manufacturer device report 1220648-2024-12831.
24948 | 1220648-2024-12832 | 19644644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644644&pc=OZD | 06/23/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025507390 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that high purge pressure and purge system blocked alarm occurred on the impella cp. tpa (tissue plasminogen activator) protocol was suggested. purge flow normalized so tpa will be on hold at this time. motor current will be monitored closely. there was no harm to the patient. | It was further reported that the patient updates revealed that there was bleeding from all access sites on day 4 of support. the patient was treated with packed red blood cells (prbcs) administered. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. no product was returned for evaluation. the data logs confirmed high purge pressure and purge system blocked alarms. the logs show a gradual rise in purge pressure which started 4 hours into support. the rise continued for about an hour before the alarms were triggered. the high purge pressure remained for 2 hours before quickly returning to normal values. there were no other occurrences for the rest of the procedure. the root cause of the high purge pressure was most likely biomaterial related. the root cause of the access site bleeding - major was most likely patient condition related as bleeding was reported from multiple sites. the instructions for use referenced in the initial report is for impella cp with smartassist for use during cardiogenic shock and high risk cpi. it further states: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b5 additional narrative, h6 clinical code 1888, h6 impact code 4643 h6 med dev prob code 2993 changed to 1491.
24949 | 1220648-2024-12833 | 19644655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644655&pc=OZD | 06/23/2024 | 08/09/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455999 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for a patient admitted in cgs (cardiogenic shock). the pci (percutaneous coronary intervention) was done with impella support and remained on afterwards. the access site developed a hematoma which required a femostop and gave a unit of blood. the pump remained on for support for 49 hours and was then explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding/hematoma has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. d4 model number. e4 should have been left blank.
24950 | 1220648-2024-12834 | 19644676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644676&pc=OZD | 06/23/2024 | 08/16/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486536 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 72-year-old male patient with acute myocardial infarction / cardiogenic shock had an attempted implant with an impella cp for mechanical circulatory support. the complainant reported that via an echocardiogram, medical stuff found the cp in malposition and unsuccessfully tried to reposition the cp. attempts in the catheterization lab were also unsuccessful, so medical staff removed the cp. the patient remains stable. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique. added- h6 impact code 4628, 4627 d4-corrected model number.
24951 | 1220648-2024-12835 | 19644700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644700&pc=OZD | 06/23/2024 | 08/22/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485087 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 19-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support was taken to the operation room for pseudo aneurysm in right groin. the vascular surgeon successfully achieved hemostasis via surgical wound closure including sutures. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the vascular damage has been completed. the patient was taken to the operating room for pseudo aneurysm in the right groin. hemostasis was achieved. no product was returned. the root cause of the vascular damage - peripheral vessel was not determined as no product was returned and limited clinical information was provided. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12835 in accordance with updated procedures. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12835 as it is unknown the initial reporter also sent the report to the fda. h.6 added code 4641 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12835. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12835 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
24952 | 1220648-2024-12836 | 19644706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644706&pc=OZD | 06/23/2024 | 08/20/2024 |  | 07/01/2024 |  |  | 06/23/2024 | 07/01/2024 |  |  | 09/30/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024401640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had oozing at the access site and a femstop was placed and heparin was paused. additionally, the impella had high purge pressure alarms. the purge tubing checked for kinks and purge cassette changed. the doctor decided to infuse tissue plasma activator through the purge and the alarm resolved. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the high purge pressure has been completed. the impella device was not returned for evaluation. field data logs were reviewed and the high purge pressure alarm was confirmed. there was a motor current spike to 807 ma followed by a 4 minute rise in purge pressure until the alarm was triggered. data logs also show the purge cassette being replaced, and within 15 minutes, the purge pressure started trending downward until it normalized. there is potential that this played a role in resolving the complications in addition to tissue plasminogen activator (tpa) infusions. additionally, the change in anticoagulation management to resolve the access site bleed the day before could have contributed to the cause of high purge pressure to begin with. without patient acts at the time, however, this cannot be confirmed. the root cause of the high purge pressure was determined to be biomaterial ingestion. this is indicative of another biomaterial ingestion event). the root cause of the access site bleed could not be determined as not enough clinical information was provided. added- h6 impact code 4629, 4644.
24953 | 1220648-2024-12837 | 19644728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644728&pc=OZD | 06/24/2024 | 08/20/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025480338 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported the patient arrived at the critical care unit bleeding from the insertion site with blood pooling under the femstop, manual pressure was held. the bleeding was unable to be resolved for four hours. angle matching, several suture techniques, femstop, surgisil, and manual pressure were utilized. thrombin and epinephrine were given to help with bleeding. two units of packed red blood cells (prbc) cryoprecipitate and 1 fresh frozen plasma (ffp) were administered. bleeding finally resolved with 80¿90-degree angle and forward pressure. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding details was determined to be an use issue because the repo sheet was inserted deep into the skin causing the bleed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12837. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12837as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12837 in accordance with updated procedures.
24954 | 1220648-2024-12838 | 19644772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644772&pc=OZD | 06/24/2024 | 08/26/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474996 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the implantation was unsuccessful. after multiple attempts to implant the pump through the axillary/subclavian territory, a dissection was noted at the anastomosis.the physician noted the anastomosis appeared to be at a bifurcation with calcium present. the 5.5 was removed and the dissection was repaired. there was also a kink noted on the catheter shaft. the implant was aborted and a new pump was inserted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance. | The investigation for the vascular damage and the delivery issues has been completed. no product was returned for evaluation. clinical details noted that patient had peripheral vascular disease, stenosis and calcified arteriotomy. excessive force was applied in attempt to insert the pump without success and a kink in the catheter just below the motor was seen after the pump was removed. additionally, a dissection of the artery at the graft anastomosis was noted and repaired with vascular surgery. the root cause of the vascular damage - peripheral vessel was due to patient condition as patient's vessels had stenosis and calcification. the root cause of the delivery issues was due to patient's diseased vessel condition. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as information was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12838. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12838 as it is unknown the initial reporter also sent the report to the fda. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12838 was submitted. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12838 in accordance with updated procedures.
24955 | 1220648-2024-12839 | 19644889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644889&pc=OZD | 06/25/2024 | 08/13/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502090 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding and a hematoma. the patient's partial thromboplastin time was over 400. heparin was withheld for several hours. the patient remained on impella cp support. | It was further reported that the patient expired on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values; heparin was stopped to achieve hemostasis as per the clinical description. added-b5 mso review h6 component code 4755 changed to 925.
24956 | 1220648-2024-12840 | 19644943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644943&pc=OZD | 06/23/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  | 10/31/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024412201 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 54-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient developed a localized infection at the insertion site. medical staff drained and cultured the infection. medical staff then took the patient to the operation room for washout of the site and wound vac placement at the cutdown site of the drainage. the patient also complained of some arm pain. echocardiography showed the impella position to be angled toward the mitral apparatus. the patient survived the event and remains stable. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of positioning issues and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. e4 should have been left blank. g1 reporting contact fax number missing.
24957 | 1220648-2024-12841 | 19644966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644966&pc=OZD | 06/24/2024 | 09/06/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  | 12/31/2024 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306405 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed access site bleeding and thrombocytopenia during impella support. medical staff performed a blood transfusion for the patient and later noticed the patient developed hemolysis. medical staff then removed the cp and replaced it with the impella 5.5. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, thrombocytopenia, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12841. d6a and d6b incorrect. e4 should have been left blank. g1 reporting contact fax number missing. g2 report source: foreign inadvertently was not selected. h6 component code revised from the initial submission in accordance with updated procedures.
24958 | 1220648-2024-12842 | 19644980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19644980&pc=OZD | 06/24/2024 | 08/16/2024 |  | 07/01/2024 |  |  | 06/24/2024 | 07/01/2024 |  |  | 11/30/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431420 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 had low purge pressure and leaking at the tubing connection to the sidearm was observed. a decision was made to wait until the following day to explant due to the purge pressure maintaining above 300 and flow appropriate throughout the night. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12842. e4 should have been left blank. h5 was inadvertently selected as no.
24959 | 1220648-2024-12843 | 19645007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645007&pc=OZD | 06/25/2024 | 08/28/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  | 09/30/2024 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023272972 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had unresolved hemolysis. the pump was repositioned; however, the hemolysis did not resolve. the physician thinks the cp is the cause of the hemolysis. the patient remains on support despite the event and is reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. the product and data were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ e.4 should have been left blank in initial manufacturer device report number as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number in accordance with updated procedures.
24960 | 1220648-2024-12844 | 19645059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645059&pc=OZD | 06/25/2024 | 08/06/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  | 09/30/2025 |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025444861 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had limb ischemia. the decision was made to leave the placement sheath in place and place an antegrade and attach the sheath to the pigtail from placement sheath to serve as distal perfusion cannulation (dpc). poor perfusion was observed and the dpc was clotted. the patient was not on anticoagulation. the impella cp was removed. there were still non-palpable pulses to the affected limb after explant. the staff continued to monitor, and patient was reported as stable | [5 empty] | The investigation for the limb ischemia and thrombus has been completed. the investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. added- h6 component code 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
24961 | 1220648-2024-12845 | 19645061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645061&pc=OZD | 06/25/2024 | 08/19/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025471267 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 47-year-old male of unknown ethnicity with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while the patient was walking down the hall, impella in aorta alarm occurred. current motor remained the same. recrossing was attempted at bedside under echocardiogram but unsuccessful. the 5.5 was exchanged. there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely due to patient movement. e4 should have been left blank in manufacturer device report 1220648-2024-12845.
24962 | 1220648-2024-12846 | 19645075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645075&pc=OZD | 06/25/2024 | 08/13/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320433 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | No patient involvement: the complainant reported that the blue plastic purge disc slot on the impella aic (automated impella controller) detached. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of aic - purge disc/flag issues has been completed. no data review is needed for this complaint because it was only reported that the retainer was broken. the purge disc retainer was found to be broken (retainer completely broken off). the root cause of the issue was a broken purge disc retainer. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12846: d4 model number e4 should have been left blank
24963 | 1220648-2024-12847 | 19645158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645158&pc=OZD | 06/11/2024 | 08/06/2024 |  | 07/01/2024 |  |  | 06/11/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 31-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that on day two of support, the pump stopped. the heart was visualized under imaging and a large left ventricle thrombus was present and they were unable to successfully restart the pump. the impella cp was explanted. the patient was reported as stable. | [5 empty] | The investigation for the pump stop and thrombus has been completed. no product was returned for evaluation. data logs were reviewed and a pump stop was observed with an alarm #119. pump was attempted to be restarted multiple times without success. clinical details noted that a large lv thrombus was observed when imaging the heart. the root cause of the pump stop was most likely due to an open electrical line. the root cause of the thrombus could not be determined without the product or additional clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk. percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
24964 | 1220648-2024-12848 | 19645166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645166&pc=OZD | 06/07/2024 | 07/26/2024 |  | 07/01/2024 |  |  | 06/09/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025452692 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support went into ventricular fibrillation and required shocking to return to normal sinus rhythm. the arrhythmia occurred while the patient was having a coughing fit while being turned. echocardiogram verified the correct pump position. two days later, it was reported that the patient was shocked about 20 times from last night into the morning for ventricular tachycardia storm. the patient required added support so the patient was intubated and placed on extracorporeal membrane oxygenation. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for the impella cp with smartassist system lists ventricular arrythmia has an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." added-h6 component code 925. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24965 | 1220648-2024-12849 | 19645454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645454&pc=OZD | 06/11/2024 | 07/25/2024 |  | 07/01/2024 |  |  | 06/11/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that they attempted an impella 5.5 implant. however, the implant was aborted due to the patient's anatomy. the patient had a very torturous axillary artery and the device was having trouble passing through tortuosity of the axillary artery leading into the aorta. the implant was aborted. | [5 empty] | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition due to tortuous anatomy of the patient. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12849: e4 should have been left blank. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
24966 | 1220648-2024-12850 | 19645597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645597&pc=OZD | 06/25/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495526 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention (pci). the impella cp had low flows but rant for the entire pci and was explanted. the team also noted that at the time of the impella cp explant, the perclose failed to achieve hemostasis. the team made the choice to give protamine to reverse the anticoagulation, hold for 50 minutes, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow and access site bleed has been completed. the cp pump was returned for evaluation and was visually inspected. no biomaterial was observed. no broken blades founds. no cannula kink observed. pump connected to aic in investigation lab to reproduce reported low pump flow. pump run for over 15 minutes at p-9. no low pump flow reproduced. flow was within specification limit at p-9. data log were reviewed finding suction alarms occurred a few minutes after impella initiated. no positioning alarms occurred. no clear indication confirm the cause of the reported issue. the cause of the low pump flow issue most likely was patient condition related. the cause of the access site bleeding most likely was anticoagulation management related due to high act. added- d9 device return date, h6 impact code 4627 added- d9 device return date, h6 impact code 4627.
24967 | 1220648-2024-12851 | 19645830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645830&pc=OZD | 06/25/2024 | 08/15/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474105 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to left ventricle assist device. during support, had an event of nausea, slight chest pain, cuff pressures 180¿-200¿s systolic, correlating with the aortic (ao) placement signal on automated impella controller (aic). the impella was briefly turned down to p-7, then back up to p-9 where the patient was supported. motor current and purge unchanged. ao placement signal was not correlating with cuff pressures. nipride was started for the patient and blood pressure was down to 115¿s, and patients symptoms resolved, but aic still read high. the physician did not want to replace the impella 5.5 and wanted to monitor vitals and labs. echo also obtained and pump was in good position. hours later, the ao placement signal was back to normal and correlating with bp cuffs. it was further reported that the impella sterile sleeve ripped. tegaderm dressing was applied around the impella catheter. the patient is stable. | [5 empty] | The investigation of optical signal issue and product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24968 | 1220648-2024-12852 | 19645943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19645943&pc=OZD | 06/25/2024 | 08/30/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463851 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 38-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that, after five minutes of impella support, waveforms appeared to indicate pump saturation, and a purge pressure alarm sounded shortly after. medical staff then decided to replace the pump. no patient harm has been reported. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation has been completed. the impella 5.5 s2 pump was not returned for investigation. data logs were reviewed and saturated flow was observed after purge pressure rise trend and alarms. no motor current spike but motor current variation found without p-level change along with gradual purge pressure rise of 6 minutes. the cause of the high purge pressure issue was most likely biomaterial.
24969 | 1220648-2024-12853 | 19646028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19646028&pc=OZD | 06/25/2024 | 08/12/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438166 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support as a bridge to ventricle assist device. the patient was known to have a severely tortuous anatomy. during implantation, the impella would not cross the axillary. the implant was aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella would not cross the axillary as per the clinical description provided.
24970 | 1220648-2024-12854 | 19646064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19646064&pc=OZD | 06/12/2024 | 08/05/2024 |  | 07/01/2024 |  |  | 06/12/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504181 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male patient in st-elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after the pump was implanted, low pump flows were noted. volume was given and the pump was repositioned. however, the pump flows did not improve. the impella was removed and replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs confirmed the failure mode and clinical narrative. logs showed after pump started for about 5 minutes, low flow occurred with motor current (mc) spiked that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
24971 | 1220648-2024-12855 | 19646091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19646091&pc=OZD | 06/25/2024 | 08/05/2024 |  | 07/01/2024 |  |  | 06/25/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458090 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 40-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after insertion of the repositioning sheath and an act (activated clotting time) of over 400, the impella cp site was oozing. overnight the patient had hemoglobin (hgb) drop of 5 points. two units of rbc (red blood cells) were given and hgb returned to above 10. oozing continuing after multiple attempts to fix, the physician decided to remove the impella. the cp was successfully removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely anticoagulation management since the act were above the limit range. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12855. new code was added to component code
24972 | 1220648-2024-12856 | 19648574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19648574&pc=OZD | 06/12/2024 | 09/30/2025 |  | 07/01/2024 |  |  | 06/12/2024 | 07/01/2024 |  |  |  |  | 07/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463954 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella 5.5 encountered low flows alarm. a clot was seen on the tip of the impella 5.5 after having suctions alarms and intermittent placement signal not reliable alarms, regardless of repositioning. the impella was removed and replaced successfully. | [5 empty] | The investigation of ow pump flow has been completed. the impella device was not returned for evaluation. data logs confirmed the failure mode & clinical narrative. logs showed after pump started and several suction alarms observed at the end of case with low flows. per clinical doctor visualized clot on inlet of pump with suction alarms. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the product was not returned. data logs indicate intermittent placement signal not reliable alarms with placement signal spiking to 3000. optical 5s parameters trend indicates likely air gap failure. this failure resolved soon later. the root cause of the optical signal failure was most likely an air gap from between the console and the patient cable plug. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-12856. g1 reporting contact email updated.
24973 | 1220648-2024-12801 | 19640683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640683&pc=PYX | 06/10/2024 | 08/05/2024 |  | 06/30/2024 |  |  | 06/10/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025442050 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that during a bed bath, the patient had the impella rp flex pulled out 6 centimeters. crossing the pulmonary vessel was unsuccessful. the impella rp flex was replaced. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on clinical description, the root cause of positioning issue was most likely patient movement.
24974 | 1220648-2024-12802 | 19641007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19641007&pc=OZD | 06/10/2024 | 08/12/2024 |  | 06/30/2024 |  |  | 06/10/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484025 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 71-year-old male patient post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that when inserting the impella 5.5, the cannula was unable to pass through the patient¿s aortic valve. after multiple troubleshooting attempts, the medical staff aborted the insertion. the medical staff instead put the patient on centrally cannulated venous-arterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition because they were unable to pass the cannula through the aortic valve after multiple troubleshooting attempts.
24975 | 1220648-2024-12803 | 19641016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19641016&pc=OZD | 06/10/2024 | 10/21/2024 |  | 06/30/2024 |  |  | 06/10/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490789 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Respiratory Failure | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 69-year-old male patient was implanted with an impella cp due to the patient having ventricular tachycardia ablation. after the implant, the flows decreased, and the patient became hypotensive. an echocardiogram was performed and revealed pericardial effusion. the impella was slowly weaned, and the patient tolerated it well. after the explant, the effusion was able to be tapped. the patient was reported as stable. | It was further reported by medical safety team via abiomed¿s long-term outcome and quality indicator impella registry (loqi) that the patient experienced acute hypoxic respiratory failure with an unknown relationship to the impella cp device used on this patient. the treatment for the respiratory failure is unknown. |  |  |  |  | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The provided logs verified the decreased placement signal pulsatility and drop in flow to 3 l/min at p-9 in the field. the product was returned and evaluated. no abnormalities were found. the low pump flow could not be reproduced in the lab. therefore, the root cause of the low pump flow was most likely patient condition related. vascular damage - peripheral vessel. the introducer was not returned, but the pump with repositioning sheath was returned and evaluated. no abnormalities were found with the repositioning sheath. the root cause of the pericardial effusion was not determined as insufficient clinical evidence was provided. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings and cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | Added details for the initial mfr report -b5 additional adverse event details, h6 clinical code 2484, h6 impact code 4648.
24976 | 1220648-2024-12804 | 19641021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19641021&pc=PYX | 06/10/2024 | 08/07/2024 |  | 06/30/2024 |  |  | 06/10/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male patient with cardiomyopathy was implanted with an impella rp flex for mechanical circulatory support. it was reported that the automated impella controller (aic) had a brief unexpected controller shut down. the aic was exchanged and there were no further issues. there was no patient harm reported. | [5 empty] | The investigation for the console (aic) shutdown or restart issue has been completed. data logs confirm that there was unexpected controller shutdown of the console on the day of the complaint. the console was visually inspected in the lab with no revelation of an anomaly to the internal circuitry of the console. the console was run at maximum p-level for more than 4 hours in unsuccessful attempt to recreate the complaint failure. the root cause for the unexpected controller shutdown could not be determined due to the failure not being able to be reproduced. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
24977 | 1220648-2024-12805 | 19641025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19641025&pc=OZD | 06/10/2024 | 07/23/2024 |  | 06/30/2024 |  |  | 06/10/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient had left superficial femoral artery (sfa) limb ischemia. a right leg bypass to left leg¿s sfa antegrade sheath for perfusion was performed. additionally, the patient had gastrointestinal bleeding and anticoagulation was withheld. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for limb ischemia and vascular damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added 3038, 925.
24978 | 1220648-2024-12806 | 19641039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19641039&pc=OZD | 06/08/2024 | 08/20/2024 |  | 06/30/2024 |  |  | 06/10/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493726 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 84-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient's heel hit the purge line resulting in damage to the purge filter. the purge flow became blocked, and the pump subsequently stopped. the patient went into cardiac arrest, and despite cardiopulmonary resuscitation, a return of spontaneous circulation was not achieved, and the patient passed away. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Product was not returned for review. data logs were reviewed, and the purge leak was confirmed. initially, the purge pressure drops to zero and purge flow rises along with a purge system open alarm, indicating a break in the purge line. then, the purge pressure rises over a 4 second span until the high purge pressure alarm triggers. this cycle between hpp and 0 mmhg repeats several times. this indicates that the team was troubleshooting by attempting to patch the leak but was not successful. the root cause of the purge leak was determined to be damage to the reservoir to filter joint during patient movement. during the 20-minute span that the team was attempting to troubleshoot the purge leak, there was very little or no purge reaching the motor. with no purge fluid in the purge gap, blood was likely able to ingress in, causing high purge pressure and a motor current trend upward. alarm 13 (high motor current) was ultimately triggered, as seen in the data logs. the pump stop was determined to be due to damage to the reservoir to filter joint during patient movement.
24979 | 1220648-2024-12807 | 19641052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19641052&pc=OZD | 04/09/2024 | 08/05/2024 |  | 06/30/2024 |  |  | 06/11/2024 | 06/30/2024 |  |  |  |  | 06/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461241 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Unspecified Kidney or Urinary Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 26-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported a hospital peer review and/or morbidity and mortality conference determined that, for this patient, the local team believed the impella caused a long-term need for dialysis. the complainant reported that during impella support, positioning and suction alarms did not occur as they thought they should. the patient survived the event and has been reported to have had a successful heart transplant post-impella placement. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of renal failure, positioning issues, and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
24980 | 1220648-2024-12794 | 19640583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640583&pc=OZD | 06/08/2024 | 08/01/2024 |  | 06/29/2024 |  |  | 06/08/2024 | 06/29/2024 |  |  | 09/30/2025 |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398810 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention (pci). upon implant of the pump, there were issues with the flow. however, the physician reported the patient did not require higher flow for support. in addition, the patient experienced an access site bleed which was treated by an injection of epinephrine which resolved the bleed issue. the pci was performed and the patient was successfully weaned from the impella cp. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The pump was started on auto with flows ~2.5 l/min. the motor current appeared steep and lvd pressure was very negative, but correct position was verified on fluoro. patient only required minimal support, so lower-than-expected flow was not a concern. low flow while on p-9 support was confirmed on the provided logs. the product was returned and evaluated. a cannula kink was found. low flow was not reproduced in the lab. the root cause of the low pump flow was most likely a cannula kink. upon admission to icu, the patient had consistent bleeding at rfa groin site. the team administered subcutaneous lidocaine with epinephrine, which stopped the bleeding. the patient had known cad and dm. the repositioning sheath was placed, and 4x4 gauze and forward suturing was utilized. there was patient movement during support, and no patient conditions noted. logs were not applicable for this failure mode. the product (without the introducer) was returned and evaluated. no abnormalities were found with the repositioning sheath. therefore, the root cause of the access site bleeding was most likely patient movement.
24981 | 1220648-2024-12795 | 19640594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640594&pc=OZD | 06/02/2024 | 07/31/2024 |  | 06/29/2024 |  |  | 06/02/2024 | 06/29/2024 |  |  |  |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465362 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Asystole | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. complications were encountered during impella cp support. upon initial insertion of the device, the impella hit the ventricle wall and interfered with the left bundle branch. as the right bundle branch had a pre-existing block, the patient went into asystole for twenty seconds. compressions were performed until the patient's native conduction returned. the following day, suction alarms suddenly appeared and the impella flow dropped to 0.0. ventriculography noted a thrombus in the left ventricle that had not been previously viewed. the patient platelet levels started to drop rapidly and platelets were given to the patient. the decision was ultimately made to explant the pump. upon removal, a thrombus was observed occluding the inlet of the pump. the patient remained on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation vf/vt/af/at, low pump flow, and hemolysis has been completed. returned complaint cp pump visually inspected. no cannula kink observed. no broken blades found. a big chunk of biomaterial observed in the inlet cage and pigtail. no other abnormalities. the cause of the low pump flow issue was biomaterial. as the necessary information was not provided, the root cause of the vt/vf and hemolysis was not determined.
24982 | 1220648-2024-12796 | 19640605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640605&pc=OZD | 06/09/2024 | 08/05/2024 |  | 06/29/2024 |  |  | 06/09/2024 | 06/29/2024 |  |  |  |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479269 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 62 year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient was ongoing support for bridging to transplant. the patient developed an access site injury, a hematoma, which prompted surgical exploration. the patient remained and still remains on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
24983 | 1220648-2024-12797 | 19640614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640614&pc=OZD | 06/09/2024 | 08/20/2024 |  | 06/29/2024 |  |  | 06/09/2024 | 06/29/2024 |  |  |  |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025464127 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the doctor reported bleeding from the patient¿s impella access side. the patient¿s hemoglobin dropped from 12 to 7 and the patient needed blood bags, volume was running and also catecholamines. the patient¿s activated clotting time was still high, 195 seconds, due to low preload and suctions have resulted and low mean arterial pressure. the patient expired on support. | [5 empty] | A.5 ethnicity and race are unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. no product was returned for evaluation. clinical details noted that patient's act at time of bleeding was 195. the root cause of the access site bleeding - major was most likely due to anticoagulation management as patient's act values were high at time of bleeding. added- component code 925.
24984 | 1220648-2024-12798 | 19640621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640621&pc=OZD | 06/09/2024 | 07/22/2024 |  | 06/29/2024 |  |  | 06/09/2024 | 06/29/2024 |  |  |  |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460810 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 82-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp was implanted via the femoral artery to allow for percutaneous intervention. the pump successfully supported the patient for stenting. when the pump was explanted, the medical staff found a bleed at the access site due to femoral vascular damage. the medical staff treated the site with balloon angioplasty and covered the stenting. the medical staff also administered blood products to the patient. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. while finishing a pci, an arterial bleed was noted after the repositioning sheath's angle of insertion was flattened a couple times and an right common femoral artery (rcfa) tear was confirmed. the impella cp was removed, so manual pressure could be applied. the tear was treated with a cover stent placed, and hemostasis was achieved with manual compression afterwards. logs were not applicable. the root cause of the vascular damage was most likely use related.
24985 | 1220648-2024-12799 | 19640640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640640&pc=OZD | 06/10/2024 | 03/25/2025 |  | 06/29/2024 |  |  | 06/10/2024 | 06/29/2024 |  |  |  |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186693 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 53-year-old male patient with electrophysiology/ablation was implanted with an impella cp for mechanical circulatory support. it was reported that the associated automated impella controller (aic) did not recognize the purge cassette during preparation. the aic was rebooted. after the reboot, a failure alarm, "change console" occurred. therefore, the aic was exchanged and the procedure was successfully completed. no patient harm was reported. | [5 empty] | A.5 ethnicity and race are unknown. the console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the data review of the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. the root cause of this issue was a broken purge disc flag. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-12799 was submitted in accordance with updated procedures. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-12799 was submitted in accordance with updated procedures.
24986 | 1220648-2024-12800 | 19640679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19640679&pc=OZD | 06/10/2024 | 04/01/2025 |  | 06/29/2024 |  |  | 06/10/2024 | 06/29/2024 |  |  |  |  | 06/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1751714 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during initiation of tissue plasminogen activator (tpa) for elevated purge pressures, the new purge cassette prompted "cassette failure" notification that would not clear. the cassette was removed and the previous purge cassette was replaced with much difficulty sliding the purge disc into position. the decision was made to perform an automated impella controller (aic) to aic transfer to ensure it was no the issue. after the aic exchange, a new purge cassette and bag change were performed to tpa 2 milligrams per 50 milliliters and that cassette has no issues being placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) -other has been completed. the console operated to specification and no problem was found during testing. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-12800 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-12800 in accordance with updated procedures.
24987 | 1220648-2024-12753 | 19632194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632194&pc=OZD | 05/07/2024 | 07/30/2024 |  | 06/28/2024 |  |  | 06/12/2024 | 06/28/2024 |  |  | 09/30/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444594 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured recurring non-sustained ventricular tachycardia with a "definite" relationship to the impella device used: ¿the patient developed recurring short runs of non-sustained ventricular tachycardia including frequent premature ventricular complexes.¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. added- h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
24988 | 1220648-2024-12754 | 19632357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632357&pc=OZD | 11/24/2023 | 08/02/2024 |  | 06/28/2024 |  |  | 06/13/2024 | 06/28/2024 |  |  | 09/30/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401577 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening svt with a "possible" relationship to the impella device used: ¿patient has history of svt. hrs 130s-140s. patient received amiodarone bolus without conversion. patient was cardioverted (b)(6) and hr was in nsr. patient entered 140s-150s again and was cardioverted again on (b)(6) 2023. patient was monitored and treated with po amiodarone until death". it was reported that the patient/event has not recovered/not resolved. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening atrial fibrillation with a "possible" relationship to the impella device used: ¿patient has history of a.fib. patient that night entered a.fib with hr 140s. two doses of amiodarone given and drip started. patient cardioverted (b)(6) 2023 and received 2 unit rbc since hypotensive. monitored for rhythm until death.¿ it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation of vt/vf has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. per medical safety officer review use of the device cannot conclusively be associated with the outcome. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12754. h5 was inadvertently reported no. h.6 code 4755 was reported incorrectly. new code was added to component code. | The device was not returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
24989 | 1220648-2024-12755 | 19632445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632445&pc=OZD | 02/25/2024 | 08/06/2024 |  | 06/28/2024 |  |  | 06/13/2024 | 06/28/2024 |  |  | 10/31/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used: ¿pt with baseline creatinine of 1.64 mg/dl on (b)(6) 2024 prior to impella placement. creatinine rose after impella placed with peak value of 3.85 mg/dl on (b)(6) 2024. nephrology consulted, pt on crrt from (b)(6) 2024- (b)(6) 2024 and intermittent hemodialysis until (b)(6) 2024. last round of hemodialysis occurred on (b)(6) 2024. creatinine continued to rise after hemodialysis was discontinued. on day of discharge creatinine was 5.71 mg/dl. per discharge summary ¿aki: he reports robust urine output. dialysis access has been removed for his icd placement. appreciate nephrology consultation. they do not perceive any future dialysis and have cleared him for discharge.¿. it was reported that the patient/event has not recovered/not resolved.we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿on (b)(6) 2024 pt with evidence of hemolysis. ldh on (b)(6) 2024 at 16:42 was 5426 u/l and urinalysis on (b)(6) 2024 at 14:01 showed color=brown, clarity=extra turbid, blood=3+, rbc >100. impella lowered from p8 to p7. (b)(6) 2024 repeat ldh was 1988 and it continued to downtrend from that point on.¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the renal failure and hemolysis has been completed. the pump was not returned for investigation. data logs show no signs of motor current spikes, positioning issues, or placement signal trends. according to the clinical report, hemolysis and renal failure were reported during impella support. the cause of the renal failure issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ . section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
24990 | 1220648-2024-12756 | 19632581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632581&pc=OZD | 06/14/2024 | 08/06/2024 |  | 06/28/2024 |  |  | 06/14/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1636031 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing for 3 weeks when the aic (automated impella controller) had to be replaced due to a yellow console alarm. no harm or injury and the support continues on a 2nd console. | [5 empty] | The investigation for the console (aic) yellow alarm has been completed. data logs were reviewed which confirmed the reported failure. data logs showed that an "controller error (opm comm crc invalid) [optical pump connected] ¿ alarm,¿ event #1045 occurred during use, which resolved by itself in 2 minutes 13 seconds. real-time (rt) logs confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the aic was returned for evaluation was tested finding an optical bench fw communication protocol anomaly. the root cause of the console issue was due to the controller error and loss of placement signal is due to which is likely due to misalignment of data communication packets received by epc. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
24991 | 1220648-2024-12757 | 19632601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632601&pc=OZD | 06/17/2024 | 08/06/2024 |  | 06/28/2024 |  |  | 06/17/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485814 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an unsuccessful attempt to insert an impella 5.5 pump in a patient with cardiomyopathy for mechanical circulatory support. the physician was unable to advance the impella 5.5 out of the graft despite multiple attempts. after pulling the pump across the hemostasis valve, it was further reported there was visual evidence of damage to the catheter where it joined to the motor housing. additionally, the catheter appeared crimped or creased and a small nick was evident. this pump was discarded and new impella 5.5 device used for insertion. there was no report of harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since successful insertion with other pump. added-d6a/d6b.
24992 | 1220648-2024-12758 | 19632731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632731&pc=OZD | 06/18/2024 | 07/24/2024 |  | 06/28/2024 |  |  | 06/18/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450096 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The access site bleeding prompted the team to additionally infuse 1 unit of red blood cells as treatment. | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient admitted in with angina and was placed ont he cp for the mvd cad (mitral valve disease , coronary artery disease) and planned cabg (coronary artery bypass grafting) . the pump support had an access site bleed that was treated by femostop and manual pressure. the cp was explanted and upgraded to the 5.5 pump.. no lasting harm or injury or extensive intervention was needed. |  |  |  |  | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report: b5 additional treatment information was not included in the narrative h.6 code 4755 was reported incorrectly. new codes were added to health effect impact code and component code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
24993 | 1220648-2024-12759 | 19632766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632766&pc=PYX | 06/18/2024 | 10/28/2025 |  | 06/28/2024 |  |  | 06/18/2024 | 06/28/2024 |  |  | 07/31/2024 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2023281650 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 69-year-old male patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the complainant reported that shortly after the insertion of the rp, the placement signal became unreliable and disappeared. medical staff believed this to be likely positional, as the inlet was angled into the right atrium and may be occluding the pressure differential sensor. the flow and impella were working correctly. the patient remains on impella support. | [5 empty] | The investigation of positioning issues has been completed. the pump was not returned for investigation. upon reviewing the data logs, no problems were found with the pump, and it is apparent that the console was the potential cause of the issue. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
24994 | 1220648-2024-12761 | 19632832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632832&pc=OZD | 06/19/2024 | 08/05/2024 |  | 06/28/2024 |  |  | 06/19/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486821 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 85-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that during insertion, the 5.5 would advance beyond the patient¿s graft anastomosis. medical staff made several unsuccessful attempts before aborting the procedure. medical staff then implanted an impella cp instead. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella 5.5 would not advance beyond the graft anastomosis as per the clinical description.
24995 | 1220648-2024-12760 | 19632841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632841&pc=OZD | 06/07/2024 | 08/16/2024 |  | 06/28/2024 |  |  | 06/07/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482086 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old female patient in non-st elevated myocardial infarction was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during insertion there was significant difficulty noted when making the turn to the ascending aorta. there was bending and buckling of the impella. the patient's anatomy was very challenging. the device would not advance, and the implant was aborted. a new impella 5.5 was prepped and successfully implanted. | [5 empty] | The investigation delivery issues has been completed. the device was received for investigation. a kink was found on the catheter below the motor housing and cannula under visual inspection. no other issues were found on the rest of the pump. the root cause of delivery issue was most likely due to patient condition. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿¿
24996 | 1220648-2024-12762 | 19632890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632890&pc=OZD | 06/20/2024 | 08/26/2024 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025490983 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that prior to patient implant, while priming impella 5.5 device it was plugged into console and impella catheter defective alarm occurred. another console was used, and it had the same alarm. that impella catheter was removed from the table and a new 5.5 device opened with success in powering up, priming, and was to be inserted for patient support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D.9 updated as the pump was returned for investigation. the investigation for the pump not detected issue has been completed. the product was returned for investigation. when connected to the automated impella controller, the pump went to case start window and no "impella catheter defective" or related alarms were observed. data logs were investigated and the "impella defective alarm" was confirmed. per the product investigation and data log review, the root cause of the pump not detected issue was not determined since the issue was unable to be reproduced in the lab and the pump performed to specification. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-12762 in accordance with updated procedures.
24997 | 1220648-2024-12763 | 19632969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19632969&pc=OZD | 05/31/2024 | 08/26/2024 |  | 06/28/2024 |  |  | 06/07/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2025492564 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device cp for mechanical circulatory support. it was reported that after removal of the impella cp, thromboembolism occurred in the iliac artery, requiring additional treatment. the causal relationship with the impella is unknown. the patient is stable. additional information was received four days later. after the removal of the impella cp, embolism of the mesenteric and iliac artery was observed. recanalization was confirmed by stent placement and endovascular thrombectomy. the patient is currently recovering without any symptoms of the embolism in the intestines or lower limbs. the blood vessel in which the stent was placed could not be identified. | [5 empty] | The investigation for the thrombus/embolism has been completed. after removal, thromboembolism occurred in the iliac artery, requiring additional treatment. no product was returned. the root cause of the embolism was not determined as no product was returned and insufficient clinical information was provided. as the necessary information was not provided, the root cause of the thrombus was not determined. h.6 added code 925 since the initial submission of manufacturer device report number 1220648-2024-12763 in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
24998 | 1220648-2024-12764 | 19633070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633070&pc=OZD | 06/13/2024 | 08/30/2024 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC |  |  | US |  | 2025486913 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a percutaneous coronary intervention (pci) with intra-aortic balloon pump (iabp) support was attempted but the patient deteriorated and was switched to impella cp support. extra-corporeal membrane oxygenation (ecmo) was also added but the patient¿s condition did not improve. an aortic dissection was discovered, a decision was made to make the patient a do not resuscitate (dnr). the patient expired. the attending physician stated that the impella did not cause or contributed to the dissection. | [5 empty] | The investigation for the vascular damage has been completed. clinical details states that the aortic dissection discovered during impella cp case. no product was returned. the root cause of the vascular damage - great vessel was not determined as no product was returned and limited clinical information was provided. added- h6 impact code 4642 d4-corrected model number. h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿.
24999 | 1220648-2024-12765 | 19633120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633120&pc=OZD | 06/07/2024 | 08/16/2024 |  | 06/28/2024 |  |  | 06/07/2024 | 06/28/2024 |  |  | 08/31/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386040 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old female patient of unknown ethnicity and race post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the pump would not cross from the innominate to the aorta due to the sharp angulation of the artery. the pump was removed, and a dissection was appreciated in the innominate artery. the case was aborted, and decision was made to use intra-aortic balloon pump. the patient was reported as stable. | [5 empty] | The investigation for vascular damage - peripheral vessel and delivery issues has been completed. no product was returned for evaluation. clinical details noted that tortuosity was present in patient's vasculature and the pump was unable to pass through. the root cause of the vascular damage was most likely due to patient's tortuous anatomy. the root cause of the delivery issues was most likely due to patient's tortuous anatomy. | A.5 ethnicity and race are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25000 | 1220648-2024-12766 | 19633127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633127&pc=OZD | 06/20/2024 | 08/01/2024 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp peel away sheath was leaking from the bottom portion of the peel away sheath where the impella catheter met the hemostatic valve. it was noticed that it was only leaking when both the impella and companion sheath were in place. the patient is stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12766. new code was added to component code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25001 | 1220648-2024-12767 | 19633291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633291&pc=OZD | 06/17/2024 | 08/13/2024 |  | 06/28/2024 |  |  | 06/17/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support the severe bleeding of the access site was observed. the impella cp was explanted and surgical closure of the access site was required. escalation to an impella 5.5 was noted, and the patient's outcome was noted as stable. | [5 empty] | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. e.4 should have been left blank on the initial submitted manufacturer device report number 1220648-2024-12767 as it is unknown if the initial reporter also sent the report to the fda. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial submitted manufacturer device report number 1220648-2024-12767. g.2 foreign was selected incorrectly on the previously submitted manufacturer device report number 1220648-2024-12767. h.6 codes 4627 was reported incorrectly on the initial submitted manufacturer device report number 1220648-2024-12767. a new code was added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25002 | 1220648-2024-12768 | 19633323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633323&pc=OZD | 06/07/2024 | 12/04/2025 |  | 06/28/2024 |  |  | 06/07/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1717562 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and was being transferred. during transport, the patient was requiring increasing numbers of drugs to support blood pressure. when the impella was switched to the new automated impella controller (aic), the patient arrested, coded for 30 minutes and expired. the relationship between the cause of death and switching of the aic's is unknown. | [5 empty] | H.5 and h.8 revised as previously entered incorrectly. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation has been complete. ltlog and rtlog data do not show any anomalies during support. there were no alarms and no issues that would suggest equipment failure. additional ltlog and rtlog data did not provide sufficient evidence to suspect the console as the cause of alarms. console was tested, and found to be operating within specification. the cause of positioning alarms on was not determined, but could have been related to patient condition. d2 common device name updated. g1 reporting contact email updated.
25003 | 1220648-2024-12769 | 19633347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633347&pc=OZD | 06/17/2024 | 08/01/2024 |  | 06/28/2024 |  |  | 06/17/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170631 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a purge cassette pressure disk holder that was broken. the aic was removed from service to be returned. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The automated impella controller was received for investigation. data analysis was not applicable. visual inspection showed the purge disc retainer was found to be broken. the root cause of the issue was a broken purge disk retainer. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12769: d4 model number. d9 date device returned to manufacturer incorrect. e4 should have been left blank. g1 reporting contact fax number missing.
25004 | 1220648-2024-12770 | 19633360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633360&pc=OZD | 06/17/2024 | 09/30/2025 |  | 06/28/2024 |  |  | 06/17/2024 | 06/28/2024 |  |  | 02/28/2026 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that a placement signal not reliable alarm occurred during the impella cp procedure. the patient remained stable. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the impella cp procedure the impella rp was inadvertently opened, and 23 fr sheath was used for the right femoral artery access. impella cp was advanced to the left ventricle and support started. the sheath was occlusive. the decision was made to attempt cp wean. the left femoral artery access was obtained, and advancement of wire was attempted from contralateral side but was unsuccessful. the impella cp was explanted. the patient was stable. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal. the pump was not returned. there are no other details related to the placement signal not reliable (psnr). data logs were not provided. the root cause of the optical signal issue was not determined since logs/product were not returned and insufficient clinical evidence was provided. an rp pump was inadvertently opened and a 23fr introducer was used for rfa access. the desired pump, impella cp, was successfully advanced through the 23fr sheath to the lv pre-pci. due to concern for limb ischemia with 23fr sheath, a groin angio was performed, and the sheath was found to be occlusive. the pump was explanted. the root cause of the delivery issue was most likely use related because the incorrect introducer was used for femoral access with an impella cp. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1. product problem updated. b5. updated to include placement signal issue description. h6. updated to include placement signal issue coding. corrections that were made from the initial manufacturer device report number 1220648-2024-12770. b7. other relevant history, including preexisting medical conditions updated. d.10. concomitant medical products and therapy dates updated. e4. should have been left blank in the initial report. g1. reporting contact email and reporting contact fax number updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.).
25005 | 1220648-2024-12771 | 19633440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633440&pc=OZD | 06/07/2024 | 07/31/2024 |  | 06/28/2024 |  |  | 06/07/2024 | 06/28/2024 |  |  | 11/30/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430991 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support as escalation of care. during the insertion of impella 5.5, the patient innominate artery vessel was very small and the impella was unable to pass. the impella was removed and attempted two more times without success. a decision was made to attempt with 24 french peel-away cook telescoping into graft but unable to pass sheath pass narrowing. the impella 5.5 procedure was aborted, and the patient will remain on impella cp support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, innominate artery vessel was very small and impella was unable to pass through as per the clinical description.
25006 | 1220648-2024-12772 | 19633441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633441&pc=OZD | 06/06/2024 | 09/06/2024 |  | 06/28/2024 |  |  | 06/19/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP- JP |  |  | ABIOMED, INC. |  |  | US |  | 2025496919 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 96 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support as escalation of care from an intra-aortic balloon pump. the pump was placed to support the patient after the surgical closure of the ventricular septal defect. the patient expired after 3 days but there was no malfunction or harm/injury/adverse event attributed to the use of the impella. days later, the team was updated regarding the cause of death and patient adverse event. the abiomed japan team was notified that the patient expired from a ventricular-septal wall perforation. the impella use may have caused/contributed to the outcome. | [5 empty] | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12772. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system for use during: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smart assist systemfor use during cardiogenic shock: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25007 | 1220648-2024-12773 | 19633546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633546&pc=OZD | 06/19/2024 | 07/29/2024 |  | 06/28/2024 |  |  | 06/19/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123077 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was not displaying a placement signal. the aic was exchanged. patient demographics are not available, however it was reported there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for automated impella controller (aic) - loss of opm data has been completed. the impella device was received for investigation. the reported issue of ¿no placement signal¿ (placement signal not reliable condition) was caused by weak optical signal due to contamination and damage of optical pump connector in the console. the root cause of the aic issue was a defective blue receptacle. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12773: d4 model number, e4 should have been blank, g1 reporting contact fax number missing, h.6 code 4629 was reported incorrectly.
25008 | 1220648-2024-12774 | 19633576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633576&pc=OZD | 06/19/2024 | 08/13/2024 |  | 06/28/2024 |  |  | 06/19/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474777 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after impella 5.5 was inserted and flowing for approximately 30 minutes, the patient began to decompensate. flows and waveforms were lost on the aic (automated impella controller). position was checked on tee (transesophageal echocardiography) and a large amount of air was noted in the lv (left ventricle) requiring emergently going back on bypass to allow the patient to recover. the bypass then slowly was weaned and impella level was raised to p6. the patient is stable. | [5 empty] | The investigation for the embolism has been completed. product was not returned for investigation. clinical information mentions that the pump prepped per ifu and after 30 minutes air was noted in lv. data logs were investigated. there was no "air in purge" alarm noticed at the time of failure or throughout the case. it is not possible to investigate further without the product being available and come to a conclusion. therefore, the root cause of the embolism issue was not determined since product was not returned and data logs and clinical information provided was inconclusive. d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
25009 | 1220648-2024-12775 | 19633612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633612&pc=OZD | 06/20/2024 | 01/15/2025 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1122739 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) when booting up for a case received a system self-check failed with replace immediately alert. the aic was not used. no patient was on support at the time of error. | [5 empty] | The investigation for the console (aic) has been completed. imc logs were reviewed finding no evidence of a controller error was found in the logs during this time range, however this does not preclude the occurrence of a controller error as reported. the only relevant evidence found in the imc logs was repeated failures in communication between the epc and the pmd and/or cpld. this pattern repeated for the entire duration of the imc logs. failure in communication between the epc and the pmd and/or cpld could have resulted in a controller error, as well as triggered the system self-check failure. console was sent back for evaluation. the reported complaint was reproduced through testing; the console issued a system self-check failure upon attempting to boot. the power delivered to the cpld was measured and confirmed to be 3.3v, the expected value. the communication between the epc and the pmd was monitored with an oscilloscope. no immediate anomaly was identified, however the communication did appear different than that of a known-good console. the impellatronic board was replaced with a known-good impellatronic board and the console was able to boot up as expected and run a known-good pump. the cause of the controller and self check error was a defective impellatronic board. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25010 | 1220648-2024-12776 | 19633678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633678&pc=OZD | 06/21/2024 | 08/13/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  | 05/31/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024368964 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an (b)(6)-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a small hematoma, it was expressed, and the issue resolved. it was also reported that when the peel away sheath was flushed, there was a cracked noted in the first 4 centimeters. the physician though the crack happened when they placed the 18g needle when accessing sheath for single access. the patient was reported as stable. | [5 empty] | The investigation for the access site bleed and introducer damage has been completed. the product was not returned and data logs are not relevant. the root cause of the access site bleed could not be determined as not enough clinical evidence was provided to confirm when the bleed started or if the introducer damage could have caused the bleed. the root cause of the introducer damage was determined to be most likely a use issue during the single access procedure. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25011 | 1220648-2024-12779 | 19633814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633814&pc=OZD | 06/21/2024 | 08/01/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482058 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site bleeding. the tech applied pressure on the sheath and tented the arteriotomy multiple times. the doctor decided that the bleeding would not resolve and the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12779: e4 should have been left blank. g1 reporting contact fax number missing.
25012 | 1220648-2024-12777 | 19633861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19633861&pc=OZD | 06/09/2024 | 08/12/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1363622 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the patient on impella support was transferred to another facility and when the automated impella controller (aic) was returned to the first facility the aic bracket was broken. the aic damage was after it was removed from the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) damage has been completed. console logs from the device did not contain information relevant to the damage, but ambient pressure variations observed in data log for pump were consistent with console transport. the console was returned for evaluation. upon visual inspection, damage to the aic support bracket was confirmed. results of visual inspection of the bracket suggested inferior quality of the welds. the most likely root cause was due to a manufacturing defect. no other mechanical damage was observed, and console appeared to be fully operational. added- d9 device return date.
25013 | 1220648-2024-12778 | 19634077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19634077&pc=OZD | 06/21/2024 | 08/13/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient¿s right leg was cold and ultrasound showed no flow. the following day after the impella cp was explanted, a ct scan showed possible clot on femoral head. the patient was taken to the operation room for a thrombectomy and fasciotomy. the patient was stable. the heart team decided to escalate the patient to impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. added- h6 clinical code 1942.
25014 | 1220648-2024-12781 | 19634139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19634139&pc=OZD | 06/08/2024 | 08/01/2024 |  | 06/28/2024 |  |  | 06/08/2024 | 06/28/2024 |  |  | 07/31/2025 |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433849 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 82-year-old female patient in non-st elevated myocardial infarction was to be implanted with an impella cp device for mechanical circulatory support. it was reported that following implant, the device immediately began to experience suction and flows would not exceed 0.5 liters per minute (lpm). the pre-scheduled procedure was completed, after which the left ventricle appeared to be under filled and decompressed. a subcostal view identified a large pericardial effusion, and a cardiac tamponade was suspected. a pericardial centesis was completed and impella flows increased to 2 lpm. based on the noted condition, the doctor initially suspected a left ventricle perforation, but after no visual evidence of a perforation was found, the physician later concluded the effusion was caused by cardiopulmonary resuscitation. at this point in time, it was also documented that no pulses were present in the impella leg. due to the noted conditions, the decision was made to remove the pump. the patient was reported to be in critical condition. | [5 empty] | The investigation for the limb ischemia and the low pump flow has been completed. the root cause of the low pump flow issue cannot be determined since the complaint device not returned for investigation and insufficient clinical details provided. no data logs were available for review. the root cause of the limb ischemia could not be determined without additional clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25015 | 1220648-2024-12780 | 19634270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19634270&pc=PYX | 06/21/2024 | 08/22/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1108700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male of unknown race and ethnicity in st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. it was reported that the patient was transferred to another facility and in the flight controller error alarm occurred. once arrival at the new facility the automated impella controller (aic) was exchanged. no patient harm was reported. | [5 empty] | The investigation for the automated impella controller (aic) - controller error (yellow alarm) has been completed. the console displayed ¿controller error (rp dual sensor, optical bench ambient pressure out of range) ¿ alarm¿. this aligns with the reported failure mode. the log confirmed that the ambient pressure was out of range below 550 mmhg for more than two minutes, which triggered the controller error alarm per the spec requirement. this confirms that the aic is working per the intended design. this console was tested. after running the optical test pump for more than 10 hours, the console operated as intended at the ambient pressure of 755mmhg. no hardware issue was found with the console hardware. during testing, the console displayed ¿purge system blocked - alarm¿, most likely due to a kinked purge connecting tube unrelated to the reported failure mode. upon visual inspection, no dextrose residue was found inside the aic. the controller error was triggered per the design, and no product malfunction was found. the controller error was triggered per the design, and no product malfunction was found. the device functioned/operated to specification. b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12780. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12780 as it is unknown the initial reporter also sent the report to the fda. d.6a and d.6b should have been left blank in initial manufacturer device report number 1220648-2024-12780as automated impella controllers are not implanted. h.6 code 2199 and 4114 were entered incorrectly on the initial manufacturer device report number 1220648-2024-12780. a new code was added. h.10 additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-12780 stated that the device was not returned but the device was returned. section d.9 was revised to reflect this information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25016 | 1220648-2024-12782 | 19634490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19634490&pc=OZD | 06/20/2024 | 08/20/2024 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175897 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed an error message stating the purge disc was not detected. the aic was replaced. there was no patient harm related to this event. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) purge disc/flag issues has been completed. data logs were reviewed. the reported purge disc not detected issue was confirmed. no logs were found from the reported. however, there was one instance of a purge disc not detected alarm (event set #402) found. this alarm was quickly resolved after the purge disc was inserted. the aic was returned for evaluation. the reported issue was unable to be reproduced through testing and no visible damage/anomalies found after inspecting the purge assembly. the reported purge disc not detected issue was confirmed in the imc logs but since it was quickly resolved after the purge disc was inserted and was unable to be reproduced, it is likely that this was a use related issue. added- mso review, h6 med dev prob code 2994.
25017 | 1220648-2024-12783 | 19634561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19634561&pc=OZD | 06/20/2024 | 08/05/2024 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | Thrombosis/Thrombus | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 79-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient went into ventricular fibrillation (vf) following the implant. medical staff started the cp on auto with flows of three l/min along with the persistent vf. medical staff repositioned multiple times to attempt to break the arrhythmia, and despite four shocks and cardiopulmonary resuscitation, the patient remained in vf. medical staff then decided to pull the cp back into the descending aorta, which broke the vf. angiograms performed showed the left circumflex artery completely reoccluded. medical staff tried to use penumbra to retrieve the clot, but the clot migrated to the most distal obtuse marginal, and the left circumflex artery returned to good flow. medical staff then successfully reinserted the cp into the left ventricle, with flows of 3.6 l/min. before the sheath was sutured, the patient went back into vf, and medical staff decided to explant the cp. the patient survived the event. | [5 empty] | The investigation of vf/vt/af/at and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 codes 4755 was reported incorrectly on manufacturer device report 1220648-2024-12783. new code was added to component code. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25018 | 1220648-2024-12784 | 19634613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19634613&pc=OZD | 06/20/2024 | 08/21/2024 |  | 06/28/2024 |  |  | 06/20/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025479143 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was access site bleeding and suction was used to remove any material located at anastomosis and fogarty balloon was inflated and withdrawn into graft to ensure there was nothing present. bleed back attempted again with expected flow rate. one unit of blood products were administered. the patient remains on support and is stable. | [5 empty] | D.6b explant date was added. the investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be an use issue because an additional procedure was performed to make sure of no blockages. the loss of blood has occurred during the procedure relating the cause of bleed to be use issue. h.6 code 4755 and 4641 was reported incorrectly on manufacturer device report 1220648-2024-12784. a new code was added to component code. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12784 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25019 | 1220648-2024-12785 | 19635127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19635127&pc=OZD | 06/21/2024 | 08/01/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485284 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 84-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient developed limb ischemia and the pump was removed. the patient was reported to be in critical condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12785: g1 manufacturing site name and address.
25020 | 1220648-2024-12786 | 19635220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19635220&pc=OZD | 06/13/2024 | 08/15/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025482158 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age and gender in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella placement was aborted due to patient¿s peripheral artery disease and calcification. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of calcified condition of vessels. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk pci in the section: warnings & cautions: warnings. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25021 | 1220648-2024-12789 | 19635724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19635724&pc=OZD | 06/21/2024 | 08/08/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025471095 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old female patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that multiple attempts were made to implant the impella and they were unsuccessful. the implant was aborted due to extreme tortuosity of the right axillary artery and the pump was unable to pass and be placed. the impella cp was prepped and successfully placed instead. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump was aborted after multiple attempts and vessels were tortuous. the instructions for use referenced in the initial report is the impella 5.5 with smartassist for use during cardiogenic shock in the following sections: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves.
25022 | 1220648-2024-12790 | 19635995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19635995&pc=OZD | 06/21/2024 | 08/12/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468739 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 79-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported when removing the peel-away sheath and advancing the repositioning sheath, medical staff inadvertently advanced the cp too far into the left ventricle. the cp then folded back on itself, creating a kinked cannula and a pump failure. medical staff then replaced the cp with another one. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the cannula kink has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the product damage (cannula kink) issue most likely was use related issue due to improper technique as impella was inadvertently advanced too far in the left ventricle and folded back on itself while attempted to reposition.
25023 | 1220648-2024-12787 | 19636239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19636239&pc=OZD | 06/21/2024 | 08/13/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468743 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and they had access site bleeding. an air vac was used and 8 units of blood products were administered. support continued. additionally, low pump flows were noted. 2 units of blood products were given and albumin and suction issues have resolved. reportable due to access site bleeding and low pump flows. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the low pump flow has been completed. the impella was not returned for evaluation. clinical details were not sufficient to determine the root cause. the root cause of the access site bleeding cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. data logs were reviewed finding no motor current spikes occurred. no positioning issue observed. many suction alarms occurred during the case and change in p-levels and flows observed. based on the clinical details best practice applied and no issue stated during the impella insertion. the cause of the low pump flow issue most likely was patient condition related that led to drop in flows at p-1 with multiple suction alarms. added-b5 mso review, d6b g1-corrected mfr site name and address
25024 | 1220648-2024-12788 | 19636304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19636304&pc=OZD | 06/21/2024 | 09/11/2024 |  | 06/28/2024 |  |  | 06/21/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693473 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 53-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) showed a yellow controller error. medical staff then replaced the aic. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) yellow alarm has been completed. data logs were reviewed which showed that the controller error 1040 is triggered when the ambient pressure is out of range for at least two minutes, and according to the ifu, that range is -1000ft to 8000ft which is approximated as 550mmhg. the ambient pressure falling below the 550 mmhg threshold for two minutes triggers the 1040 controller error (ic10222 ambient pressure 550mg.png) and resolves when the pressure came back above 550 mmhg. the console was returned for investigation but the failure mode was not reproduced since the ambient pressure of the lab was above 550 mmhg. therefore, the cause of the issue was the console going above 9000ft dropping ambient pressure below 543 mmhg as a use-related issue. added- d9 device return date, h6 impact code 4629 h5 changed to no. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25025 | 1220648-2024-12791 | 19637843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19637843&pc=OZD | 06/22/2024 | 10/01/2024 |  | 06/28/2024 |  |  | 06/22/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1769928 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that automated impella controller (aic) had an unexpected shutdown. the aic shutdown and restarted with no interaction from staff. the aic was exchanged. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) shutdown has been completed. data logs were reviewed which showed the first complaint occurred when the io-graphics process crashed. immediately after, the console rebooted in response followed by an io-graphics core dump and an unexpected shutdown alarm confirming what was reported clinically. therapy was resumed before the pump was unplugged and moved over to (b)(6). the second failure was an unreliable placement signal alarm that continued through the end of the case. the lt logs show that the placement signal declines over 100 mmhg before jumping to 3000mmhg triggering the unreliable placement signal alarm. console with serial number (b)(6) was returned for investigation which was the incorrect console. (b)(6) does not need to be returned as this is a 3rd party software failure that is unlikely to reproduce. (b)(6) also does not need to be returned as the failure was caused by pump degradation. the cause of the unexpected reboot was the 3rd party hardware/software qnx that runs io-graphics. there was no issue found during the case. the device functioned/operated to specification. added-d9 device return date, h6 impact code 4629.
25026 | 1220648-2024-12792 | 19638004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19638004&pc=OZD | 06/21/2024 | 09/03/2024 |  | 06/28/2024 |  |  | 06/22/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2025477676 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. shortly after the impella was implanted and started up, the pump stopped. the impella was explanted and the patient was supported with intra-aortic balloon pump and extracorporeal membrane oxygenation. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of pump stop has been completed. the pump was returned for investigation. significant catheter kink and cannula kink were found on the returned product. the pump passed electrical testing and laser testing. during a test in a bucket of water, the pump stoppage was reproduced. further investigation revealed damage to the outflow cage inside the cannula, which was obstructing the impeller from rotating. according to clinical details, pump stop was reproduced at the start of the case. the data log recorded only impella plug connect prompt. the cause of the pump did not start issue was a damaged outflow cage due to excessive force applied during pump implant.
25027 | 1220648-2024-12793 | 19638433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19638433&pc=OZD | 06/23/2024 | 04/11/2025 |  | 06/28/2024 |  |  | 06/23/2024 | 06/28/2024 |  |  |  |  | 06/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103511 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 44-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) did not show a placement signal. medical staff exchanged the aic to resolve the issue, and no patient harm has been reported. | [5 empty] | The impella device has been received from the customer, upon completion of the investigation of the device, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. console logs were consistent with what was reported in the complaint. the reported issue was only able to be reproduced during testing by partially inserting the pump into the pump receptacle. the root cause of the reported placement signal issue was likely an air gap between the pump and pump receptacle. d2 common device name revised on manufacturer device report 1220648-2024-12793 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-12793 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12793.
25028 | 1220648-2024-12716 | 19623099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623099&pc=OZD | 06/14/2024 | 07/30/2024 |  | 06/27/2024 |  |  | 06/14/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491908 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Ischemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 66-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed limb ischemia during impella support ¿ medical staff observed no distal pulses in the lower extremity of the impella side and could not establish flow via doppler. medical staff rapidly weaned the patient off the cp, and, when removing the cp, the patient went into arrest, requiring resuscitation. the patient went in and out of ventricular tachycardia before eventually expiring. the cause of death is unknown and the cause of death in relationship to the impella remains unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and ventricular tachycardia arrest has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and ventricular tachycardia arrest could not be determined. h6 component code 4755 changed to 925
25029 | 1220648-2024-12659 | 19623162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623162&pc=OZD | 06/03/2024 | 07/24/2024 |  | 06/27/2024 |  |  | 06/03/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced rhythm issues and multiple shocks were applied. the pump also had persistent lows flows/suction and so the pump was explanted and an intra-aortic balloon pump was used. patient care was withdrawn and the patient expired. the medical safety office determined abiomed does not have enough information to associate use of the device with the outcome. | [5 empty] | The investigation for low pump flow and vf/vt/af/at has been completed. the impella device was not returned for evaluation. data logs reviewed: many suction alarms occurred. motor current spikes observed at first hour of impella support. low pump flow observed during the case. no positioning related alarm observed. the cause of the low pump flow issue most likely was biomaterial ingestion. as the necessary information was not provided, the root cause of the vt/vf was not determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12659: a5 ethnicity incorrect. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25030 | 1220648-2024-12717 | 19623265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623265&pc=OZD | 06/14/2024 | 08/16/2024 |  | 06/27/2024 |  |  | 06/16/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025508128 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age, gender, ethnicity, and race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support there was kinking in the cannula that led to low flow. the patient is noted to have expired. no alternative cause for the death event was identified. | [5 empty] | The investigation of product damage (cannula kink) has been completed. data was not returned for review. visual inspection found a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. based on clinical description and product analysis, the root cause of low flow was most likely kinked cannula. the root cause of low flow was most likely kinked cannula. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12717. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25031 | 1220648-2024-12718 | 19623295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623295&pc=OZD | 06/15/2024 | 08/09/2024 |  | 06/27/2024 |  |  | 06/16/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479171 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The complainant reported a 20-year-old male patient with acute cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported high purge pressure and blocked purge system alarms. medical staff added tissue plasma activator (tpa) to the purge. the purge pressures remained in the 1000-1050 range with flows of 1-2 ml/hr. medical staff then planned to use the remainder of the 50 ml bag of tpa for infusion, then switching back to d5 with bicarb. medical staff also withheld anticoagulation due to a gastrointestinal bleed. the patient remains on impella support. |  |  |  |  | The investigation of high purge pressure and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the high purge pressure issue was most likely biomaterial based on data log review and given clinical details. the cause of the vascular damage issue was not determined due to insufficient clinical information and unknown relation to impella. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25032 | 1220648-2024-12719 | 19623414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623414&pc=OZD | 06/16/2024 | 03/19/2025 |  | 06/27/2024 |  |  | 06/16/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1329223 |  | 0042-0000-AU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 42-year-old male patient with acute cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) had a battery failure. the aic alarmed constantly even when unplugged and shipped with the patient¿s escalation transfer. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) battery failure (red alarm) has been completed. the impella device and data was not returned for evaluation. the root cause of this issue was the aic being booted up while the battery transport switch was disengaged. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12719. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-12719 was submitted in accordance with updated procedures. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-12719 was submitted in accordance with updated procedures.
25033 | 1220648-2024-12720 | 19623462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623462&pc=OZD | 05/27/2024 | 02/25/2025 |  | 06/27/2024 |  |  | 06/09/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1462548 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 15-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were purge pressure alarms and it was found the yellow-to-yellow connection on the purge sidearm was wet. the purge cassette and tubing were replaced, and the issue resolved. on a different day the alarms returned, and the automated impella controller (aic) was replaced, which did not resolve the issue. the purge cassette was replaced again, and the issue resolved. the aic was sent for investigation. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the data logs confirm that prior to purge cassette change, multiple purge error messages and alarms were issued. during the inspection of the purge assembly of the console revealed no abnormalities. the console was able to operate to specification with a known good purge system and pump. the cause of the purge issue was not determined, as the issue could not be reproduced during the testing. d2 common device name revised on manufacturer device report 1220648-2024-12720 in accordance with updated procedures. d6a and d6b should have been left blank on manufacturer device report 1220648-2024-12720 e4 should have been left blank on the initial manufacturer device report number 1220648-2024-12720 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-12720 in accordance with updated procedures.
25034 | 1220648-2024-12722 | 19623518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623518&pc=OZD | 06/16/2024 | 08/05/2024 |  | 06/27/2024 |  |  | 06/16/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450592 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported the pump and wire interacted, which failed the delivery. medical staff successfully replaced the failed pump on the same side access on the patient. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to use issue since the doctor was unable to advance due to wire entangled on catheter pigtail because of improper technique.
25035 | 1220648-2024-12723 | 19623556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623556&pc=OZD | 06/17/2024 | 11/20/2024 |  | 06/27/2024 |  |  | 06/17/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103167 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon start up, an aic (automated impella controller) displayed a yellow "controller error" message. the aic was swapped for the planned support. there was no harm to the patient. | [5 empty] | The investigation for the console (aic) controller alarm-yellow alarm has been completed. data logs were reviewed and confirm that the yellow controller alarm was issued on the day of the complaint. the console was internally inspected and ribbon cables were found to be intact, suggesting no physical damage. power supply was inspected and found to be within specification. the failure mode could not be replicated after recycling 200 times.the root cause for the "loss of communication with keypad or between the kbd and i2c for kbd hw revision 1" alarm could not be determined due to not being able to reproduce. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25036 | 1220648-2024-12725 | 19623589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623589&pc=OZD | 06/15/2024 | 12/04/2025 |  | 06/27/2024 |  |  | 06/15/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1375165 |  | 0042-0000-AU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 42-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) had a battery malfunction. medical staff reported constant alarming even when unplugged and the screen showed nothing. no patient harm has been reported, and the patient remains on impella support. | [5 empty] | The investigation has been completed. the console displayed ¿battery failure ¿ alarm,¿ event. this confirms the reported failure mode. after the service, the console was booted on june 15, 2024 - likely at the customer site - and at initial boot the console displayed the battery failure alarm. this most likely indicates the battery was disengaged, which resulted in the alarm. the root cause of the battery-failure alarm was most likely a disengaged battery switch. there was no log evidence to confirm false / buzzer alarm. it is most likely the battery-failure alarm was misreported as an additional failure mode: "aic ¿ false / buzzer alarm." d2 common device name updated. g1 reporting contact email and reporting contact fax number updated. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25037 | 1220648-2024-12724 | 19623596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623596&pc=OZD | 06/15/2024 | 08/07/2024 |  | 06/27/2024 |  |  | 06/15/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488515 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a female patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported that, while the cp was being primed for insertion, the automated impella controller showed an impella defective alarm. the alarm would not clear, so a different cp was used successfully. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of pump not detected has been completed. the impella device was received for investigation. the root cause of the pump not detected alarm could not be determined as the failure mode did not reproduce. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12724: d3 manufacturer zip code, country, and fax number. d4 model number. e4 should have been left blank. g1 manufacturing site name and address. g2 report source: foreign was inadvertently selected. h.6 code 765 was reported incorrectly.
25038 | 1220648-2024-12726 | 19623605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623605&pc=OZD | 06/15/2024 | 08/07/2024 |  | 06/27/2024 |  |  | 06/15/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493242 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported an 81-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported an impella defective alarm during preparation for a right femoral access implant. medical staff removed the cp, and no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the pump not detected has been completed. clinical details state that the console had an "impella defective" alarm, so the team switched to another pump for insertion. the logs and product were returned and evaluated. an "impella defective" alarm was not found in any of the data logs for pump. there were no abnormalities found with the pump or logs; the pump ran to specification in the lab with no alarms being reproduced. there was no problem found while investigating the reported failure mode of pump not detected. added- d9 device return date g1- corrected mfr site name and address.
25039 | 1220648-2024-12727 | 19623626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623626&pc=OZD | 06/15/2024 | 08/20/2024 |  | 06/27/2024 |  |  | 06/15/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported an 81-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported positioning issues ¿ after the pump was inserted, the pump was deep. suction and a drop in blood pressure occurred when medical staff pulled the pump back. the inlet appeared to be papillary or mitral valve. medical staff were unable to reposition the pump; the pump was pulled back across the aortic valve and unable to re-cross. medical staff then explanted the pump. the patient survived the event. | The patient experienced pulseless electrical activity (pea) arrest and resuscitation was attempted but unable to achieve return of spontaneous circulation (rosc) and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and low pump flow has been completed. data logs confirmed pump was in ventricle position (about 1.25 hours into the case) corresponded with clinical description that triggered alarms impella position in ventricle. logs also showed low flow after that. the positioning issue was resolved within 2 minutes then happened again 40 minutes later. next, pump was in aorta and stopped manually. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related due to improperly technique inserting the repo unit resulting in the pump going deep. the root cause of the low flow was most likely due to pump was in improper position. b5 updated with additional information the following have been reported incorrectly or missing from manufacturer device report: e4 should have been left blank. g1 reporting contact fax number missing.
25040 | 1220648-2024-12728 | 19623631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623631&pc=OZD | 06/15/2024 | 08/07/2024 |  | 06/27/2024 |  |  | 06/15/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485851 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported even though the goal was to place both the 5.5 and extra-corporeal membrane oxygenation, the 5.5 pump delivery failed due to the patient¿s anatomy. medical staff then planned to implant an impella cp instead of the 5.5. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through patient anatomy. added- h6 impact code 4629.
25041 | 1220648-2024-12729 | 19623650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623650&pc=OZD | 06/17/2024 | 07/24/2024 |  | 06/27/2024 |  |  | 06/17/2024 | 06/27/2024 |  |  | 08/31/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398211 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 35-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was a large hematoma noted. the doctor ordered to stop angiomax and a fem-stop device was applied proximal to the cp insertion site. additionally, there was bleeding noted from the insertion site. the fem-stop was removed, angle matching was performed and the bleeding subsided. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12729: g1 manufacturing site name and address. h.6 code 4755 was reported incorrectly. new code was added to component code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25042 | 1220648-2024-12730 | 19623651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623651&pc=OZD | 06/15/2024 | 08/17/2024 |  | 06/27/2024 |  |  | 06/15/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced access site bleeding. a femstop was placed and one unit of packed red blood cells was provided. the patient was noted to be in stable condition. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12730 as it is unknown the initial reporter also sent the report to the fda. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12730. g.2 foreign was selected incorrectly on the previously submitted manufacturer device report number 1220648-2024-12730. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25043 | 1220648-2024-12731 | 19623691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623691&pc=OZD | 06/16/2024 | 08/07/2024 |  | 06/27/2024 |  |  | 06/16/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482639 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported the access site was oozing but required no intervention. medical staff observed suction due to poor positioning and could not reposition to a good left ventricle position, so they explanted the pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the positioning issue has been completed. data logs were reviewed which confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related.
25044 | 1220648-2024-12721 | 19623693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623693&pc=OZD | 06/13/2024 | 09/01/2024 |  | 06/27/2024 |  |  | 06/13/2024 | 06/27/2024 |  |  | 05/31/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET, CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2024375133 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 36 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Aortic Valve Insufficiency/ Regurgitation | Tricuspid Valve Insufficiency/ Regurgitation | Renal Impairment | [7 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Additional information was received from the study team. microcytic anemia was diagnosed. hemoglobin dropped post procedure and the patient was iron deficient. the team treated the patient with 300mg venofer every other day for three doses. the relationship was also "possible". significant oozing was noted to the right groin. manual pressure was applied and the impella was removed per protocol for resolution of the bleeding. another adverse event of vascular complication in which the superior mesenteric artery occlusion and embolism lead to splenic and renal infarcts with the relationship to be "probable" to impella. | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in canada had a cp placed to support a stemi (st elevation myocardial infarction) patient. the patient was enrolled in the stemi dtu protocol. the pump support was successful for the pci (percutaneous coronary intervention) and just under 7 hours. weeks after explant the crf form was shared in which 2 ae with possible and probable relationship to the use of the cp was shared by the hospital. there was a possible relationship for the vf (ventricular fibrillation) arrest and shock, and the other ae of aortic valve damage at a cusp was probable relationship to the impella. |  |  |  |  | B.5 additional information was received and was added. h.6 health effect - clinical code and health effect - impact codes were updated to reflect new information provided. h.10 instructions for use were updated with new information to reflect all adverse events reported. the investigation of ventricular fibrillation (vf), heart valve damage, and access site bleeding ¿ major has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vf was not determined. the root cause of the heart valve damage issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. as noted in the impella instructions for use: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿. section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 revised as required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12721. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12721 as it is unknown if the initial reporter also sent the report to the fda. g.1 revised reporting contact facility name and contact fax number as they inadvertently omitted from the initial manufacturer device report number 1220648-2024-12721. g.2 revised as study was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12721. h.6 codes 4453 and 4114 were reported incorrectly on the initial manufacturer device report number 1220648-2024-12721. new codes were added. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
25045 | 1220648-2024-12732 | 19623701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623701&pc=OZD | 06/16/2024 | 08/08/2024 |  | 06/27/2024 |  |  | 06/16/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025493735 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was placed but after 30 minutes the pump came out of position. the malpositioned pump was removed and replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the pump was not returned for investigation. data logs show a pump position in aorta alarm at the end of the case. according to the clinical details, the pump was pulled back into the aorta, and the doctor replaced the pump. the cause of the positioning issue was use related since doctor pulled pump back to aorta. added- h6 impact code 4628.
25046 | 1220648-2024-12733 | 19623702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623702&pc=PYX | 06/17/2024 | 08/12/2024 |  | 06/27/2024 |  |  | 06/17/2024 | 06/27/2024 |  |  | 05/31/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024346886 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Noise, Audible | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the patient was also implanted with an impella cp. during support, the impella rp was making sounds and vibrations with high purge flow errors. there has been no harm or injury to date from the noise and vibration and the patient remained on support for five additional days. | [5 empty] | The investigation for the reported noise and high purge pressure has been completed. according to clinical details, nurse was concerned about a humming noise coming from the pump catheter. additionally, they noted that purge flow rates were "high" but within specification. the pump was not returned for investigation. field data logs were investigated and there were no signs that purge flow or pressure were abnormal. the only purge related alarms coincide with purge bag changes. additionally, no other metrics were out of specification. the root cause of the cosmetic defect was determined to be noisy or audible sound. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
25047 | 1220648-2024-12735 | 19623733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623733&pc=OZD | 06/18/2024 | 08/13/2024 |  | 06/27/2024 |  |  | 06/18/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491254 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and they had access site bleeding after the implant. two units of blood products were provided. the surgeon attempted to cinch 1 perclose suture tighter and it broke. repo sheath of impella was also advanced too far causing blue suture hub to enter arteriotomy. bleeding was not controlled despite mitigation efforts including manual pressure, fem stop, additional sutures and heparin reversal. the impella was explanted. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system. g1-corrected mfr site name and address h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25048 | 1220648-2024-12734 | 19623781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19623781&pc=OZD | 06/17/2024 | 08/13/2024 |  | 06/27/2024 |  |  | 06/17/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1598659 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) that while the device was being primed and prepped, there were "purge system open" and "purge system blocked" alarm. the staff replaced the purge cassette but this did not resolve the problems. the aic was replaced. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - purge issue has been completed. after reviewing the logs, the reported purge system open and purge system blocked alarms were confirmed. the reported purge system open and purge system blocked alarms were unable to be reproduced. the root cause of the issue was use related. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12734. d4 model number. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25049 | 1220648-2024-12736 | 19624177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624177&pc=OZD | 06/18/2024 | 08/13/2024 |  | 06/27/2024 |  |  | 06/18/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486095 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The supraventricular tachycardia was treated by pump repositioning. | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed to support a stemi (st elevation myocardial infarction) patient who needed care escalated from the iabp (intra-aortic balloon pump) . the 5.5 pump was alleged to be the cause of svt (supraventricular tachycardia). |  |  |  |  | D6.b explant dated was added. the investigation ventricular tachycardia/atrial fibrillation has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia/atrial fibrillation was not determined. b.7 revised as information was inadvertently omitted from the initial manufacturer device report 1220648-2024-12736 submitted. b.5 revised as information was inadvertently omitted from the initial manufacturer device report 1220648-2024-12736 submitted. h.6 code 4611 and 4755 were reported incorrectly on the initial manufacturer device report 1220648-2024-12736 submitted. new codes were added. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted initial manufacturer device report number 1220648-2024-12736 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25050 | 1220648-2024-12737 | 19624261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624261&pc=OZD | 06/18/2024 | 08/08/2024 |  | 06/27/2024 |  |  | 06/18/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025469441 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 39-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported purge system blocked alarms with purge flows less than 1 ml/hr and purge pressures over 1,000 occurred. medical staff checked the purge line for kinks and changed the purge cassette, which did not resolve the alarm issue. medical staff noted normal limits for motor current and decided to initiate tissue plasminogen activator (tpa) protocol. the patient has been on impella support for 49 days and is on biventricular assist device support, working up for a transplant. trans-thoracic echocardiograms showed worsening left ventricle (lv) function. the lv waveform increased over the aortic waveform with equalization of maximum/minimum flows. medical staff expected pump saturation and noted slowly increased motor currents over an hour¿s time. medical staff then replaced the 5.5, and the patient remains stable on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the saturated flow issue has been completed. data logs were investigated and there are several periodic raises in purge pressure until a final instance where the purge pressure crosses 1000 mmhg and stays high. that being said, there are no motor current spikes/trends that coincide with the rises in purge pressure. returned product was investigated and there was a kink in the catheter/purge line at the junction between the red handle and catheter. additionally, there appeared to be biomaterial in the purge gap. a window was cut in the catheter and blood had gotten into the purge line. when running the pump in the lab, high purge pressure (hpp) reproduced until a cut was made in the catheter near the motor housing that lead to a release of blood/biomaterial and pressure. further investigation showed that when flexing the existing kink location near the red handle, purge pressure spikes could be reproduced and then resolved when letting go. this indicates that the build up of biomaterial and blood ingress in the purge gap was likely a symptom of the purge line being kinked farther up the catheter. the field data logs further suggest that the short periods of purge pressure rise and drop before the sustained event could likely correlate to instances of patient movement that kinked the purge line blocking off flow repeatedly. this would create time for blood to ingress into the purge line. the root cause of the high purge pressure was determined to be a kinked purge line. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25051 | 1220648-2024-12738 | 19624360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624360&pc=OZD | 06/18/2024 | 07/25/2024 |  | 06/27/2024 |  |  | 06/18/2024 | 06/27/2024 |  |  | 05/31/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356517 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 85-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient experienced ventricular tachycardia. while remaining on impella support, medical staff repositioned the pump after the tachycardia runs. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12738: h.6 code 1721 and 4755 were reported incorrectly. new code was added to component code.
25052 | 1220648-2024-12739 | 19624440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624440&pc=OZD | 06/06/2024 | 07/30/2024 |  | 06/27/2024 |  |  | 06/07/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487315 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2024 |  |  |  |  | The complainant reported a 69-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient¿s right leg, the site of insertion, became ischemic. the medical staff removed the pump and reported that the patient¿s leg was cold. the medical staff continued to monitor the patient, and the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- h6 component code 925.
25053 | 1220648-2024-12740 | 19624453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624453&pc=OZD | 06/19/2024 | 08/12/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485614 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support when there was an increased output from the jp(jackson-pratt drain) and a large hematoma noted. a hematoma evacuation was performed. additionally, the impella was repositioned due to the patient's ventricular tachycardia (vt). no arrhythmias after repositioning. | [5 empty] | The investigation for the ventricular tachycardia (vt) and the access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and the vt could not be determined. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25054 | 1220648-2024-12741 | 19624548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624548&pc=OZD | 06/19/2024 | 08/13/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025501497 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old female patient in in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that flow saturation was noted on impella cp. the cp was exchanged for another cp. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of saturated flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
25055 | 1220648-2024-12742 | 19624607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624607&pc=OZD | 06/19/2024 | 07/31/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025480583 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old male patient of unknown race and ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the repositioning sheath became dislodged. there was "a lot of blood" around the area and the patient was taken emergently to the operating room for removal and vascular repair. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
25056 | 1220648-2024-12744 | 19624702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624702&pc=OZD | 06/19/2024 | 08/12/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458092 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) was displaying low pump flows. the aic showed decreased flow rates, decreased map, and aortic/left ventricle waveform separation. the echocardiogram showed the left ventricle had been sucked down completely. an impella rp was implanted to help support the patient. the low pump flows were resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. product was not returned for investigation. data logs were investigated. there were about 30 suction alarms seen during the time of the complaint. as the pump was on auto mode, decreasing the p-level and therefore the flow at the time of suction is an automated response of the pump. as per the clinical the "lv was sucked down" which suggests that the patient needed right heart support and had lower blood volume. this was resolved by adding an rp pump to the treatment. therefore per the clinical information and data log review, the root cause of the low pump flow issue was most likely patient condition related.
25057 | 1220648-2024-12743 | 19624706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624706&pc=OZD | 06/07/2024 | 07/30/2024 |  | 06/27/2024 |  |  | 06/07/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510277 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and upon removal of the pump, the patient did not have distal pulses noted. an angiogram revealed a clot to the right femoral artery. an ekos catheter was placed. | [5 empty] | The investigation for the reported limb ischemia and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and thrombus could not be determined. | A.5 race is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25058 | 1220648-2024-12745 | 19624889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19624889&pc=OZD | 06/07/2024 | 08/16/2024 |  | 06/27/2024 |  |  | 06/07/2024 | 06/27/2024 |  |  | 09/30/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401641 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during the impella cp procedure after insertion of the repo sheath, the physician performed a groin shot through the re-access port. extravasation was noted and the physician elected to remove the impella and sheath. manual pressure was held but bleeding could not be controlled. attempted to deploy perclose devices but unsuccessful. the vascular surgeon elected to insert a 16 french gore dryseal sheath into the right common femoral artery (rcfa) to maintain hemostasis until transfer to the operating room for closure and vessel repair. the patient was reported as stable. | [5 empty] | The investigation for vascular damage - peripheral vessel has been completed. after pump insertion, the physician performed groin shot and extravasation was noted. no product was returned. the root cause of the vascular damage - peripheral vessel was not determined due to lack of sufficient clinical information. h.6 code 925 was added since the information of the initial manufacturer device report number 1220648-2024-12745 was submitted in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25059 | 1220648-2024-12746 | 19625452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19625452&pc=OZD | 06/19/2024 | 08/26/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  | 03/31/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359591 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 34 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 34-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in australia placed a cp to support a patient admitted to tertiary center. first hospital did not have impella to support the 34yo with heart disease and cad (coronary artery disease). the patient was supported by ecmo (extracorporeal membrane oxygenation) and iabp (intra-aortic balloon pump) when crashed and added the cp support. the cp caused lv (left ventricle) obliteration when the p level was at p2. the team reduced to p1 and remained on with ecmo and consulted the csc for assistance. (b)(6) 2024, "extent of cva suggests recovery unlikely. plan was to wean off ecmo and onto impella and then off impella." patient passed (b)(6) 2024, with unknown relation with impella. | [5 empty] | The investigation of low pump flow and stroke/cva has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. data logs confirmed the failure mode. log analysis show several suction alarms throughout the case with placement signal trend and low flows. also, impella position unknown alarms were seen. the root cause of low flow was most likely patient condition related. mso statement: the decision was made to withdraw care, and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. d3 manufacturer us zip code missing from manufacturer device report 1220648-2024-12746. | The device was not returned for evaluation. after completion of the investigation, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25060 | 1220648-2024-12747 | 19625512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19625512&pc=OZD | 06/19/2024 | 08/06/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  | 06/30/2024 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023237206 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old female noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported a few minutes after implant the flows dropped to 1.0, it was noted on fluoroscopy there was a kink in the impella cannula possibly due to heavy calcification of the aortic arch. the physician was unsure if the low flows were from the kinking of the cannula or from a potential clot aspiration. the impella cp was removed and a new one placed. there was no patient harm reported. | [5 empty] | The investigation for low pump flow has been completed. the pump was not returned for investigation. data logs show low pump flow from the beginning of the case run without a noted motor current spike, placement signal trend, or positioning issue. according to the clinical details, low pump flow was noted from the start of the case. fluoroscopy revealed a kink in the cannula at proximal to the black radiopaque marker. patient had a heavily calcified aortic arch. the cause of the low pump flow issue was most likely patient condition related, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25061 | 1220648-2024-12748 | 19625565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19625565&pc=OZD | 06/19/2024 | 08/27/2024 |  | 06/27/2024 |  |  | 06/19/2024 | 06/27/2024 |  |  | 08/31/2025 |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024382470 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella cp for mechanical circulatory support. the impella bent due to the angle of the patient's aorta and an aortic arch replacement and could not reach the left ventricle. the implant was aborted. |  |  |  |  | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
25062 | 1220648-2024-12749 | 19626021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19626021&pc=OZD | 06/20/2024 | 03/19/2025 |  | 06/27/2024 |  |  | 06/20/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1205056 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller was turned on and battery failure alarm occurred. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) battery failure (red alarm) has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of this issue was the console being booted up while the battery transport switch was disengaged. d2 common device name revised on manufacturer device report 1220648-2024-12749 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-12749 in accordance with updated procedures.
25063 | 1220648-2024-12750 | 19626113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19626113&pc=OZD | 06/20/2024 | 08/30/2024 |  | 06/27/2024 |  |  | 06/20/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025509564 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Physical Resistance/Sticking | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male patient of unknown ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the sterile sleeve was very difficult to pull back and required a lot of force. the physician was unable to pull the sleeve back all the way. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage was complete. the pump was returned for investigation. there was no abnormalities found with the pump. the sterile sleeve and the tuhoy were able to move along the catheter just as they should and no force was required. clinical mentions a record of difficulty moving pump which could not be reproduced during the investigation. the root cause was that the user was not using the correct technique to move the reposition sheath.
25064 | 1220648-2024-12751 | 19626468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19626468&pc=OZD | 06/01/2024 | 07/30/2024 |  | 06/27/2024 |  |  | 06/01/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504419 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported low flows after medical staff initiated cardiopulmonary resuscitation ¿ the cp couldn¿t reach higher flows than p-2. medical staff then exchanged the cp, and the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed about 3.5 hours into the case corresponded with the time post cpr per clinical description, low flow occurred with motor current (mc) spike that triggered auto suction response and suction alarms. this event remained to the rest of the case. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi. in section: entering cardiac output states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ added- d9 device return date.
25065 | 1220648-2024-12752 | 19627614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19627614&pc=OZD | 06/07/2024 | 07/31/2024 |  | 06/27/2024 |  |  | 06/07/2024 | 06/27/2024 |  |  |  |  | 06/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482978 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 69-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the medical staff could not drop the impella 5.5 into the ascending aorta over the 0.018 guidewire. the medical staff attempted multiple times unsuccessfully before removing both the catheter and the wire. upon removal, the medical staff noticed a kink on the cannula, near the outlet. the medical staff then successfully placed another 5.5. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was vascular patient anatomy resulting in a kinked cannula.
25066 | 1220648-2024-12684 | 19613890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19613890&pc=OZD | 05/30/2024 | 07/24/2024 |  | 06/26/2024 |  |  | 05/30/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1446166 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) controller started with a system self-check failure. there was no report of harm to a patient. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) a/c power issues has been completed. the data logs show pbm1 communication errors. the console was returned for evaluation and the failure mode was reproduced. when the console was turned on, it immediately output a system self check fail. the super cap on the pbm was confirmed to have corroded the pbm communication trace next to it. the cause of the aic issue was contamination of a communication line on the pbm. g1-corrected mfr site name and address
25067 | 1220648-2024-12685 | 19613947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19613947&pc=OZD | 07/13/2023 | 07/17/2024 |  | 06/26/2024 |  |  | 05/31/2024 | 06/26/2024 |  |  | 02/28/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337600 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient expired after the procedure. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured vt (ventricular tachycardia) with a "possible" relationship to the impella device used: ¿patient tolerated the procedure well initially although got more hypotensive towards the end. had to restart vasopressors. patient went onto a wide complex rhythm afterwards which appeared to be a rate-related bundle however heart rate was all the way up to 220 beats per minute and could not differentiate it from ventricular tachycardia. patient was shocked with 150 joules and after that patient was maintaining a heart rate of close to 90s with a rate related bundle. at this time patient was transferred back to the icu¿. the patient recovered with sequelae. |  |  |  |  | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25068 | 1220648-2024-12686 | 19613995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19613995&pc=OZD | 06/04/2024 | 08/09/2024 |  | 06/26/2024 |  |  | 06/06/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493834 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported the impella cp pump was preoperatively placed via the right subclavian vein for mechanical circulatory support of a patient admitted in for ventricular septal defect surgery. however, during support, the impella cp pump was causing upper extremity ischemia, and therefore, after just 22 hours, it was explanted. the team instead placed an impella cp pump via the femoral with no report of further complications. |  |  |  |  | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12686. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25069 | 1220648-2024-12688 | 19614251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614251&pc=OZD | 06/13/2024 | 08/21/2024 |  | 06/26/2024 |  |  | 06/13/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1275035 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during surgical preparation, the automated impella controller¿s (aic) purge disc was not recognized. medical staff exchanged the aic to resolve the issue, and the procedure was completed successfully. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12688. h4 labeled for single use yes. | D.9 updated as the console was returned for investigation. the investigation into automated impella controller - purge disc/flag issues has been completed. after reviewing the logs the reported issue that purge disc could not be detected was confirmed. visual investigation showed the purge disc flag was found to be broken (missing at the blue purge disc retainer). the root cause of the issue was a broken purge disc flag. a.3 revised as unknown selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12688. e.1 name prefix/title, first and last name, and facility name were inadvertently left off the previously submitted manufacturer device report number 1220648-2024-12688 and were added. e.4 should have been left blank on the initial submitted manufacturer device report number 1220648-2024-12688 as it is unknown if the initial reporter also sent the report to the fda. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12688. g.2 report source health professional was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12688. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2024-12688. a new code was added.
25070 | 1220648-2024-12690 | 19614338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614338&pc=OZD | 06/13/2024 | 07/30/2024 |  | 06/26/2024 |  |  | 06/13/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention. the patient experienced bleeding from their impella site during support. despite holding pressure and maintaining the angle of entry, the bleeding persisted and the decision was made to remove the pump. an additional perclose was used during removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12690: e4 should have been left blank. g1 reporting contact fax number missing.
25071 | 1220648-2024-12691 | 19614345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614345&pc=OZD | 05/16/2024 | 09/03/2024 |  | 06/26/2024 |  |  | 06/13/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025488860 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was pulled out by the extracorporeal membrane oxygenation (ecmo). the physician therefore explanted the pump and used the impella access site to place new ecmo and implanted a new pump. there was no reported patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella was pulled out by extracorporeal membrane oxygenation. a new pump was placed. no product was returned. the root cause of the positioning issue was most likely a use-related issue. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12691. e.1 revised address line 1 as it was entered incorrectly on the initial manufacturer device report number 1220648-2024-12691. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12691 as it is unknown if the initial reporter also sent the report to the fda. g.1 revised manufacturing site address line 1 and 2, city and zip code were entered incorrectly on the initial manufacturer device report number 1220648-2024-12691. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25072 | 1220648-2024-12693 | 19614370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614370&pc=OZD | 06/06/2024 | 08/02/2024 |  | 06/26/2024 |  |  | 06/06/2024 | 06/26/2024 |  |  | 10/31/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414766 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient expired after explantation. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The complainant reported a 63-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported an arterial doppler showed no flows to the left lower extremity. the medical staff consulted vascular surgery about a distal perfusion catheter. the patient was transferred to the intensive care unit. the patient remained on impella support for 11 days. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added b5 mso review, h6 component code 925. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25073 | 1220648-2024-12694 | 19614394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614394&pc=OZD | 06/14/2024 | 08/08/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486823 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 70-year-old patient with elective placement was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported a clot was observed on the catheter above the motor in a donut shape. the circumference was on the catheter. the patient remained stable on impella support before medical staff weaned and then explanted the 5.5. the patient survived the event. | [5 empty] | The investigation of thrombus has been completed. the root cause of the thrombus cannot be determined since insufficient clinical details were provided. no issues were found with the returned product. instructions for use as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12694. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25074 | 1220648-2024-12695 | 19614436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614436&pc=OZD | 06/14/2024 | 08/07/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486442 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that after the 5.5 insertion, the sterile sleeve was torn. medical staff used tegaderm to resolve the issue. the patient remains stable on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12695: e4 should have been left blank. g1 reporting contact fax number missing. h.6 code 765 was reported incorrectly.
25075 | 1220648-2024-12689 | 19614444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614444&pc=OZD | 06/13/2024 | 07/24/2024 |  | 06/26/2024 |  |  | 06/13/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506804 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had some oozing from the impella access site. perclose was tightened and epi lido was injected for hemostasis. it was further reported that access site bleeding occurred from peel away while md was repositioning the cp. systemic heparin was stopped. the patient remains on impella cp support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation the root cause of the access site bleeding was most likely due to anticoagulation management since hemostasis was achieved by stopping heparin. the following have been reported incorrectly or missing from manufacturer device report. b2 outcomes attributed to adverse event missing. e4 should have been left blank. g1 reporting contact fax number missing.
25076 | 1220648-2024-12696 | 19614491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614491&pc=OZD | 06/14/2024 | 03/20/2025 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684102 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) continued to prime the pump during the priming of the purge cassette - the blue progress bar never appeared, but fluid continued to exit. the 5.5 continued to operate without a replacement cassette, but medical staff exchanged the aic regardless, with success. no patient harm has been reported, and the patient remains on impella support. | [5 empty] | The investigation of automated impella controller (aic) purge issue ¿ other has been completed. the reported issue was confirmed in the logs. there were numerous instances of the console continuously priming device analysis: the reported issue could not be reproduced in the lab. while testing, the aic primed normally after multiple attempts to reproduce the issue. conclusion: suspected to be a use related issue. purge cassette was likely inserted but not flush with the purge assembly. this causes intermittent detection of the purge cassette which leads to constant priming of the purge because the console keeps thinking that a new purge cassette was inserted. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-12696 in accordance with updated procedures. d6a should have been left blank on manufacturer device report 1220648-2024-12696 g1 reporting contact email revised on manufacturer device report 1220648-2024-12696 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12696. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25077 | 1220648-2024-12687 | 19614504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614504&pc=OZD | 06/04/2024 | 08/22/2024 |  | 06/26/2024 |  |  | 06/13/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025476486 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old female was pre-operatively implanted with an impella cp device for mechanical circulatory support. it was reported that air was discovered in the left ventricle (lv) following start-up of the implanted impella cp pump. it was stated that proper priming and air removal was performed during setup and a cause for the air in the lv was unknown. the condition was monitored in the intensive care unit with no known intervention documented. the patient was reported stable. | [5 empty] | The investigation of embolism has been completed. no visual abnormalities were noted on pump or purge cassette. data logs were also reviewed, and purge pressure and purge flow were within a normal range at beginning of lt logs before purge cassette change was attempted. the root cause of the embolism cannot be determined as limited clinical information was provided and no issues with return pump was noted. h6 component code added e4 should have been left blank on manufacturer device report. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: section: potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
25078 | 1220648-2024-12697 | 19614526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614526&pc=OZD | 06/14/2024 | 08/07/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514435 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Detachment of Device or Device Component | [5 empty] | Foreign Body In Patient | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 70-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the cp migrated deep into the patient¿s left ventricle. the doctor attempted to pull it back, and it got stuck and looped in on itself. the doctor then attempted to straighten the catheter unsuccessfully and pulled the pump into the aorta to straight the pump, which was also unsuccessful. when trying to pull the pump back, the doctor broke the cp broke from the pigtail and inlet window, which left pieces in the patient¿s aortic arch. medical staff removed the rest of the cp and snared the pieces in the aortic arch. medical staff then implanted a new cp. no patient harm has been reported. | [5 empty] | The investigation for the positioning issue and product damage has been completed. clinical details states that the impella was inserted too deep in left ventricle while advancing repo sheath. impella stuck when attempting to remove leading to detachment of pigtail and inlet. both parts of the pump were returned and the cannula showed detachment location and product damage with evidence of stretching. the data logs show impella position unknown alarms and impella position in ventricle which both resolved. the root cause of the positioning issues was determined to be use related due to improper technique when advancing the repositioning sheath. the root cause of the product damage (detached inflow/outflow - great vessel) was determined to be use related as evidenced by clinical details g1-corrected mfr site name and address. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
25079 | 1220648-2024-12698 | 19614549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614549&pc=OZD | 06/15/2024 | 08/08/2024 |  | 06/26/2024 |  |  | 06/15/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495396 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 70-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported after the implant, the motor current was in the 200s with no motor current pulsatility. medical staff then removed the cp. the patient¿s left ventricle was completely akinetic with a large apical thrombus. the patient survived the event. | [5 empty] | The investigation for the thrombus and saturated flow has been completed. no product was returned for evaluation. data logs were reviewed and motor current decreased almost immediately after start-up with limited pulsatility. pump flow during this time was variable with multiple spikes in pump flow. suction alarms were triggered when motor current decreased. clinical details noted a very large thrombus in the lv under imaging and the lv was found to be akinetic. the root cause of the saturated flow was most likely due to biomaterial found in lv. added- h6 impact code 4629. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25080 | 1220648-2024-12692 | 19614569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614569&pc=OZD | 06/14/2024 | 08/07/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  | 10/31/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412916 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown age and gender with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported a cerebral hemorrhage was found. medical staff decided to proceed with a do-not-resuscitate but noted no hemodynamic problems with the patient. the patient was weaned off impella support before explantation. the relationship between the medical prognosis and impella support is unknown, and the patient survived the event. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25081 | 1220648-2024-12699 | 19614724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614724&pc=OZD | 06/15/2024 | 07/26/2024 |  | 06/26/2024 |  |  | 06/15/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232826 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Additional information received that the decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. | The complainant reported a 20-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) showed a purge disc not detected alarm. a nurse had been performing a routine purge cassette and bag change when the alarm occurred. the aic did not recognize both the old and new purge discs. medical staff then replaced the aic successfully. no patient harm has been reported, and the patient remains on impella support. |  |  |  |  | The investigation for automated impella controller (aic) - purge disc/flag issues has been completed. after reviewing the logs reported issue that purge disc couldn't be detected was confirmed. the purge disc flag was found to be broken (missing at the blue purge disc retainer). the root cause of the issue was a broken purge disc flag. b5 updated with additional information. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12699: d6a should have been blank. e4 should have been blank. g1 reporting contact fax number missing. h.6 code 4629 was reported incorrectly. new code was added to health effect - impact code. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25082 | 1220648-2024-12700 | 19614756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614756&pc=OZD | 06/06/2024 | 08/06/2024 |  | 06/26/2024 |  |  | 06/06/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469490 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 69-year-old female patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed ventricular tachycardia. the medical staff decided to explant the impella cp and continue percutaneous coronary intervention without impella support. once the impella cp was removed, the patient¿s heart rhythm returned to baseline. the patient remained stable. | [5 empty] | The investigation for ventricular tachycardia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issue could not be determined. b.7 revised as information was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-12700. g.3 date received by manufacturer was reported incorrectly on the initial submission of manufacturer device report number 1220648-2024-12700. the correct date is 09-jun-2024. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2024-12700 in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25083 | 1220648-2024-12701 | 19614968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19614968&pc=OZD | 08/01/2023 | 07/23/2024 |  | 06/26/2024 |  |  | 05/31/2024 | 06/26/2024 |  |  | 03/31/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347273 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿patient became anemic likely secondary to hemolysis from the impella, so the device was adjusted, hemolysis labs (ldh) were trended down, and patient received 1 unit of packed red cells. 8/2: ldh: 2162 8/3: 1768 8/4: 1414¿. the patient recovered with no sequelae. | [5 empty] | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 manufacturing site name and address. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12701. new code was added to component code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25084 | 1220648-2024-12653 | 19615167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19615167&pc=OZD | 06/04/2024 | 08/12/2024 |  | 06/26/2024 |  |  | 06/05/2024 | 06/26/2024 |  |  | 09/30/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024419317 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The complainant reported a 75-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the medical staff observed an abrupt decrease in flow when performing an introducer exchange. the medical staff then performed a fluoroscopy and observed a cannula kink near the outlet. the medical staff tried to rectify the kink with rotational maneuvers without success. the medical staff then decided to replace the impella cp with another without complications. the patient survived the event and remained stable. | [5 empty] | A.5 ethnicity and race are unknown. the impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. data was not returned for review. the impella was returned for evaluation. visual inspection found a kink on can about 15 mm from outflow cage can cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was kinked cannula due to improper technique while exchanging the introducer. d4-corrected model number f76/f8 should have been left blank in the initial report.
25085 | 1220648-2024-12652 | 19615189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19615189&pc=OZD | 05/29/2024 | 08/01/2024 |  | 06/26/2024 |  |  | 05/29/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025481761 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support there was suction and low impella flow alarms. the pump position was verified via echocardiogram. a clot was suspected to be in the canula and the pump was exchanged. the patient¿s overall condition continued to worsen, and the decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. | [5 empty] | The investigation of low pump flow has been completed. the returned complaint 5.5 pump visually inspected. no cannula kink observed. biomaterial observed around impeller and inflow and outflow cages. the cause of the low pump flow issue was biomaterial ingestion. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12652: e4 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25086 | 1220648-2024-12702 | 19615191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19615191&pc=OZD | 12/11/2023 | 07/02/2024 |  | 06/26/2024 |  |  | 06/06/2024 | 06/26/2024 |  |  | 08/31/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST EU |  |  | ABIOMED, INC. |  |  | US |  | 2024391147 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The complainant reported that during support for a 56-year-old male patient, the impella 5.5 stopped due to clot formation. the clot formation was due to improper heparinization of the patient resulting in too low activated partial thromboplastin time. the pump was explanted. there was no patient harm reported. | [5 empty] | Literature citation: przybylski, r., kosowski, m., bochenek, m., reczuch, k., barteczko-grajek, b., kupiec, a., zakliczynski, m., zimoch, w., blaziak, m., & kuliczkowski, w. (2024). impella 5.5 and mitral transcatheter edge-to-edge repair as a bridge to heart transplantation in a patient with cardiogenic shock. polish heart journal, 82(3), 341¿342. https://doi.org/10.33963/v.kp.98041 g.2 revised as literature was inadvertently omitted from the previously submitted report. h.10 revised as literature citation was inadvertently omitted from the previously submitted report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
25087 | 1220648-2024-12703 | 19615515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19615515&pc=OZD | 06/06/2024 | 08/02/2024 |  | 06/26/2024 |  |  | 06/06/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025482694 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that during support, the patient was taken to the operating room to remove an access site hematoma and received an unspecified amount of blood bags as treatment. the patient remained on support, despite the event. the patient was reported as stable. | [5 empty] | The investigation of access site bleeding/hematoma has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient condition since the patient had coagulation disorder. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12703. d4 model number, h6 health effect - impact code 4624 was not included. | A.5 ethnicity and race are unknown. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25088 | 1220648-2024-12704 | 19615830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19615830&pc=OZD | 06/05/2024 | 08/15/2024 |  | 06/26/2024 |  |  | 06/07/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025482037 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (age unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had signs of hemolysis. the lactate dehydrogenase levels had risen. the impella cp was explanted. the patient is stable. | [5 empty] | The investigation for the hemolysis has been completed. product and data logs were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. added- h6 component code 925. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25089 | 1220648-2024-12707 | 19616649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19616649&pc=PYX | 06/04/2024 | 08/02/2024 |  | 06/26/2024 |  |  | 06/04/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025485073 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with a heart transplant rejection was implanted with an impella rp device for mechanical circulatory support. after implant, the patient experienced an access site bleed. an extra purse string suture was added to the access site. additionally, while on support, the patient went into ventricular fibrillation and was shocked 4 times. echocardiogram was used to check the position of the pump. the patient remained on impella rp support. | [5 empty] | The investigation for access site bleed and ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and vf could not be determined. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿. impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ d.1 brand name and d.4 revised model number were entered incorrectly on the initial manufacturer device report number and were revised. h.6 codes 4627 and 4755 were entered incorrectly on the initial manufacturer device report number. new codes have been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smart assist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smart assist system catheter: use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smart assist system with automated impella controller: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25090 | 1220648-2024-12705 | 19616692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19616692&pc=OZD | 06/04/2024 | 07/25/2024 |  | 06/26/2024 |  |  | 06/04/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462191 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and next post implant the patient developed oozing at the access site. the site was angle matched, pressure was applied, heparin was withheld, and two units of blood were transfused to manage the condition. support with the implanted pump was maintained. further, it was noted the patient was in multi system organ failure, shock liver and acute kidney injury, and continued to be very vasoplegic. care was withdrawn by the family, and the patient expired. medical safey review was completed and noted there is not enough information to associate use of the device with the patient's demise. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient condition. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12705: h.6 code 4755 was reported incorrectly. new code was added to component code.
25091 | 1220648-2024-12706 | 19616964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19616964&pc=OZD | 06/04/2024 | 07/24/2024 |  | 06/26/2024 |  |  | 06/04/2024 | 06/26/2024 |  |  | 05/31/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELL CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352092 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was in profound cardiogenic shock, acidotic and declining. while on support, it was noted the patient experienced persistent hemolysis. repositioning of the pump was completed, but the condition continued. one unit of blood was given to help treat the condition and the pump was eventually explanted later that same day. further information noted the patient expired. medical safety review was completed and noted there is not enough information to associate use of the device with the patient's demise. | [5 empty] | The investigation for hemolysis has been completed. the product was not returned for review. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12706: h.6 code 4755 was reported incorrectly. new code was added to component code. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25092 | 1220648-2024-12708 | 19617431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19617431&pc=OZD | 06/09/2024 | 08/16/2024 |  | 06/26/2024 |  |  | 06/09/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489742 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported attempt to insert an impella cp pump was aborted due to the patient's severe aortic stenosis. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump was aborted due to severe aortic valve stenosis. instructions for use as follows. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12708. g1 manufacturing site name and address.
25093 | 1220648-2024-12709 | 19617731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19617731&pc=OZD | 01/21/2024 | 07/18/2024 |  | 06/26/2024 |  |  | 06/05/2024 | 06/26/2024 |  |  | 11/30/2025 |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426068 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ventricular tachycardia with a "probable" relationship to the impella device used: ¿on (b)(6) 2024 after pt returned to icu from cath lab, pt had 2 separate incidents of vt (ventricular tachycardia) without a pulse (approximately. during each incident pt was given 1 shock at 200 joules with immediate rosc (return of spontaneous circulation). stat echo obtained which showed impella had become displaced. impella repositioned under echo guidance with return of nsr¿. the patient recovered with no sequelae. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured premature ventricular contractions with a "probable" relationship to the impella device used: ¿on 23jan2024 pt developed new onset of pvcs (premature ventricular contraction). a bedside tte (transthoracic echocardiogram) showed that the impella had become displaced overnight. impella repositioned with return to nsr. (normal sinus rythm)¿. the patient recovered with no sequelae. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured premature ventricular contractions with a "probable" relationship to the impella device used: ¿on 24jan2024 pt developed pvcs for a second time. a bedside tte showed that the impella had become displaced overnight. impella repositioned with return to nsr.¿. the patient recovered with no sequelae. | [5 empty] | The investigation for vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source added. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25094 | 1220648-2024-12710 | 19618036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19618036&pc=OZD | 04/03/2024 | 07/30/2024 |  | 06/26/2024 |  |  | 06/11/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ischemic leg with a "definite" relationship to the impella device used: ¿there was significant concern for ischemic leg, decision was made to transition to ecmo (extracorporeal membrane oxygenation) and remove impella to protect leg¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. g1-corrected mfr site name and address h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25095 | 1220648-2024-12711 | 19618163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19618163&pc=OZD | 03/11/2024 | 07/31/2024 |  | 06/26/2024 |  |  | 06/11/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438128 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening right heart failure with a "probable" relationship to the impella device used: ¿patient continued to feel poorly following impella implant, repeat rhc (right heart catheterization) completed on (b)(6) 2024 showed a decreased papi falling from 1.9 to 1.1 and an increase in rap from 11 to 19 mmhg.¿. the patient recovered with no sequelae. | [5 empty] | The investigation has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia/ventricular fibrillation was not determined. b.7 information was added as it was inadvertently omitted on the initial submission of manufacturer device report number 1220648-2024-12711. h.6 codes 4611, 4755 and 4114 were entered incorrectly on the initial manufacturer device report number 1220648-2024-12711. new codes have been added. h.10 additional manufacturer narrative on the previously submitted initial manufacturer device report number 1220648-2024-12711 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted initial manufacturer device report number 1220648-2024-12711 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25096 | 1220648-2024-12712 | 19618529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19618529&pc=OZD | 06/12/2024 | 08/16/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506028 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown age and gender with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that, during impella support, medical staff noted reduced perfusion in the patient¿s leg. medical staff had the patient on a combination of extra-corporeal membrane oxygenation (ecmo) and impella support, where the anterograde flow from ecmo was connected to the leg, which resolved the patient¿s reported issue. medical staff noted no bleeding and the leg well-perfused. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12712. g1 manufacturing site name and address.
25097 | 1220648-2024-12713 | 19618654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19618654&pc=OZD | 06/14/2024 | 03/19/2025 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1160647 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Pressure Problem | Application Program Freezes, Becomes Nonfunctional | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 54-year-old patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that, during support, the physician reported the associated automated impella controller (aic) showed an air in purge system alarm, which was selected and started in the purge menu. medical staff then reported the aic's screen froze, no selection could be made with soft buttons, and there was no touch response. medical staff then exchanged the aic, which resolved the issue. no patient harm has been reported, and the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) purge issue ¿ other has been completed. after reviewing the logs, the reported air in purge system alarm and unresponsive softkey issues were confirmed. there was one instance of air in purge system alarm observed in the logs from the reported complaint. there were also numerous keyboard-(kbd) error entries observed in the logs. the console was tested and the reported issues could not be reproduced. ribbon cable between the kbd and 4-in-1 was inspected but there were no abnormalities found. conclusion: the air in purge system issue was likely use related. prior to this alarm being triggered, it was observed that the purge cassette and disc were removed from the aic. when the purge cassette and disc is removed from the aic, the purge pressure can¿t increase because the purge motor can¿t turn the purge cassette piston which can lead to the air in purge system alarm. the unresponsive soft key issue could not be determined due to the inability to be reproduced but was likely caused by a random loss of communication between the kbd and 4-in-1. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-12713in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-12713 in accordance with updated procedures g1 manufacturing site name and address was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2024-12713.
25098 | 1220648-2024-12714 | 19618836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19618836&pc=OZD | 06/14/2024 | 08/01/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480657 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that, when the patient arrived at the intensive care unit, thrashing and access site bleeding occurred. medical staff held manual pressure, matched angles, forwarded tension sutures, and applied topical tranexamic acid-soaked gauze to obtain hemostasis successfully. medical staff also reported that bleeding occurred again after the patient had been turned. medical staff then managed the patient with manual pressure and hemostatic gauze dressing. the patient remained stable until potential escalation to an impella 5.5 and survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): limb ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the bleeding was controlled, and hemostasis was achieved by applying manual pressure and forward tension sutures. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12714. new code was added to component code.
25099 | 1220648-2024-12715 | 19619425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19619425&pc=OZD | 06/14/2024 | 09/19/2024 |  | 06/26/2024 |  |  | 06/14/2024 | 06/26/2024 |  |  |  |  | 06/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124467 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 56-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the associated automated impella controller (aic) had a controller error alarm soon after powering on. medical staff exchanged the aic to resolve the issue. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) controller alarm yellow has been completed. review of data logs confirmed the reported complaint. there were multiple occurrences, across multiple boots, of a controller errors due to a defective optical sensor lamp. immediately preceding each controller error, the console performed a reset of the optical bench in an attempt to resolve recognized malfunctions, including contrast values that were outside of specification. the console was returned for evaluation. upon evaluation, the optical bench was tested, and the 5s parameters were found to be within specification. the reported complaint could not be reproduced. however, physical inspection of the console also revealed that the ribbon cable connecting the 4-in-1 to the optical bench was not fully inserted into the ribbon cable header on the optical bench carrier pca, causing the ribbon cable to become partially disconnected when the console is exposed to vibration or movement. the reported complaint was determined to potentially have been caused by a defective optical bench, specifically potentially an intermittent failure of the lamp. added- d9 device return date h5 changed no. h8 changed to reuse. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
25100 | 1220648-2024-12648 | 19604535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19604535&pc=OZD | 05/28/2024 | 07/25/2024 |  | 06/25/2024 |  |  | 05/28/2024 | 06/25/2024 |  |  | 10/31/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support during high risk percutaneous coronary intervention (hrpci). during support, the patient experienced pulsatile oozing at the groin site. additionally, there were suction alarms present with impella dropping to p-1. volume and blood products were administered. the patient experienced bleeding from the chest incision during hrpci. the bleeding continued and the patient eventually expired. abiomed does not have information to associate use of the device with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and low pump flow has been completed. the impella was not returned for evaluation. data logs confirmed the low flow when pump started and remained the rest of the case that triggered auto suction response and suction alarms. based on clinical description, the root cause of access site bleeding was most likely due to patient condition. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
25101 | 1220648-2024-12650 | 19604695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19604695&pc=OZD | 05/28/2024 | 07/31/2024 |  | 06/25/2024 |  |  | 05/28/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494493 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circlulatory support. during support high purge pressure was encountered and tissue plasminogen activator (tpa) was added to the purge to address the issue with no harm to the patient. further information noted the patient was later made "do not resuscitate" and passed away from their existing condition. medical safety review was completed and noted there is not enough information to associate use of the impella device with the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: troubleshooting the purge system ¿if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen." | The investigation for high purge pressure has been completed. data logs were reviewed, and lt logs of the full case showed a slow increase in purge pressure with "high purge pressure" alarms being triggered three hours into support, and high purge pressure was able to be resolved six hours into support. real time logs from the beginning of case showed a motor current spike with simultaneous speed deviation and drop in pump flow as the pump was ramping up. the purge pressure was observed to steadily increase after motor current spike with purge flow slowly dropping. clinical details noted that tissue plasminogen activator protocol was initiated when "high purge pressure" alarms was triggered in attempt to resolve the alarms. the root cause of the high purge pressure was most likely due to ingestion of biomaterial. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12650 as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12650 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
25102 | 1220648-2024-12649 | 19604702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19604702&pc=OZD | 05/28/2024 | 08/08/2024 |  | 06/25/2024 |  |  | 05/28/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453289 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention |  |  |  |  | VA ECMO | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient postcardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. upon insertion of the impella 5.5, the cannula perforated the patient's left ventricle. the physician pulled the pump back and repaired the perforation. support continued. the patient had a washout of the chest while on support and the team found a thrombus in the patient's atrium. care was withdrawn after 35.40 hours of support and the patient expired. abiomed does not enough information to exclude use of the device from the outcome. | [5 empty] | The investigation for ventricle perforation and thrombus have been completed. the impella device was not returned for evaluation. the cause of the ventricle perforation issue was not determined, since the product was not returned and due to insufficient clinical information. as the necessary information was not provided, the root cause of the thrombus was not determined. e.4: should have been left blank on the initial manufacturer device report number. as it is unknown, if the initial reporter, also sent the report to the fda. h.6: code 4755, was reported incorrectly on the initial manufacturer device report. new codes was added to component codes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25103 | 1220648-2024-12651 | 19605094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605094&pc=OZD | 05/28/2024 | 08/12/2024 |  | 06/25/2024 |  |  | 05/29/2024 | 06/25/2024 |  |  | 06/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024416666 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | 505 | Foreign | Health Professional | Company representation | Distributor |  |  | 24 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient with biventricular failure and on impella cp support had a clot in the left ventricle. the file is noted: patient met accident at work, inhaled poisonous fumes. heart failure due to gas poisoning. the physician was planning to do left ventricle venting but the family didn¿t want to proceed with the surgery. the patient expired. abiomed does not have enough information to associate use of the device with the outcome. | [5 empty] | The investigation thrombus has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. b.2 revised as the incorrect field was selected on the initial manufacturer device report number. e1 revised as name prefix/title was inadvertently omitted from initial manufacturer device report number. e.4 should have been left blank in initial manufacturer device report number as it is unknown the initial reporter also sent the report to the fda. h.6 code 4755 was reported incorrectly on initial manufacturer device report number. a new code was added to component codes. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25104 | 1220648-2024-12654 | 19605544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605544&pc=OZD | 05/31/2024 | 08/08/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  | 04/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337948 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 15 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was experiencing torsade de pointes type ventricular tachycardia which caused hemodynamic failure, and the impella cp device was introduced. the implanted impella cp pump experienced a positioning issue during patient support. it was noted an "impella position in aorta" alarm appeared, and the pigtail was found in the left ventricle. repositioning proved to be difficult, and the decision was made to explant the pump. an intra-aortic balloon pump was subsequently placed for ongoing support. there was no patient harm resulting from the positioning issue. information subsequently received noted the patient expired. however, medical safety review of the information noted there is not enough information to associate use of the impella device with the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor data logs were returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of positioning issue was most likely patient condition related.
25105 | 1035166-2024-00070 | 19605647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605647&pc=DYB | 05/23/2024 | 05/28/2024 |  | 06/25/2024 |  |  | 05/28/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3046 |  |  | OSCOR.INC |  |  | US |  |  |  | 0052-3046 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 73 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that hemostatic valve failed immediately after insertion causing bleed back requiring opening of a new sheath for impella placement. there was blood loss reported but no medical or surgical intervention required. patient outcome is stable. device discarded by customer .no additional information reported. | [5 empty] | No product was provided for analysis, nor were any pictures provided by the customer. customer provided multiple lot and could not confirm which lot was used during procedure and had malfunction. there was blood loss reported, however no surgical or medial intervention required. the procedure was completed successfully. dhr review is not possible as lot number is unknown. a review of the manufacturing processes indicates the production operators are instructed to 100% visually inspect for any obvious defect prior to shipment. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. no corrections or corrective actions will be taken. device was not returned for analysis. this complaint is non-verifiable. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
25106 | 1220648-2024-12656 | 19605709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605709&pc=OZD | 05/31/2024 | 07/24/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  | 11/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424861 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced limb ischemia. the impella extremity was without a pulse and was cold and dusky. a fem-fem bypass was performed and the pulse was able to be dopplered after. additionally, there was bleeding noted from multiple sites including the impella access site. a femoral hematoma was noted at the impella access site as well. one unit of blood products were provided to the patient. the patient's family decided to withdraw care. the impella cp was left in place and turned off. the patient expired. abiomed does not have enough information to associate use of the device with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and limb ischemia could not be determined. f6/f8 should have been left blank in the initial report.
25107 | 1220648-2024-12655 | 19605712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605712&pc=OZD | 06/05/2024 | 07/30/2024 |  | 06/25/2024 |  |  | 06/05/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was taken to the catheterization lab as the impella cp was difficult to reposition under transthoracic echocardiogram (tte). there was concern for right ventricle dysfunction. using fluoroscopy and tte, impella pigtail appeared to be caught under a papillary muscle. several attempts to reposition the impella cp were made. it was further reported that the patient had signs of hemolysis. labs showing lactate dehydrogenase over 1700 and plasma free hemoglobin greater than 40 for two days. the patient was given a blood transfusion. the plan is for high risk percutaneous coronary intervention with a goal of weaning and explanting the device post procedure. the patient is stable. | [5 empty] | The investigation for the positioning issue and hemolysis has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started and ran for about 1.5 hours, low flow occurred that triggered auto suction response and suction alarms. no positioning issue was found on the log. this event remained for 13 hours then resolved. after that, pump ran without issue to the rest of the case. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely pump was in improper positioning. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. f6/f8 should have been left blank in the initial report. | A.5 race is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25108 | 1220648-2024-12657 | 19605734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605734&pc=PYX | 06/03/2024 | 08/06/2024 |  | 06/25/2024 |  |  | 06/03/2024 | 06/25/2024 |  |  | 11/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2024432999 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The user facility reported a 70-year-old female patient postoperative cardiothoracic surgery was to be implanted with an impella rp device for mechanical circulatory support. it was reported during the placement the rp would not track into the right ventricle and would not continue into the inferior vena cava (ivc). placement was attempted three more times before switching to a lunderquist 0.035 wire to advance impella successfully into the left pulmonary artery (lpa). the impella was ramped up to p7 without complication. the rp continued to point into the ivc. the rp was removed and placed a total of five times with an estimated blood loss of 1500cc from the impella insertion and removal. the patient required a transfusion. the patient¿s overall condition continued to worsen, and the decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. | [5 empty] | The investigation for access site bleeding and delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was determined to be use issue because by switching to a lunderquist 0.035 wire they were able to advance the device successfully into left pulmonary artery. hence, the root cause was determined to be a use issue. the root cause of access site bleeding was determined to be a use issue because of the difficulty in delivering the pump. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp flex with smartassist system with automated impella controller section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ d.1 brand name and d.4 model number revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12657. e.1 revised as last name was accidently placed in the middle name field on the initial manufacturer device report number 1220648-2024-12657. the middle name was left blank as it is unknown. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12657 and should not have been. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12657 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open. make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: warnings and cautions: do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
25109 | 1220648-2024-12658 | 19605895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19605895&pc=OZD | 06/06/2024 | 07/18/2024 |  | 06/25/2024 |  |  | 06/06/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464840 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy implanted with an impella 5.5 device for mechanical circulatory support experienced an access site complication. the patient was admitted in with cardiogenic shock. the impella 5.5 pump was placed, but patient developed a hematoma. the hematoma was treated by surgical intervention. the site was evacuated, a bleeding vessel was closed by clip, and the jp drain had drainage. the pump remained on for several days to support the patient. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | A.5 race in unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25110 | 1220648-2024-12660 | 19606017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606017&pc=OZD | 06/08/2024 | 08/01/2024 |  | 06/25/2024 |  |  | 06/08/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025465398 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/08/2024 |  |  |  |  | The user facility reported a 62-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was causing signs of hemolysis. the pump was positioned deep. the pump was repositioned, but the condition did not improve, and family made choice to withdraw care and patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. | [5 empty] | Data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. f6/f8-should have been left blank. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12660. new code was added to component code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
25111 | 1220648-2024-12661 | 19606047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606047&pc=OZD | 06/11/2024 | 08/05/2024 |  | 06/25/2024 |  |  | 06/11/2024 | 06/25/2024 |  |  | 07/31/2024 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there were positioning issues. the staff were unable to reposition the pump and it was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12661.
25112 | 1220648-2024-12662 | 19606137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606137&pc=OZD | 04/20/2021 | 11/19/2025 |  | 06/25/2024 |  |  | 04/20/2021 | 06/25/2024 |  |  | 03/31/2022 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1474641 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Sepsis | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | EXTRACORPOREAL MEMBRANE OXYGENATION | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2021 |  |  |  |  | The user facility reported a male patient of unknown age, race, and ethnicity was implanted with an impella cp device for mechanical circulatory support to unload the left ventricle. it was reported that while on support, the purge flow dropped, and pressure was increased over the last few hours. it was resolved by lowering the glucose concentration to 2.5. it was reported that the patient had ongoing hemolysis. the patient¿s conditioned worsened due to sepsis and expired after 27,877.31 hours on impella support. | [5 empty] | A.5 ethnicity and race are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instruction for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email updated.
25113 | 1220648-2024-12663 | 19606224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606224&pc=OZD | 05/28/2024 | 08/01/2024 |  | 06/25/2024 |  |  | 05/28/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025505584 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (race and age unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient bled from the puncture site in the icu with loss of hemoglobin. the patient¿s partial thromboplastin time was initially 60, compress was placed under the catheter and fixation suture was also fixed in the direction of the patient's head. the patient was treated with a sandbag and the bleeding stopped. six days later, the patient expired on support. abiomed does not have enough information to associate use of the device with the outcome. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- b5 mso review g1-corrected mfr site name and address.
25114 | 1220648-2024-12664 | 19606244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606244&pc=OZD | 05/28/2024 | 07/16/2024 |  | 06/25/2024 |  |  | 05/28/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a pulseless, cold, mottled right leg. originally pulses were present on admission to icu. this complication led to the impella cp being removed. the patient decompensated after explant and ultimately expired. abiomed does not have enough information to associate use of the device with the outcome. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. e4 should have been left blank in the initial report. g1 reporting contact fax number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿
25115 | 1220648-2024-12665 | 19606267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606267&pc=OZD | 05/31/2024 | 08/01/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488820 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, hemolysis was diagnosed. red blood cell concentrate was transfused for a second time, but the issue did not resolve. it was further reported that the pump was not stable in a good left-ventricular position; it slipped when the patient was repositioned. further information received noted the patient expired on support. the physician, however, stated due to the severity of the illness, the patient passed, and patient¿s death is not impella related. the physician explained at some point there was a hemolysis present due to a brief malposition of the pump and otherwise the team had no problems with the impella and felt well supported. additionally, medical safety reviewed was completed and noted there is not enough information to associate use of the device with the outcome. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12665: e4 should have been blank. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25116 | 1220648-2024-12667 | 19606287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606287&pc=OZD | 06/06/2024 | 07/24/2024 |  | 06/25/2024 |  |  | 06/06/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479292 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The complainant reported a 53-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed hemolysis following an extracorporeal membrane oxygenation decannulation. the medical staff gave the patient volume and a high dose of systemic heparin. the afterload was then reduced. the patient remains stable. | [5 empty] | The impella device has not been received from the customer. therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. data logs showed multiple suction alarms were triggered on the day the ecmo was decannulated and cvp was 6-7. the suction alarms were resolved next day. no other issue was found on the log. trending placement signal observed with likely deteriorating patient condition. the root cause of hemolysis was most likely patient condition related per log analysis and clinical review. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12667: f6/f8-should have been left blank.
25117 | 1220648-2024-12669 | 19606296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606296&pc=OZD | 06/17/2024 | 08/05/2024 |  | 06/25/2024 |  |  | 06/18/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490052 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when they had access site bleeding. the nurse noticed discoloration and firmness of abdomen superior to impella site. the patient was rolled over and after, there was external oozing from site. a blood transfusion was performed. additionally, the patient was in atrial fibrillation (af). the impella was repositioned and pulled back to 3.8 cm. | [5 empty] | The pump was returned for investigation and no problem was found with the pump. access site bleeding: the root cause of access site bleeding is determined to be patient movement because the bleeding started upon rolling the patient. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12669: d4 model number e4 should have been left blank g1 manufacturing site name and address g2 report source: foreign was inadvertently selected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25118 | 1220648-2024-12666 | 19606316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606316&pc=OZD | 06/04/2024 | 07/24/2024 |  | 06/25/2024 |  |  | 06/04/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488478 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post percutaneous coronary intervention (pci). the pci was performed and cp support was successful until the groin site was seen to have a hematoma. the hematoma was treated by femostop. the patient was becoming hypotensive and so the team consulted with vascular surgery. the pump was explanted, perclosed, and new pump placed via contralateral leg. the patient's care was withdrawn and the patient expired. abiomed's medical safety office review determined abiomed does not have enough information to associate use of the device with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
25119 | 1220648-2024-12670 | 19606319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606319&pc=OZD | 06/16/2024 | 12/03/2025 |  | 06/25/2024 |  |  | 06/19/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1764867 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | IMPELLA 5.5 | [17 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 69-year-old male patient in in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day 17 of support the automated impella controller (aic) failed and was replaced. the replacement of the aic was done without harm or injury and the 5.5 pump remains on for support. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. d9 return date has been added. updated h3 to device evaluation anticipated. h6 type of investigation code added. | Data analysis: before the complaint date, the last case performed with console ic11031 was from may 30 to jun 15, 2024, using pump sn (b)(6). there was no occurrence of ¿controller error¿ in the logs, but the console shut down unexpectedly on june 15, 2024. the pump was then moved to the second console, ic1605. there were no occurrences of ¿controller error¿ in the second console¿s log either. after the console ic11031 was manually turned on aug 21, 2024, it displayed ¿unexpected controller shutdown - alarm,¿ event #603. note: the unexpected controller shutdown was likely misreported as ¿controller error.¿ device analysis: this console was tested after arriving at fs. unable to reproduce either the ¿controller error¿ or the ¿unexpected controller shutdown¿ alarms. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). root cause: there was no evidence in the logs to confirm the ¿controller error.¿ the unexpected controller shutdown was likely misreported as ¿controller error.¿ the root cause of the unexpected controller shutdown could not be determined due to the inability to reproduce it. most likely, either a 4-in-1 malfunction or a pbm malfunction caused the unexpected shutdown. d6a and d6b should have been left blank. d10 concomitant medical products and therapy dates updated. e4 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25120 | 1220648-2024-12668 | 19606325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606325&pc=OZD | 06/13/2024 | 08/20/2024 |  | 06/25/2024 |  |  | 06/13/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025488260 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during patient support, the impella cp had "sensor failure". the device was explanted as a result. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical signal issue has been completed. an impella cp was inserted through the femoral artery with no complications pre-pci. during the case, there was a sensor failure, and the pump was removed as a result 6 hours after implant. no other clinical details were given. product and logs were not returned. the root cause of the optical signal issue was not determined since product/logs were not returned and insufficient clinical evidence was provided.
25121 | 1220648-2024-12671 | 19606333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606333&pc=OZD | 06/20/2024 | 08/12/2024 |  | 06/25/2024 |  |  | 06/20/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025490261 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support and there was access site bleeding noted. a light femostop was placed. additionally, there were high purge pressure alarms noted. the purge cassette was replaced to no avail. eventually, the pump stopped. the impella was removed. | It was further reported that the pump stopped and while preparing to exchange it out, the patient coded and expired. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure and the pump stop has been completed. the pump was not returned for investigation. data logs show a long purge bag change time of about 3 minutes, followed by a gradual rise in purge pressure and a pump stop with saturated pump flow at the end of the case run. the pump failed to restart, with several motor current spike alarms. heparin was not added in the purge solution. the cause of the pump stop issue was related to use involving a long purge bag change time, which led to clot formation and a rise in purge pressure. according to the impella ifu, the recommended time for purge cassette/bag changes is less than 90 seconds. b1-selected adverse events b2-selected death and added date of patient death. h1 type of reportable event changed to death. h6 impact code changed to 1802.
25122 | 1220648-2024-12672 | 19606356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606356&pc=OZD | 06/06/2024 | 07/31/2024 |  | 06/25/2024 |  |  | 06/06/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma requiring blood transfusions. a total of two units of red blood cells, two units of fresh frozen plasma, one unit of platelets, and one unit of cryo were given. a femostop was placed. the patient remained on impella cp support for two additional days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12672 and should not have been. h.6 code 4641 was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12672 and has been added. code 925 was added since the initial manufacturer device report number 1220648-2024-12672 was submitted in accordance with updated procedures.
25123 | 1220648-2024-12673 | 19606440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19606440&pc=OZD | 06/20/2024 | 08/14/2024 |  | 06/25/2024 |  |  | 06/20/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025491247 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old female patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported the 5.5 pump was prepped but failed all priming and so the pump was replaced. the new 5.5 pump was delivered, and support initiated. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the priming issue has been completed. pump was returned for investigation. purge cassette was not returned. when connected to the aic the pump went to the case start wizard step 1. the pump was then successful primed and was not able to reproduce the failure. as per the clinical information provided, the error code received was "prime not completed, check purge system tubing for kinks" and since the pump was successfully primed during the investigation, suggests that it was a user issue. therefore, the root cause of the priming issue was most likely use related to improper technique. added- d9 device return date, h6 impact code 4629.
25124 | 1220648-2024-12674 | 19607074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19607074&pc=OZD | 06/06/2024 | 08/20/2024 |  | 06/25/2024 |  |  | 06/06/2024 | 06/25/2024 |  |  | 11/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421786 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | *the complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the patient experienced massive groin bleeding externally and internally. the source of the bleed was noted to be at the puncture site and connection between blue suture hub and white connection area. the patient expired while on support due to fulminate cardiogenic shock. per the medical safety office review: abiomed does not have enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system; section: general patient care considerations; ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related. bleeding at multiple sites was noted as per the clinical description provided. d.4 unique identifier (udi) #, d.6a and d.6b, e.1 facility name, and g.1 revised manufacturing site name and address section were revised as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-12674. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12674 and should not have been. h.6 codes 1762 and 4655 were reported incorrectly on initial manufacturer device report number 1220648-2024-12674. code 925 was added.
25125 | 1220648-2024-12675 | 19607437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19607437&pc=OZD | 06/14/2024 | 08/06/2024 |  | 06/25/2024 |  |  | 06/14/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487386 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella 5.5 placement and upon starting case air bubble was noted coming in. there was a concern if there is way bubble came from the impella. the 5.5 was primed more than 15 minutes before placement and visual was noted that the purge was coming out prior. the patient is stable. | [5 empty] | The investigation for the embolism has been completed. pump was not returned for investigation. clinical details mention that the pump was primed outside the body for 15 minutes before insertion suggesting good insertion practice was followed. it is not possible to investigate the complaint without the product back. therefore, the root cause of the embolism issue was not determined since product was not returned for investigation and limited clinical details were provided. f6/f8 should have been left blank in the initial report. h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
25126 | 1220648-2024-12676 | 19607525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19607525&pc=OZD | 11/06/2023 | 08/01/2024 |  | 06/25/2024 |  |  | 05/30/2024 | 06/25/2024 |  |  | 09/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399767 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured mediastinal bleeding from impella surgery with a "possible" relationship to the impella device used: ¿after impella implanted, due to patient's history of thrombocytopenia (setting ongoing coagulopathy), there was ongoing bleeding at the mediastinal site. sternal wires applied to stop bleeding along with wound vac was placed. h/h baseline is unknown. h/h is 11.1 and 31.8. h/h peak is 12.2/34.3. acute care surgery was following. patient continued to have wound vac post discharge¿. it was reported that the patient/event has not recovered/not resolved. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening ventricular tachycardia with a "possible" relationship to the impella device used: ¿on evening (b)(6) 2023, patient had increased ventricular ectopy on telemetry. patient had epicardial pacer on but undersensing and not capturing. while examining external pacer wires, patient went into monomorphic vt (ventricular tachycardia) (220s). patient received 6 shocks 200j unipolar. brief return nsr between shocks followed by another vt (approx 15 mins). amiodarone and lidocaine boluses given prior to start of continuous drips. bedside echo confirmed impella was positioned correctly. settings for external pacer changed. patient didn't have ct (computed tomography) after. patient continued to be monitored for vt post discharge in inpatient skilled facility¿. it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation for access site bleeding and ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be patient condition because of the patient's history of thrombocytopenia where the patient has abnormally low levels of platelets in their blood. as the necessary information was not provided, the root cause of the vt was not determined. b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12676. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12676 and should not have been. h.6 code 4755 was entered incorrectly on the initial submission of manufacturer device report number 1220648-2024-12676. a new component code has been added. code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12676. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-12676 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12676 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25127 | 1220648-2024-12677 | 19607651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19607651&pc=OZD | 04/23/2024 | 07/30/2024 |  | 06/25/2024 |  |  | 05/30/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured left femoral artery occlusion with a "definite" relationship to the impella device used: ¿left femoral cfa (common femoral artery¿¿) impella was weaned then removed using manta which lead to left common femoral artery occlusion requiring ballooning using 6x60mm armada balloon which lead to left femoral artery bleeding requiring prolonged balloon inflation using 6x60mm armada. final peripheral angiogram was performed from contralateral r cfa with no evidence of bleed/contrast extravasation.¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as information was inadvertently omitted on the initial submission of manufacturer device report number 1220648-2024-12677. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12677 and should not have been. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer device report number 1220648-2024-12677. h.6 code 4641, 4114 and 4755 were entered incorrectly on the initial submission of manufacturer device report number 1220648-2024-12677. new codes were added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12677 in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-12677 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25128 | 1220648-2024-12678 | 19607696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19607696&pc=OZD | 05/02/2024 | 07/12/2024 |  | 06/25/2024 |  |  | 05/30/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494996 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Perforation of Vessels | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured retroperitoneal hematoma with a "definite" relationship to the impella device used. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured perforation in right external iliac artery with a "definite" relationship to the impella device used. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hypotension requiring levophed with a "definite" relationship to the impella device used. | [5 empty] | The investigation for the access site bleed, vascular damage, and stroke has been completed. no product was returned for investigation. the root cause of the access site bleeding major was not determined since product was not returned and there is a lack of details. the root cause of the vascular damager peripheral vessel issue was not determined since product was not returned and there is a lack of clinical details. the root cause of the stroke was not determined since product was not returned and there is a lack of clinical details. d4-updated udi number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25129 | 1220648-2024-12679 | 19608253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19608253&pc=OZD | 12/15/2023 | 07/30/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  | 09/30/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401508 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿elevated ldh in setting of impella along with plasma hgb. peak ldh was 1003 and plasma hgb was 79.0. patient also had new hemogloburia. team wasn't sure if hemolysis was due to impella or absorption of hematoma. tee on (b)(6) 2023 r/o thrombus. on (b)(6) 2023, ldh and plasma hgb were high but hgb, u/a and total bil were stable. labs were monitored throughout this admission¿. the patient recovered with no sequelae.we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured cervical internal carotid arteries (ica) occlusion with a "possible" relationship to the impella device used: ¿patient c/o of blurred vision in r eye noticed several days in hospital with some dizziness. neuro consulted. cta head and neck on (b)(6) 2023 showed l ica above level of bifurcation with reconstitution at supraclinoid level from contralateral aca. patient on iv bilval drip. vascular surgery confirmed ica is chronic. optho consult confirmed no evidence of retinal artery occlusion. patient will follow up with team outpatient post discharge. (b)(6) 2023 ct scan confirmed no acute infarct/mass effect or developing infarct¿. it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation for the limb ischemia and hemolysis has been completed. product was not returned for review. data logs were reviewed which showed pump ran without issue during support. there were suction alarms triggered after pump started and at the end of the case but resolved quickly. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. f6/f8 should have been left blank in the initial report. h6 component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25130 | 1220648-2024-12680 | 19608419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19608419&pc=OZD | 04/25/2024 | 07/23/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  |  |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454011 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured marked right femoral hematoma with a "probable" relationship to the impella device used: ¿4/24/24: patient arrived to p5 hds hypertensive with iabp in place. not in shock. home meds re-ordered. noted to have r femoral hematoma which was marked. no le numbness/weakness and doppler pulses intact. 1 day later hematoma unchanged, firm. observed overnight in cicu without need for pressors or other invasive needs, blood pressure stable. removal of iabp without complications. 4/26/24: r groin hematoma firm, marking unchanged from yesterday, pulses doppler strong multiphasic moves extremity and warm denies numbness¿. the patient recovered with no sequelae. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12680. new code was added to component code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25131 | 1220648-2024-12681 | 19608685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19608685&pc=OZD | 10/04/2023 | 10/14/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  | 09/30/2024 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured thrombotic vascular complication with a "probable" relationship to the impella device used: ¿4/7 thrombectomy of the right common femoral was necessary due to thrombosis from the impella placement.¿. the patient recovered with no sequelae. we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured limb ischemia with a "probable" relationship to the impella device used: ¿the patient had an attempted impella on 4/4 through the right femoral artery, which was unsuccessful, and which caused ischemia of the right lower extremity, necessitating a reperfusion cannula in the right superficial femoral artery.¿. the patient recovered with sequelae. | It was also noted that the patient endured hemolysis, which was noted as ongoing through discharge. action taken to resolve the issue are unknown. |  |  |  |  | The investigation for the limb ischemia and the thrombus has been completed. neither the device nor sufficient clinical details were provided to determine the root cause. therefore, the root cause of the reported limb ischemia and the thrombus could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in section: potential adverse events (united states). | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | Additional information for the initial mfr 1220648-2024-12681 was provided in sections b1, b2, b5, h1, and h6. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. delivery issue: per the clinical information provided, the root cause of the delivery issue was patient condition related to the replaced valve. thrombus: "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined". limb ischemia: "limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined." --- impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
25132 | 1220648-2024-12682 | 19608893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19608893&pc=OZD | 06/16/2024 | 08/07/2024 |  | 06/25/2024 |  |  | 06/16/2024 | 06/25/2024 |  |  | 02/28/2025 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024331624 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 33-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the implant an echocardiogram was performed and showed the pump¿s placement was too deep and was repositioned. there was concern for hemolysis and the pump was pulled back additional 3.4 centimeters. the patient remained on support despite repositioning and was reported as stable. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the data logs showed multiple suction alarms were triggered at early of the case then resolved. no other issues were found on the logs. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25133 | 1220648-2024-12683 | 19609637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19609637&pc=OZD | 05/29/2023 | 07/23/2024 |  | 06/25/2024 |  |  | 05/31/2024 | 06/25/2024 |  |  | 12/31/2024 |  | 06/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292396 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured thrombocytopenia with a "possible" relationship to the impella device used: ¿thrombocytopenia suspected mostly related to mechanical devices impella/ecmo (extracorporeal membrane oxygenation). 1st hit (heparin-induced thrombocytopenia) antibodies were positive, repeat hit was negative. 2 units platelets transfused on 6/2 and 4 units of rbcs (red blood cells) 5/29 64 5/30 41 5/31 29 6/1 51 6/2 44 6/3 51 6/4 61 6/5 76 6/6 79 6/7 82 6/8 97 6/9 117¿. the patient recovered with no sequelae. | [5 empty] | The investigation for thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25134 | 1220648-2024-12627 | 19593194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593194&pc=OZD | 04/23/2024 | 06/30/2025 |  | 06/24/2024 |  |  | 06/05/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481800 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 71 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Pericardial Effusion | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | As reported under the impact eu study, a patient was fitted with an impella 5.5 heart pump to provide hemodynamic support during and after high-risk cardiac surgery. impella support completed without any issues. two weeks after impella 5.5 removal, a deep wound infection occurred at the axillary impella access site. purulent secretion emerged and enterobacter cloacae was detected. the infection was treated with medication. according to the current assessment by the study doctor, the incident is not related to the impella 5.5 heart pump or to the heart surgery but is related to the impella procedure. | As reported under the impact eu study, a patient was fitted with an impella 5.5 heart pump to provide hemodynamic support during and after high-risk cardiac surgery. it was reported that the patient was treated with heparin for management of impella. heparin likely contributed to coagulopathic bleeding on reoperation. the patient also had thrombocytopenia which also may be device related. no additional information was provided. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the infection/sepsis has been completed. the investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ g1-corrected mfr site name and address h6 component code 4755 changed to 925. | Added the following: b5-updated narrative to include additional information for new adverse events pericardial effusion and thrombocytopenia. h6 clinical codes: 3271, 4431.
25135 | 1220648-2024-12628 | 19593255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593255&pc=OZD | 06/12/2024 | 07/31/2024 |  | 06/24/2024 |  |  | 06/12/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced substantial bleeding around the impella cp repositioning sheath following implant. attempts to control the bleeding via mattress sutures, angle matching, and forward tension failed. as a result, the decision was made to replace the pump. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12628: e4 should have been left blank. h.6 code 4755 was reported incorrectly. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25136 | 1220648-2024-12629 | 19593304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593304&pc=PYX | 06/12/2024 | 07/31/2024 |  | 06/24/2024 |  |  | 06/12/2024 | 06/24/2024 | 06/24/2024 |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  | IMPELLA RP | IMPELLA RP |  | PYX | PYX |  | IMPELLA RP | IMPELLA RP |  | ABIOMED, INC. | ABIOMED, INC. |  | US | US | 2025507394 | 2025507394 | 1000323 | 1000323 | No | No |  | Returned to Manufacturer | Returned to Manufacturer | Yes | No | Health Professional | Health Professional |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, and there was oozing at the access site. the sheath was upsized, platelets were ordered, extra sutures were placed at all the sheaths. the patient remained on impella rp support. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, extra sutures were placed at all sheaths, including impella and swan, to reduce bleeding as per the clinial description provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12629. e4 should have been left blank. h.6 code 4755 was reported incorrectly. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
25137 | 1220648-2024-12630 | 19593333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593333&pc=PYX | 06/16/2024 | 08/05/2024 |  | 06/24/2024 |  |  | 06/16/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025480513 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella rp support developed oozing in the right groin access site. multiple products were given to mitigate coagulopathy. one unit of rbc (red blood cells) was given. it was further reported that no pulses were noted in the same affected leg. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ | The investigation for the reported access site bleed and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and limb ischemia could not be determined. d4-corrected model number h6 component code 4755 changed to 925.
25138 | 1220648-2024-12631 | 19593437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593437&pc=OZD | 05/26/2024 | 07/16/2024 |  | 06/24/2024 |  |  | 05/26/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494039 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the next day of support, the purge flow decreased while the purge pressure spiked during patient support. the purge cassette was replaced in an attempt to troubleshoot the issue, but the issue remained. the decision was subsequently made to administer tissue plasminogen activator through the purge line to resolve the increase in purge pressure. there was no report of harm to the patient as a result of this event. further information was received regarding the patient's final outcome. approximately two days later, the decision was made with the family to withdraw care as the patient's labs/hemodynamics continued to trend in "wrong" direction. the patient expired. abiomed's medical safety officer reviewed the outcome and noted there is not enough information to associate the use of the device from the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | High purge pressure and low purge flows reported on day 2 of support. no product was returned. the data logs confirm high purge pressure and low purge flows. the logs show a gradual build up in purge pressure which resolved about 9 hours after tpa was added to the purge. the root cause of the high purge pressure was most likely biomaterial build up.
25139 | 1220648-2024-12632 | 19593639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593639&pc=PYX | 06/05/2024 | 08/02/2024 |  | 06/24/2024 |  |  | 06/05/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX WITH SMART ASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2025502663 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was implanted with an impella rp flex from an impella 5.5 for further optimization of the right ventricle in setting of worsening right ventricle. the same day of the implant, the pulmonary artery pressure spiked, and there was a reduction in flows. it was noted that the team was unable to verify the cause or if the impella rp flex was properly flowing. however, there was no spike in motor current. consequently, the impella rp flex was removed and replaced with extracorporeal membrane oxygenation support. there were no adverse effects reported as a result of this event. | [5 empty] | Data logs show that about 3 hours into the case the ps begins to drift upwards but eventually stabilizes for 5 hours before drifting more severely for the rest of the case. psnr alarms are not triggered. the ps does not recover. mc is stable with the p-level changes the entire case with no mc spikes outside of p-level changes. cvp is stable. the returned product passed electrical testing. the returned 5s was within specification. dp sensor drift was reproduced in the pressurized flow loop in the lab. the root cause of the placement signal issues was most likely sensor drift. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
25140 | 1220648-2024-12633 | 19593678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593678&pc=OZD | 06/16/2024 | 08/17/2024 |  | 06/24/2024 |  |  | 06/16/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490282 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and they had access site bleeding. multiple products given to mitigate coagulopathy. one unit of blood products were provided to the patient. additionally, the patient had no pulses in their leg. staff monitored. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12633 as it is unknown the initial reporter also sent the report to the fda. g.1 revised manufacturing site name and address section as previously entered incorrectly on initial manufacturer device report number 1220648-2024-12633. h.6 code 4755 was reported incorrectly on the initial manufacturer device report number 1220648-2024-12633. new codes were added to component codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12633 in accordance with updated procedures.
25141 | 1220648-2024-12634 | 19593714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593714&pc=OZD | 06/17/2024 | 07/11/2024 |  | 06/24/2024 |  |  | 06/17/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481086 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella 5.5 support there was an impella stopped alarm with a flat motor current reading 0/0. impella was attempted to restart on same aic (automated impella controller) multiple times without success. at one of the last attempts to restart there was a motor current spike seen. the pump was transferred to a new aic with an attempt to restart which was also unsuccessful. at the time of the alarm, the patient had stood up abruptly and all three fixation anchors had ripped from the patient¿s chest. the device was surgically removed at the bedside, patient tolerating removal well. | [5 empty] | The investigation for the pump stop has been completed. according to clinical data, the pump stop occurred on day 10 of support. patient stood up abruptly at the time of the alarm. no data logs were returned. the pump stop was reproduced in the lab with alarm #115 and the electrical reading showed an open electrical line. upon inspection, it was discovered that the motor cables in the red handle on the patient cable side were necked and the ground cable was broken. this is evidence of excessive force during support and aligns with the clinical mention of the pump stop occurring when the patient stood up abruptly. the root cause of the pump stop was determined to be an electrical open issue due to excessive force. | The impella device was received from the customer on 20-jun-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
25142 | 1220648-2024-12635 | 19593771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593771&pc=OZD | 06/17/2024 | 07/31/2024 |  | 06/24/2024 |  |  | 06/17/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474622 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was implanted in routine fashion but after advancing repositioning sheath, groin site started bleeding. manual pressure applied. angio taken through rewire access port and showed common femoral artery (cfa) dissection. the cp was weaned and removed. covered stents were deployed over the access site. groin site stable after stent deployment. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. e4 should have been left blank . h.6 code 4755 was reported incorrectly. new code was added to component code.
25143 | 1220648-2024-12636 | 19593847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19593847&pc=OZD | 06/05/2024 | 07/22/2024 |  | 06/24/2024 |  |  | 06/05/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171907 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a "controller error-switch to back-up controller" message. the event occurred during a patient transfer. there was no patient harm. the aic was replaced. | [5 empty] | The investigation into the automated impella controller error has been completed. data analysis: ¿ on the complaint date ((b)(6) 2024), only the demo pump was used in the initial run, and there was no controller error. ¿ in the second bootup, commercial pump sn: (b)(6) was used for 2 hours 51 minutes, and there was no controller error. ¿ immediately after the third bootup, the console displayed ¿controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) ¿ alarm¿ event #24. ¿ then, the console was shut down and turned manually a couple of more times. ¿ again, the console displayed controller error ¿ event #24. this confirms the reported failure mode. device analysis: this console was tested after arriving in fs. upon power cycling the console 200 times, it was unable to reproduce the reported failure mode - controller error. the root cause for the intermittent keyboard communication issue could not be determined due to the inability to reproduce. f6 and f8 should have been blank. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
25144 | 1220648-2024-12638 | 19594179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19594179&pc=OZD | 06/05/2024 | 08/13/2024 |  | 06/24/2024 |  |  | 06/05/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025441127 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The complainant reported that an impella cp was placed in a 67-year-old female patient. the impella ran without alarms or incidents. during intervention in the intensive care unit, the patient's circulatory parameters became increasingly unstable. there was a drop in hemoglobin. the patient was determined to have an aortic dissection in the abdominal area. the cause of the dissection could not be directly attributed to the impella. the patient was unable to receive treatment and the decision was made to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported vascular damage has been completed. no product or data logs were returned for evaluation. clinical details noted that an aortic dissection in the abdominal area after patients circulatory parameters became unstable. it was not noted if the impella was the cause of the aortic dissection, no issues with insertion were noted. the root cause of the vascular damage cannot be determined as limited clinical information was provided. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ added- h6 component code 925. d4-corrected model number. f6/f8 should have been left blank in the initial report.
25145 | 1220648-2024-12639 | 19594370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19594370&pc=OZD | 06/05/2024 | 08/01/2024 |  | 06/24/2024 |  |  | 06/05/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025460506 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The complainant reported that during impella 5.5 support, a patient of an unknown age and gender was in critical condition. the patient had a mitral valve replacement done the previous day. the impella 5.5 was in poor position and there was no echocardiogram image available. there was a thrombus at the level of mitral valve visible. therefore, the patient was transferred to operating room for thrombus removal. however, it was impossible to stabilize the patient with impella 5.5 only and the patient expired on support. no additional information available. | [5 empty] | A.5 ethnicity and race are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of positioning issues, unspecified (allegation) ¿ death, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12639: d4 model number f6/f8-should have been left blank.
25146 | 1220648-2024-12640 | 19595042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595042&pc=OZD | 06/17/2024 | 07/25/2024 |  | 06/24/2024 |  |  | 06/17/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495330 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when placing the impella cp, 14fr peel away was removed and the impella repo sheath was advanced into place. the physician shot a 1-leg runoff that showed the repo sheath was occlusive. the physician then placed a retrograde 5fr sheath in the lfa (left femoral artery) and a 5fr antegrade sheath to the rfa (right femoral artery) as a reperfusion strategy. the patient is stable. | [5 empty] | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-12640: a5: ethnicity missing, d4: model number, e4: should have been left blank. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25147 | 1220648-2024-12641 | 19595079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595079&pc=OZD | 06/17/2024 | 08/13/2024 |  | 06/24/2024 |  |  | 06/17/2024 | 06/24/2024 |  |  | 05/31/2025 |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024340889 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and via direct access and an aortic dissection was noted. ct (computed tomography) scan confirmed a stanford type a dissection. with the graphs being down and the aortic dissection, care team spoke with family, decision was made to withdraw care on patient. the patient expired after 15.68 hours of support. the relationship between the cause of death and the impella is unknown. r. | [5 empty] | The investigation for the vascular damage has been completed. clinical details states that an aortic dissection found during graft evaluation. no reports of excessive force or patient anatomy issues. no product was returned. the root cause of the vascular damage - great vessel was not determined as no product was returned and limited clinical information was provided d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
25148 | 1220648-2024-12642 | 19595098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595098&pc=PYX | 06/17/2024 | 08/05/2024 |  | 06/24/2024 |  |  | 06/17/2024 | 06/24/2024 |  |  | 11/30/2025 |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024435395 |  | 10000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella implant was aborted due to patient's clot. the sheath was attempted to be placed however, a clot was found and the impella implant was aborted. the surgeon advised not attempting contralateral side after examining ct (computed tomography) imagery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿handle with care. the impella rp flex with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details state there was a clot found during sheath placement. as a result of the clot, the procedure was aborted. the root cause of the clot could not be determined without additional clinical details.
25149 | 1220648-2024-12643 | 19595141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595141&pc=OZD | 06/15/2024 | 07/30/2024 |  | 06/24/2024 |  |  | 06/18/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489327 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp site is in the right leg and right pedal pulses are intermittent and right foot increasingly mottled. fasciotomy with medial and lateral incision to right calf in impella leg was done and after fasciotomy, pulses returned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. | The investigation limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12643:
25150 | 1220648-2024-12644 | 19595187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595187&pc=OZD | 06/18/2024 | 07/19/2024 |  | 06/24/2024 |  |  | 06/18/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103518 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | No patient involvement: it was reported that when powering on the aic (automated impella controller) a ¿system self check failed, change console immediately" displayed on the screen. unable to turn off the unit and the buttons are unresponsive. there was no patient involvement. | [5 empty] | The cause of the automated impella controller (aic) issue (loss of communication with pud due to loss of power to pud) was a loose connector. d9 added e1 name prefix/title. e4 should have been left blank in the initial report. g1 reporting contact fax number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25151 | 1220648-2024-12645 | 19595197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595197&pc=OZD | 06/18/2024 | 08/13/2024 |  | 06/24/2024 |  |  | 06/18/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493948 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and during support the patient had hemolysis. the pump was repositioned, and the hemolysis was still not resolved. the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed generally pump ran without issue during support. there were 2 times, flow dropped suddenly but resolved quickly that triggered auto suction respond. there was high purge pressure (pp) alarm triggered in the end of the case. the pump was returned in damage condition as the catheter was cut off. visual inspection found biomaterial in outflow cage and purge gap. no other issue was found on the impeller or pump housing wall and the rest of the pump. the root cause of the hemolysis was most likely biomaterial ingestion. added- d9 device return date d4-corrected model number. g1-corrected mfr site name and address.
25152 | 1220648-2024-12646 | 19595255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595255&pc=OZD | 06/18/2024 | 10/20/2025 |  | 06/24/2024 |  |  | 06/18/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458558 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | Positioning Problem |  |  |  |  | Respiratory Failure | Thrombosis/Thrombus | [11 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 after approximately 85 days on support, position alarm and frequent suction alarmed occurred. there was a concern for a sensor failure or pump ingestion. the impella 5.5 was exchanged. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory failure on (b)(6) 2024 with "unknown/undetermined" relationship to the impella use. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Data review: data logs showed no motor current spikes or low pump flow for the respective p-level observed in the logs. logs also showed ps that matches the sensor wear data log pattern. the placement signal shifts down and is lower with mc variability and there is gradual decrease in signal not reliable (snr) suspecting sensor wear. psnr alarm did not trigger but the pump was likely removed prior to the sensor wear reaching the level of triggering that alarm. product analysis: visual inspection found biomaterial inside the pump housing. no other issues were found on the rest of the pump. optical bench 5s was tested for complaint pump and was found to be in normal range. flow loop tested 5.5 pump for 16 hours and the placement signal(ps) was declining since the start of pump run. even though snr being within normal range, gradual decrease in snr was observed along with steady decline of ps. other 5s parameters (contrast, gain, lamp & light) had no major variations. sensor wear was likely confirmed with the reproduction testing. this pump was used for more than 83 days in support which is well beyond the design life per design traceability matrix 0550-9900. conclusion: the root cause of optical signal issue was likely sensor wear due to using impella beyond the design life per design traceability matrix 0550-9900. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined d9 date device returned was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12646-1. | B1 revised to both adverse event and product problem to reflect additional information received. b5 updated with additional information. h1 revised to serious injury to reflect additional information received. h6 health effect - clinical code and health effect - impact code revised to reflect additional information received. b7 revised to reflect medical history not included on initial manufacturer device report. d4 model number revised on initial manufacturer device report. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact facility name was revised as it was not included on initial manufacturer device report. g1 reporting contact fax number was revised as it was not included on initial manufacturer device report. g2 report source revised to reflect study which was not included on initial manufacturer device report. h6 health effect - clinical code 4440 was inadvertently selected on manufacturer device report. h6 code 2616 was reported incorrectly on manufacturer device report. new code was added to medical device problem code.
25153 | 1220648-2024-12647 | 19595464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19595464&pc=OZD | 06/18/2024 | 09/17/2025 |  | 06/24/2024 |  |  | 06/18/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478475 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and they had oozing at the access site. blood products were administered. additionally, the patient had ectopy with position changes. echo was done on the impella and it was found that the pump did appear to be moving slightly in systole. furthermore, there were placement signal not reliable alarms with absent ao and left ventricle placement signal. support continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the ectopy has been completed. the impella was returned for evaluation. upon visual inspection, there were no abnormalities with the pump/repo unit and the introducer was not returned. clinical details states there was a patient update that the access site was oozing. the oozing appeared similar to what nurses had seen from ecmo access sites previously. the root cause of the access site bleed was determined to be patient condition. without additional clinical details, therefore, the root cause of the ectopy could not be determined. clinical details state the patient's ldh increased, urine began to turn red, and a thrombus was suspected to be interacting with the impeller causing hemolysis. there were no data logs returned for investigation. product return investigation showed a large degree of biomaterial wrapped around the impeller. the root cause of the hemolysis was determined to be biomaterial. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. data logs were not returned for investigation. returned product was found to have a light leak in the red handle on the catheter side. additionally, at the joint between the red handle and catheter, the catheter had a major kink. the root cause of the optical signal issue was determined to be a damaged optical fiber line at the junction of the red handle and catheter, most likely due to excessive force. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-12647.
25154 | 1220648-2024-12637 | 19597812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19597812&pc=OZD | 06/05/2024 | 08/12/2024 |  | 06/24/2024 |  |  | 06/05/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025490013 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was implanted in a patient admitted in with severe mitral regurgitation. the pump was placed, but its location caused premature ventricular contraction, which prompted the impella cp pump to be explanted and positioning of a new pump to the left ventricle. it was noted the native anatomy was cause of the impella cp pump malpositioning. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was patient condition related due patient anatomy as the patient had severe mitral valve regurgitation. impact code 4629.
25155 | 1220648-2024-12571 | 19601696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19601696&pc=OZD | 06/12/2024 | 08/02/2024 |  | 06/24/2024 |  |  | 06/12/2024 | 06/24/2024 |  |  |  |  | 06/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 82-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella cp insertion, the md went to peel away the sheath and clear disc from introducer was still on sheath. the physician could not properly advance the repo sheath. md proceeded to use or sheers to cut disc off and proceeded to advance the repo sheath. there was no harm to the patient. |  |  |  |  | The investigation of difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12571: d4 model number. e4 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
25156 | 1220648-2024-12577 | 19580505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19580505&pc=OZD | 06/12/2024 | 11/20/2024 |  | 06/21/2024 |  |  | 06/12/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693068 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | No patient involvement: the complainant reported that the aic (automated impella controller) had a system self-check failed with change console immediately alert showing in the screen. there was no patient involvement. | [5 empty] | The investigation for the console (aic) power issues has been completed. data logs were reviewed which showed a communication issue with the epm during the initial bootup, with error messages "sau_ser6 cmd seq.101 timeout. reply err and sem_post. due to the communication issue between epc and the epm, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the console was returned and was dead upon arrival. the console was plugged into ac power and performed the battery test. battery failure upon booting up and batteries were not charging or discharging. the reported complaint was determined to have been caused by a malfunction of the impellatronic pca resulting in a loss of communication between the epc and the epm. the batteries were also found to be fully depleted and therefore unable to communicate with the pbm or charge. | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25157 | 1220648-2024-12579 | 19580539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19580539&pc=OZD | 06/04/2024 | 07/30/2024 |  | 06/21/2024 |  |  | 06/04/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487199 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a 68-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the pump was placed via the femoral artery to stabilize while awaiting surgery. it was reported that the access site developed a bleed. the medical staff intervened by using suture(s) and surgicell. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the issue was resolved upon placing additional forward tension suture. added- h6 component code 925 f6/f8 should have been left blank in the initial report.
25158 | 1220648-2024-12576 | 19580590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19580590&pc=OZD | 06/12/2024 | 08/05/2024 |  | 06/21/2024 |  |  | 06/12/2024 | 06/21/2024 |  |  | 07/31/2025 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443823 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when they had access site bleeding. a hemostatic patch was placed, dressing was changed and perclosed was cinched down (secured using stopcock and dead-end cap). | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access bleeding issue was most likely use related, a hemostatic patch was placed, dressing was changed and perclosed was cinched down (secured using stopcock and dead-end cap) to achieve hemostasis as per the clinical description provided. d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25159 | 1220648-2024-12578 | 19580709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19580709&pc=OZD | 06/13/2024 | 08/05/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491626 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and low pump flows were noted and suction alarms occurred. shortly after, flows dropped down to 0.0lpm and the impella was removed. the doctor suspected a thrombus was ingested. the patient had clotting issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow was most likely due to biomaterial as a result of patient's condition. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12578: d4 model number, e4 should have been left blank.
25160 | 1220648-2024-12582 | 19582587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19582587&pc=OZD | 05/24/2024 | 08/01/2024 |  | 06/21/2024 |  |  | 05/27/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025489728 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (race and age unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the systolic left ventricle value was noted to be as high as the systolic positioning signal value. the pump position was noted to be good via echocardiogram. however, there was a thrombus observed in the inlet. the impella was therefore explanted, and a new one was placed. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation thrombus has been completed. per clinical imaging, the pump was in proper position and there was biomaterial at the inlet. they removed the pump and replaced it. there were no data logs returned to aid the investigation. in investigating the pump, there were no abnormalities noted from visual inspection. the pump was run in the lab without any alarms and all optical parameters were within spec. the root cause of the saturated flows was most likely biomaterial. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-12582: e4: should have been left blank. g1: manufacturing site name and address. h6: code 4755 was reported incorrectly. new code was added to component code.
25161 | 1220648-2024-12583 | 19582894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19582894&pc=OZD | 05/27/2024 | 07/11/2024 |  | 06/21/2024 |  |  | 05/27/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486171 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient developed hemolysis during impella cp support. the patient, admitted in with acute myocardial infarction/cardiogenic shock, underwent percutaneous coronary intervention. the patient continued to decline and required intubation. an impella cp device was postoperatively placed for mechanical circulatory support, and a swan was placed. the patient was sent to the unit on support. and next day developed hemolysis. volume was originally given to counteract the condition, but the issue remained unresolved until explant. further information regarding the patient's status noted a move to comfort care because the patient's outlook was "grim." the patient subsequently expired. medical safety review was completed and noted there is not enough information to associate use of the impella device with the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed no positioning issue was found during the case. there were some motor current (mc) spikes and multiple suction alarms triggered during support. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
25162 | 1220648-2024-12584 | 19583103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583103&pc=OZD | 05/27/2024 | 07/22/2024 |  | 06/21/2024 |  |  | 05/27/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025440751 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (race unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support during a pulmonary vein isolation (pvi) ablation procedure. during impella cp support, the patient experienced atrial fibrillation during the pvi procedure. the impella was explanted and the patient remained stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 manufacturing site name and address corrected e4 should have been left blank in the initial report. h.6 code 4755 was reported incorrectly on manufacturer device report. new code was added to component code.
25163 | 1220648-2024-12585 | 19583482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583482&pc=OZD | 06/07/2024 | 07/18/2024 |  | 06/21/2024 |  |  | 06/07/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764240 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) sounded and displayed an alarm for battery failure despite being connected to the power supply. consequently, the aic was replaced. there was no report of harm to the unknown age patient noted to be in stable condition. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the battery error issue has been completed. data review: after reviewing the imc logs, the reported battery error issue was confirmed. there were multiple instances of battery failure alarms (transport connection blown/missing). battery failure alarm (low voltage) [v < 12v], and battery level low (50%) found on the reported occurrence date. device analysis: ic11003 was tested after arriving in field service. battery error issue was able to be reproduced. results of running batttest on telnet revealed that vcell2 in battery1 had a voltage that was significantly less than the rest of the cells in the battery. conclusion: reported battery error issue was determined to be caused by internal battery cell imbalance (found in vcell2 of battery1). g1 corrected manufacturing site name and address
25164 | 1220648-2024-12587 | 19583535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583535&pc=OZD | 06/12/2024 | 07/31/2024 |  | 06/21/2024 |  |  | 06/12/2024 | 06/21/2024 |  |  | 09/30/2025 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399152 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown age) post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. while on support the pump position was near the mitral valve. the physician tried to rotate the pump and it was too deep so it was pulled back. at that time, a large amount of bleeding was observed and the physician found the left ventricle had been perforated. the ventricle was patched and the pump was attempted to be manipulated again, resulting in another perforation. this perforation required them to lift the heart and access the backside. upon doing so, the pump developed purge pressure high and purge system blocked alarms. the pump was explanted. patient was then supported with venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ventricle perforation. impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the high purge pressure, ventricle perforation, and the positioning issue has been completed. the pump was not returned for investigation. data logs show a 5-minute rise in purge pressure with several suction alarms and motor current spikes, followed by saturated pump flow during this event. additionally, placement signals were disabled after noted impella position in the aorta. the doctor reported ventricular perforation during repositioning after the pump was found to be too deep. the damage was repaired. the doctor discovered a second perforation while repositioning. high purge pressure was also reported during this event. the cause of the ventricular perforation was most likely improper positioning of the device, possibly due to improper technique. the cause of the high purge pressure was most likely biomaterial, based on data log review.
25165 | 1220648-2024-12588 | 19583573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583573&pc=PYX | 06/12/2024 | 07/18/2024 |  | 06/21/2024 |  |  | 06/12/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025487969 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced bleeding from their access site during impella rp support. the sutures were tightened to slow the bleed and heparin was held to manage the condition. the patient experienced oozing from other sites in conjunction with the impella site leading to a transfusion of one unit of packed red blood cells. the pump was eventually removed from the femoral artery and a new impella rp pump was placed via a new access point. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since the sutures were tightened to slow the bleed and heparin was held to manage the condition. d4 model number corrected. g1 manufacturing site name and address corrected. corrected h2 type of reportable event. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25166 | 1220648-2024-12591 | 19583615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583615&pc=OZD | 06/09/2024 | 08/05/2024 |  | 06/21/2024 |  |  | 06/09/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490195 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp had been placed through right common femoral artery using 25cm peel-away sheath. waveforms and automated impella controller (aic) was showing impella in aorta, running on p-9 with flows >4.5l. the patient¿s blood pressure was decreased so the p-level was brought down to p-2. md stated that the impella cp moved when peel away was removed. the impella was removed due to poor positioning and a new cp was successfully placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely was use related due to improper technique. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12591: e4 should have been left blank, g1 reporting contact fax number missing.
25167 | 1220648-2024-12590 | 19583617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583617&pc=OZD | 05/25/2024 | 07/25/2024 |  | 06/21/2024 |  |  | 05/25/2024 | 06/21/2024 |  |  | 11/30/2025 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025442474 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the impella was probably dislocated during resuscitation, and repositioning was no longer possible. the patient was in very critical condition with high doses of catecholamines. there are no details received regarding the impella position prior to cardiopulmonary resuscitation or thuoy-borst. the impella cp was explanted. | The decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. |  |  |  |  | A.5 ethnicity and race are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for positioning issue has been completed. the impella device was not returned for evaluation. based on clinical description, the root cause of positioning issue was most likely use related due to cardiopulmonary resuscitation (cpr). the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12590: b5 included information that was inadvertently left out from the original report f6/f8-should have been left blank. g1 manufacturing site address line 2 should have been left blank.
25168 | 1220648-2024-12593 | 19583668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583668&pc=OZD | 06/13/2024 | 07/22/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1542953 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a purge disk holder that was broken. the aic was removed from service. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the automated impella controller purge issues has been completed. data review : no data review is needed for this complaint because it was only reported that the retainer was broken. device analysis : the purge disc retainer was found to be broken (retainer completely broken off). the root cause of the issue was a broken purge disk retainer.--defective component. added d9 device return and date e4 was checked no in initial mdr. it should have been blank as abiomed does not know if initial reporter also sent report to fda.
25169 | 1220648-2024-12595 | 19583799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583799&pc=OZD | 06/13/2024 | 07/25/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515093 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the day after implant the team observed the limb to be ischemic. the device was explanted and the patient was escalated to the impella 5.5 axillary access. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. g1 reporting contact fax number missing. e4 should have been left blank.
25170 | 1220648-2024-12586 | 19583847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583847&pc=OZD | 05/24/2024 | 08/01/2024 |  | 06/21/2024 |  |  | 05/24/2024 | 06/21/2024 |  |  | 04/30/2025 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024346184 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient (age, gender and race unknown) in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. following insertion of the 25 centimeters peel-away, the impella cp was unable to pass inside it. the peel away was exchanged, but the impella was still unable to be delivered. there were no kinks noted on the device. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. there were no abnormalities on the returned pump. the pump was able to be inserted into both introducer lengths in the lab. a kink was found on the long introducer, likely from the insertion attempts in the field. the root cause of the difficulty inserting pump through introducer issue was most likely use related as they were unable to insert the pump through both introducer lengths and the pump was able to be inserted through both in the lab and no damages were observed. f6/f8 should have been left blank in the initial report. | A.5 ethnicity and race are unknown. the impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25171 | 1220648-2024-12597 | 19583878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583878&pc=OZD | 06/13/2024 | 07/24/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474755 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that on day one of impella cp support the patient had no blood flow distal to impella. the patient has known to have severe pvd (peripheral vascular disease), but impella was placed emergently. external fem-fem placed in same setting to perfuse lower leg. pulse was per doppler post impella placement but during the night the pulse was no longer able to be heard on doppler. a decision was made to pull the impella into the aorta and place on p1. the patient was on p1 for about 1 hour before taken to or for explant. the patient maintained hemodynamics without the support of any gtts. once in or, rfa (right femoral artery) cutdown was performed. the impella was explanted successfully. rfa thrombectomy, rfa repair and fasciotomy to lower leg was done. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12597: e4 should have been left blank. h.6 code 4641 and 4628 were reported incorrectly. new code was added to health effect - impact code
25172 | 1220648-2024-12598 | 19583913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583913&pc=OZD | 06/13/2024 | 08/01/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474974 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when placing the impella 5.5, surgical cut down of the right axillary artery was performed and 10mm graft sewn on. attempted to cross the valve with abiomed recommended al1 and was unsuccessful. a bern catheter was also used, which was also unable to cross the aortic valve (av). after multiple attempts with different wires and catheters, the decision was made to abort placement of 5.5. the patient is stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump was unable to pass through vasculature and couldn't be delivered. | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
25173 | 1220648-2024-12589 | 19583937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583937&pc=OZD | 06/08/2024 | 07/23/2024 |  | 06/21/2024 |  |  | 06/08/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485323 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. at time of explant, clot dislodgment may have occurred when the impella cp pump was removed over the access wire. dry closure technique and manual pressure were utilized to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12589: e4 should have been left blank.
25174 | 1220648-2024-12599 | 19583950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583950&pc=OZD | 06/05/2024 | 07/24/2024 |  | 06/21/2024 |  |  | 06/05/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473977 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Puncture/Hole | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility reported a 50-year-old male was escalated to impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced access site bleeding and there was an increased output in the chest tubes and the doctor opened the patient's axillary site and the bleed was confirmed. the patient received six units of blood products and the doctor secured the site. additionally, blood was found leaking from the red impella plug. it was noted that a suture needle may have stuck the catheter. the patient remains on support and was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Pump was returned for investigation. the red handle was wrapped throughout to stop the blood leak. there were 3 holes observed on the catheter near the 25 cm mark. the red handle was full of blood when opened. clinical also mentions possibility of catheter puncture due to suturing which was confirmed during the product investigation. therefore, pre the product investigation the root cause of the product damage (hole in catheter) issue was use related to improper technique. clinical mentions that bleeding was noted at the anastomosis which was resolved by adding sutures to the graft. therefore, per clinical information provided, the root cause of the access site bleeding issue was use related to improper technique. d6b added: d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12599: f6/f8-should have been left blank . h6 component coded added.
25175 | 1220648-2024-12592 | 19583966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19583966&pc=OZD | 06/12/2024 | 07/29/2024 |  | 06/21/2024 |  |  | 06/12/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485224 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Positioning Failure | Failure to Advance | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the patient had access site bleeding. heparin was placed on hold and a figure 8 stitch was applied to the access site. additionally, the impella cp was placed in the aorta during surgical mode. upon coming off of bypass, the doctor attempted to advance the impella and it was unsuccessful. multiple attempts were made however, the pump was removed | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since product was not returned and there is a lack of clinical details. the root cause of the positioning issue was likely wireless re-access. the impella device is not indicated for wireless re-access per instructions for use the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12592: d4 model number. d4 catalog number. e4 should have been blank. h.6 code 4628 and 2524 was reported incorrectly. new code was added to health effect - impact code.
25176 | 1220648-2024-12594 | 19584018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584018&pc=OZD | 06/13/2024 | 08/06/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494601 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp when the impella was in the aorta. the physician was unable to recross and the impella was removed. | [5 empty] | The investigation for the positioning issue has been completed. an impella cp was successfully delivered through the right femoral artery. the pump was positioned in the aorta overnight, and they were unable to recross the valve the next morning. the impella was removed. logs were not applicable in this investigation because position was verified by echo during the case. the root cause of the positioning issue was not determined since no product was returned and insufficient clinical evidence was provided. d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25177 | 1220648-2024-12600 | 19584028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584028&pc=OZD | 06/14/2024 | 08/06/2024 |  | 06/21/2024 |  |  | 06/14/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella had low flows after insertion due to a kink in the cannula. the physician made choice to explant and replace the pump. there was no patient harm. | [5 empty] | The investigation for the low pump flow has been completed. the impella device was not returned for investigation. data logs were reviewed and no position issues were observed. no motor current spikes were observed. low pump flow confirmed at p-9 with suction alarm. the cause of the low pump flow issue most likely was cannula kink due to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25178 | 1220648-2024-12601 | 19584102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584102&pc=OZD | 06/13/2024 | 08/05/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515097 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 41-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after several attempts of inserting the impella, the cp was unable to cross the aortic valve. the impella placement was aborted. | [5 empty] | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of aortic valve replacement and the valve may not be working correctly due to scar tissue.
25179 | 1220648-2024-12606 | 19584123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584123&pc=OZD | 06/17/2024 | 08/13/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495368 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp sheath was inserted and when the dilator was removed a kink was noted in the body of the sheath. this limited the catheter¿s maneuverability, so the sheath was exchanged for another impella sheath. after multiple attempts with the impella trying to cross the aortic valve, the impella insertion was aborted due to the patient vessel tortuosity. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues and introducer damage has been completed. the impella device was not returned for evaluation. the root cause of the delivery issues was most likely due to patient condition as patient had transcatheter aortic valve replacement (tvar) valve. the cause of the issue was an introducer sheath kink. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12606: d4 model number. e4 should have been left blank.
25180 | 1220648-2024-12604 | 19584127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584127&pc=OZD | 06/14/2024 | 08/07/2024 |  | 06/21/2024 |  |  | 06/14/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474574 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an impella cp implant and when the doctor accessed the hemostatic valve for single access pci (percutaneous coronary intervention) they encountered resistance with the micropuncture needle. the needle lurched forward and went through the peel away sheath and the doctor was concerned of vessel puncture/perforation. the impella was removed and re-implanted in the left femoral artery. no dissection was noted and the site was free of hematoma at the end of the procedure and pulses were identified with doppler. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. the root cause of the introducer damage - mechanical was most likely due to a use issue as the angle for single access micro-puncture insertion was not correct. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12604.
25181 | 1220648-2024-12596 | 19584162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584162&pc=OZD | 06/13/2024 | 08/21/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  | 08/31/2024 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2023251630 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. during the implant procedure, the impella sheath introducer kinked due to calcification in iliac artery, the problem was solved with a competitor¿s 16fr introducer that is stiffer. the patient underwent pci procedure and was stable upon impella cp explant. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to calcium in iliac artery. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
25182 | 1220648-2024-12602 | 19584308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584308&pc=OZD | 06/14/2024 | 08/02/2024 |  | 06/21/2024 |  |  | 06/14/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support had no perfusion down the impella leg. reperfusion cannula was placed. the patient is stable. | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12602: d4 model number. e4 should have been left blank.
25183 | 1220648-2024-12612 | 19584385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584385&pc=OZD | 06/16/2024 | 08/08/2024 |  | 06/21/2024 |  |  | 06/16/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025487675 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 39-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was acute motor current abnormality and spike with inability to flow noted. the impella was repositioned, volume was given and there was no improvement. staff was concerned for pump failure and the pump was replaced. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the returned pump visually inspected. no cannula kink observed. no broken blades found. biomaterial around impeller observed. clinical details confirmed attempt repositioning issue didn't solve the low flow issue. the cause of the low pump flow was biomaterial ingestion. d9 date device returned to manufacturer added.
25184 | 1220648-2024-12611 | 19584426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584426&pc=OZD | 06/15/2024 | 08/02/2024 |  | 06/21/2024 |  |  | 06/15/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL,PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1691019 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 41-year-old male patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the associated automated impella controller (aic) needed to be changed due to a leak in the purge system and maybe a broken purge disc. after aic was exchanged the pump ran well. there was no patient harm. | [5 empty] | The investigation of aic - purge disc/flag issues has been completed. logs are irrelevant to this complaint as the clinical staff only reported that the purge disc retainer may be broken. upon examining automated impella controller (aic) for any visible defects, it was evident that the blue purge disc retainer was broken. given the aic's purge retainer was broken, it likely contributed to the reported purge leak. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report: e4 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25185 | 1220648-2024-12613 | 19584494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584494&pc=OZD | 06/16/2024 | 08/08/2024 |  | 06/21/2024 |  |  | 06/16/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485918 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 55-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that purge flows dropped to less than 1 ml per hour during impella cp support. tissue plasminogen activator was added to the purge system, but the issue persisted. the pump subsequently began to show signs of failure and the decision was made to replace the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. low purge flow reported on day 5 of support. the data logs confirm high purge pressure alarms. the data logs show a motor current spike after which the purge pressure started increasing gradually before high purge pressure alarms were triggered. there were also purge system blocked alarms. the high purge pressure did not resolve for the rest of the procedure. only the purge cassette and a severed y-connector was returned. the purge cassette had no visual abnormalities and no flow issues. the root cause of the high purge pressure was most likely biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ added- d9 device return date g1-corrected mfr site name and address.
25186 | 1220648-2024-12605 | 19584586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584586&pc=OZD | 06/16/2024 | 07/01/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1263046 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella aic (automated impella controller) was not reading the purge disc properly. multiple attempts were made to re-insert the purge disc and cassette, but the issue was not resolved. the aic was exchanged. there was no harm to the patient. | [5 empty] | This manufacturer report was submitted in error due to duplication and should be retracted. the event was filed under manufacturer report number 1220648-2024-12619. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25187 | 1220648-2024-12603 | 19584612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584612&pc=OZD | 06/13/2024 | 07/31/2024 |  | 06/21/2024 |  |  | 06/13/2024 | 06/21/2024 |  |  | 09/30/2025 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025451706 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was taken to the operation room for impella 5.5 escalation. as the 5.5 was being inserted the patient had ventricular fibrillation arrest. cpr was initiated. the physician proceeded with cannulation for ecmo (extracorporeal membrane oxygenation). axillary graft utilized for arterial return. unfortunately, the patient¿s condition continued to decline and ultimately support was withdrawn. the impella 5.5 case was aborted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. h6 component code 4755 changed to 925.
25188 | 1220648-2024-12607 | 19584614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584614&pc=OZD | 06/17/2024 | 08/12/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479391 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the insertion of impella 5.5, an attempt was made to get left ventricle access, but unable to cross due to anatomic issues. the impella then was inserted into the graft otw with hopes that impella might cross with concurrent digital manipulation of the heart. multiple attempts failed. the patient hemodynamics were sound on minimal ggt (gamma-glutamytransferase) support, so the 5.5 insertion was aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition as the pump could not pass due to angle of valve in relation to aorta and aneurysmal anatomy of ascending d4-corrected model number.
25189 | 1220648-2024-12615 | 19584662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584662&pc=OZD | 06/16/2024 | 08/07/2024 |  | 06/21/2024 |  |  | 06/16/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025487647 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory. it was reported the patient experienced bleeding at their access site during impella cp support. the patient was transfused an unspecified quantity of red blood cells to counteract the blood loss. support with the implanted pump continued following the intervention. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. isection: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12615: e4 should have been left blank. h.6 code 4755 was reported incorrectly. new code was added to component code.
25190 | 1220648-2024-12608 | 19584688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584688&pc=OZD | 06/15/2024 | 08/22/2024 |  | 06/21/2024 |  |  | 06/15/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477371 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 75-year-old male in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the perfusionist reported that a pump stop occurred; despite several attempts by the perfusionist to re-start the pump, the pump did not start again. the perfusionist had already exchanged the automated impella controller (aic) but nothing happened. patient¿s last act was around 130sec. therefore, the impella was explanted. | [5 empty] | The investigation of pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. e4 should have been left blank g1 manufacturing site name and address h3 device evaluated by manufacturer: evaluation summary attached was inadvertently selected. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
25191 | 1220648-2024-12616 | 19584689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584689&pc=OZD | 06/16/2024 | 08/17/2024 |  | 06/21/2024 |  |  | 06/16/2024 | 06/21/2024 |  |  | 12/31/2024 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024322883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male of unknown race and ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the physician placed the pump far/deep into the left ventricle (lv) and the pump kinked fully back upon itself and was removed from the valve/ventricle to the aorta where it straightened. the team evaluated to be sure no lv perforation/effusion had occurred and then replaced the kinked pump with a new cp and support initiated without any noted harm or injury from malpositioning or delay. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage/kink and the positioning issue has been completed. data logs were reviewed which confirmed the pump was in improper position corresponded with the clinical description that triggered alarms, impella position in ventricle/aorta. no other issues were found on the logs. the product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. the root cause of kinked cannula was most likely excessive force. b.7 revised as information was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-12616. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12616 as it is unknown the initial reporter also sent the report to the fda. h.6 codes 4604 and 2981 were reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12616. a new code has been added.
25192 | 1220648-2024-12617 | 19584724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584724&pc=OZD | 06/16/2024 | 08/05/2024 |  | 06/21/2024 |  |  | 06/16/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025514619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male patient undergoing electrophysiology study (ep) with ablation was implanted with an impella cp device for mechanical circulatory support. it was reported that multiple attempts were made, but the pump was unable to advance due to the tortuous anatomy at the bifurcation of the aorta. the pump was removed from the left femoral artery with the intent to implant via the right femoral artery, but the device was unintentionally contaminated while outside of the patient. as a result, a new pump was prepped and implanted for patient support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, impella was aborted due to tortuous anatomy at bifurcation of aorta as per the clinical description provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12617: e4 should have been left blank. g1 manufacturing site address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25193 | 1220648-2024-12609 | 19584744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584744&pc=OZD | 06/15/2024 | 03/31/2025 |  | 06/21/2024 |  |  | 06/15/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1332033 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the associated automated impella controller (aic) does not trigger an acoustic signal in the event of an alarm. there was no reported patient harm, and the console was not exchanged. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - buzzer/ alarm malfunction has been completed. the impella device was not returned for evaluation. logs are not informative for this complaint. the failure mode could be neither confirmed nor denied due to the nature of the alleged malfunction. the hospital was reached out for return but the console was put into decommission by the hospital so has not been returned. the cause of the alleged lack of an alarm was not determined since the product was not returned. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-12609 was submitted in accordance with updated procedures. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12609 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-12609 was submitted in accordance with updated procedures. g.1 revised manufacturing site address line 1, manufacturing site address line 2, city and zip code as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-12609.
25194 | 1220648-2024-12610 | 19584829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584829&pc=OZD | 06/15/2024 | 08/06/2024 |  | 06/21/2024 |  |  | 06/15/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 29-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella was tangled in the mitral valve apparatus. the doctor attempted to reposition the pump; however, this was unsuccessful, and the pump was explanted and replaced. no patient harm was reported. | [5 empty] | The investigation for the positioning issue has been completed. complaint device not returned for investigation. the cause of the positioning issue most likely was use related due to improper technique and unable to de-tangle cannula from mv apparatus. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25195 | 1220648-2024-12619 | 19584940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19584940&pc=OZD | 06/17/2024 | 07/25/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1263046 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility biomedical department reported the automated impella controller (aic) had a purge disc not detected error. the console was replaced. no patient harm was reported | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | After reviewing the logs, the issues with the automated impella controller (aic) not being able to detect the purge disc was confirmed. the logs have evidence of the console having trouble detecting the purge disc (numerous purge disc not detected alarms. purge disc flag was found to be broken (missing at the blue purge disc retainer). the root cause of the observed issue was a broken purge disc flag. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12619: e4 should have been left blank in the initial report.
25196 | 1220648-2024-12618 | 19585018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19585018&pc=OZD | 06/11/2024 | 08/21/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC |  |  | US |  | 2025462844 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that bleeding developed around the impella access site between the repositioning sheath and the fixation ring. additional sutures were placed to achieve hemostasis. to treat the prolonged slow bleed, ten units of blood were transfused. support continued with the implanted pump following the intervention. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, bleeding developed around the impella access site between the repositioning sheath and the fixation ring. additional sutures were placed to achieve hemostasis as per the clinical description provided. d4-corrected model number h6 component code 4755 changed to 925 | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25197 | 1220648-2024-12614 | 19585057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19585057&pc=OZD | 06/16/2024 | 08/08/2024 |  | 06/21/2024 |  |  | 06/16/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025486124 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support while undergoing cardiopulmonary resuscitation. it was reported that the pump was placed but, flows were low. the team made choice to explant and replace the pump. the patient was reported as stable. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. data logs reviewed: impella positioning alarms occurred followed by suction alarm. the cause of low pump flow issue most likely was improper positioning related issue due to improper technique. e4 should have been left blank on manufacturer device report 1220648-2024-12614. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25198 | 1220648-2024-12620 | 19586686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19586686&pc=OZD | 05/22/2024 | 03/23/2025 |  | 06/21/2024 |  |  | 05/22/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442884 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hematuria | Epistaxis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and while on support, the patient had a nose bleed. anticoagulation was withheld and one unit of blood products was provided. additionally, the patient had hematuria and anticoagulation was withheld again. furthermore, there was a spike in aic (automated impella controller) fields noted. impella in aorta alarms were present with decreased flows, maps (mean arterial pressure), motor current and purge flows. the doctor felt that the patient was intravascularly dry and the septum was occluding the intake area. staff confirmed proper position and echo did not detect a thrombus. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of low pump flow and vascular damage - peripheral vessel has been completed. the data log review shows a motor current spike and a spike in purge pressure on (b)(6), indicating a potential clot or tissue interaction with the pump and purge gap. a large motor current spike is seen on (b)(6), immediately followed by frequent impella position in aorta alarms, a decrease in motor current pulsatility, and a drop in p level. flow rates still remain within expected range for the p level, although they have decreased. clinically, is it of note that this patient was not on anticoagulation. based on this evidence, it is likely the decreased flows are due to clot ingestion. hematuria was reported while on support and heparin was removed. the cause of the vascular damage peripheral vessel, hematuria, was not determined due to lack of clinical information and the relation to impella is unknown. b1 revised from the initial manufacturer device report number 1220648-2024-12620 to reflect both adverse event and product problem d4 model number was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2024-12620. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-12620 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-12620 in accordance with updated procedures. h6 codes revised from the initial manufacturer device report number 1220648-2024-12620 to reflect the omitted health effect - clinical code and health effect - impact code.
25199 | 1220648-2024-12623 | 19586907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19586907&pc=OZD | 06/17/2024 | 08/06/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  | 11/30/2025 |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425825 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was low pump flow observed. the p level was lowered to p5 and albumin and 2 units of packed red blood cells were administered. the suction was resolved. additionally, the patient had access site bleeding. the impella was removed due to the bleeding. the femoral artery was repaired with a patch and a dissection was noted to the illiac. a self expanding stent was used to repair the dissection. | [5 empty] | The investigation for the vascular damage, access site bleed, and the low pump flow has been completed. according to clinical details, bleeding was reported after a decision was made to leave peel-away sheath in place. the patient was taken to the or to remove impella. bleeding stopped after device removal. no product was returned. the data logs show suction alarms at various points during the procedure and placement signal trending downward during low pump flow instances. the root cause of the vascular damage - peripheral vessel was most likely use related due to improper technique. the root cause of the access site bleeding - major issue was most likely use related due to improper technique as described in clinical details. the root cause of the low pump flow was most likely patient condition related b1-checked product problem. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25200 | 1220648-2024-12624 | 19586958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19586958&pc=OZD | 06/17/2024 | 08/30/2024 |  | 06/21/2024 |  |  | 06/18/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481629 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a 70-year-old female in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the pump stopped a few minutes after support. therefore, the impella was explanted, and another was implanted without complications. the patient is still on support. | [5 empty] | The investigation for the pump stop has been completed. clinical details state that the cp pump stopped after a few minutes of support and could not be restarted. the pump was removed and replaced. no logs were returned. product was returned and evaluated. biomaterial was found in the purge gap, and blood was found in the purge line in the catheter near the motor housing. the pump had high purge pressure alarms while being run in the lab, and the pump stop was reproduced. the high purge pressure was isolated to the motor housing. therefore, the root cause of the pump stop was most likely biomaterial. added-d9 device return date g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25201 | 1220648-2024-12625 | 19587162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19587162&pc=OZD | 06/18/2024 | 08/21/2024 |  | 06/21/2024 |  |  | 06/18/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1210870 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility biomedical department reported the automated impella controller (aic) purge disc did not read the purge cassette. the console was as not in patient use and there was no harm or injury possible. | [5 empty] | D.9 updated as the console was returned for investigation. the investigation of automated impella controller (aic) - purge disc/flag issues has been completed. after reviewing the logs, the reported issue not being able to read the catheter purge could not be confirmed. however, there was evidence of the console having issues reading the catheters. there were numerous instances of impella defective (error reading eeprom) found. the reported occurrence date was listed as 2024.06.18 but no logs were found from that day (occurrence date was likely misreported). the automated impella controller was tested. the console was tested in the lab and was able to reproduce the eeprom reading issue. the pump was unable to be detected and a yellow impella defective prompt appeared on the consoles interface. the root cause of the consoles inability to read the catheters eeprom was determined to be caused by a defective impellatronic pca. e.1 revised frist name as it was entered incorrectly on initial manufacture device report number. h.6 code 2917 was entered incorrectly on initial manufacture device report number. a new code was added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25202 | 1220648-2024-12621 | 19587344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19587344&pc=OZD | 06/17/2024 | 09/03/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489812 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp had a high purge pressure alarm. lysis therapy was applied and the purge pressure when down. the purge pressure high was persistent again and the impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. clinical details states that on the second day of support, a high purge pressure alarm was reported. there were no kinks in the purge line and tissue plasminogen activator (tpa) infusions did not resolve the complication. the pump was explanted due to the failure mode. data logs were not returned for investigation. returned product was investigated and biomaterial was wrapped around the impeller and covering the purge gap. when purging the pump in the lab, hpp reproduced. cutting the catheter just proximally of the motor confirmed that no blood had ingressed into the purge line. additionally, when the cut was made in the catheter, while still purging, the high purge pressure resolved immediately. this confirmed that the root cause of the high purge pressure was determined to be biomaterial in the purge gap. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi. added- d9 device return date, h6 impact code 4644 g1-corrected mfr site name and address.
25203 | 1220648-2024-12626 | 19588201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19588201&pc=OZD | 05/24/2024 | 07/16/2024 |  | 06/21/2024 |  |  | 05/24/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450903 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was implanted with an impella 5.5. device for mechanical circulatory support. the patient was admitted for cardiogenic shock and given coronary artery bypass surgery (left internal mammary artery to the left anterior descending artery, saphenous vein grafts to the obtuse marginal branch and saphenous vein grafts to ramus intermedius) and impella support. on day ten of the support, there were signs of hemolysis. plasma-free hemoglobin levels were high and continuous renal replacement therapy was performed. the patient condition deteriorated during the support and a mitraclip was done. additionally, the patient suffered a cerebral vascular accident. however, the impella 5.5 pump remained on support for over 27 days when the patient had care withdrawn and thereafter expired. medical safety review was completed and noted there is not enough information to associate the use of the impella device with the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for hemolysis and stroke has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
25204 | 1220648-2024-12622 | 19588537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19588537&pc=OZD | 06/17/2024 | 08/13/2024 |  | 06/21/2024 |  |  | 06/17/2024 | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509592 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support and there was access site bleeding. angle matching was performed, manual pressure applied, the site was redressed and one unit of packed red blood cells was administered. heparin was also withheld. additionally, the impella was frequently migrating up to the aorta. the device was repositioned multiple times and left deep to stabilize position. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the positioning issue has been completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. based on clinical description, the root cause of access site bleeding was most likely due to patient condition.
25205 | MW5156485 | 19594919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19594919&pc=OZD | 04/22/2024 | 06/12/2024 |  | 06/21/2024 |  |  |  | 06/21/2024 |  |  |  |  | 06/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. | [6 empty] | 1000100 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [8 empty] | Male | 105 KG | Non Hispanic | White | Injury | Yes | Yes | [7 empty] | Insufficient Information | [12 empty] | Hospitalization | Life Threatening | [27 empty] | Initial submission | [9 empty] | Patient experienced a significant pull of the impella 5.5 driveline causing a separation of the power wires within the driveline and subsequent device failure. patient was emergently taken to operating room for pump exchange. concern is that the device is not manufactured to be as durable externally as anticipated and what should be expected compared to similar devices with external leads. due to the nature of the event and subsequent findings, the exact location of the break within the driveline may point to a specific area of concern with current manufacturing process and durability testing.
25206 | 1220648-2024-12553 | 19574210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574210&pc=OZD | 05/26/2024 | 09/25/2025 |  | 06/20/2024 |  |  | 05/27/2024 | 06/20/2024 |  |  | 01/31/2026 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454960 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that a discrepancy was observed between the pump's placement signal and the patient's arterial line pressure. | The user facility reported an unknown age patient with post cardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support the experienced oozing from femoral impella access site around the sheath. the patient was given one unit of red blood cell concentrate and was on anti-coagulants. manual pressure was used to successfully achieve hemostasis. |  |  |  |  | The investigation for the access site bleed has been completed. pump was returned for evaluation. no abnormalities were found upon visual inspection. pump 5s parameters were found to be within range. pump was run in a bucket of water for approximately 5 minutes and no abnormalities were observed. clinical details patient had thrombocytopathy with a very high risk of bleeding unrelated to impella. the root cause of the access site bleeding - major was most likely due to patient's thrombocytopathy condition. g1-corrected reporting contact facility name and address. g1-corrected mfr site name and address. h6 component code 4755 changed to 925. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was returned for evaluation. no abnormalities were found upon visual inspection. pump 5s parameters were found to be within range. pump was run in a bucket of water for approximately five minutes, and no abnormalities were observed. clinical details noted that a discrepancy was observed between the pump placement signal issue and patient's arterial line pressure. the root cause of the optical signal issues was most likely due to patient condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. corrections that were made from the initial manufacturer device report number. b7 other relevant history, including preexisting medical conditions, updated. d10 concomitant medical products and therapy dates updated. e4 was erroneously marked "no" in the initial report. g1 reporting contact email updated. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25207 | 1220648-2024-12554 | 19574261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574261&pc=OZD | 05/27/2024 | 07/23/2024 |  | 06/20/2024 |  |  | 05/27/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male patient in acute myocardial infarction/cardiogenic shock patient suffered a stroke during impella cp support. the patient was unable to move their right upper extremity, which led to discovery that the patient suffered an ischemic infarction with a midline shift. support continued with the implanted pump following the diagnosis. the decision was made to withdraw care and the patient expired. per the medical safety officer review there is not enough information to associate the use of the device with outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.2 selection was revised due to health effect - impact code being inadvertently omitted on the initial manufacturer device report 1220648-2024-12554. b.7 revised as was inadvertently omitted from the previously submitted manufacturer device report 1220648-2024-12554. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer device report 1220648-2024-12554. h.6 code 4755 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12554. a new code was added to component codes. h.6 a health effect - impact code was added as it was inadvertently left off the previously submitted manufacturer device report number 1220648-2024-12554. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report manufacturer device report number 1220648-2024-12554 in accordance with updated procedures.
25208 | 1220648-2024-12555 | 19574321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574321&pc=PYX | 05/30/2024 | 10/17/2025 |  | 06/20/2024 |  |  | 05/30/2024 | 06/20/2024 |  |  | 11/30/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024434864 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella rp flex was prepped and primed however, because of concerns for malfunctioning optical sensor, the pump was removed and a new pump was used for the case. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D.9 updated as the pump and purge cassette were returned for investigation. the investigation has been completed. the impella device was received for investigation. during the case, the automated impella controller placement pa signal read 157 (157/157). the pump was replaced out of caution. the data logs showed no alarms and the 5s parameters were within normal range. the product was inspected and no abnormalities were found. the 5s parameters were within normal range and the pump passed the light continuity test. pa reading was as reported most likely due to the pump not being in the body. there was no problem found as the root cause of the optical signal issue as no abnormalities were found and the device functioned within specification. d.4 model number was revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12555. e.1 added name prefix/title, first name and last name as this information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12555. e.3 revised since initial submission of initial manufacturer device report number 1220648-2024-12555 to reflect the occupation of the name added in section e.1 e.4 this field should have been left blank on the initial manufacturer device report number 1220648-2024-12555 as this information is unknown. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12555. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
25209 | 1220648-2024-12556 | 19574438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574438&pc=OZD | 06/01/2024 | 07/22/2024 |  | 06/20/2024 |  |  | 06/01/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492139 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.5 support had a cva (cerebral vascular accident). the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date was reported incorrectly in the initial report.
25210 | 1220648-2024-12557 | 19574546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574546&pc=OZD | 06/07/2024 | 07/31/2024 |  | 06/20/2024 |  |  | 06/07/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp was accidentally pulled into the aorta while transporting the patient to another facility. the patient was taken to the surgical suite to attempt a new impella placement. there is no additional information on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue most likely was patient movement as the impella accidentally pulled into aorta during transport.
25211 | 1220648-2024-12558 | 19574605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574605&pc=OZD | 06/11/2024 | 08/02/2024 |  | 06/20/2024 |  |  | 06/11/2024 | 06/20/2024 |  |  | 07/31/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440333 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Retroperitoneal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp procedure oozing at the cp access site was noted. sutures were replaced and properly angle matched. once prepped to move to bed bleeding noted to site again. improved with manual pressure. it was further reported that cp was escalated to impella 5.5, and retroperitoneal bleeding (rp) developed from previous cp site surgical cutdown. the patient was taken to or and placed covered stent and repaired bleeding area. the patient is stable. | [5 empty] | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding issue was most likely use related, based on given clinical details. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12558: e4 should have been left blank h.6 code 4755 was reported incorrectly. new code was added to component code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
25212 | 1220648-2024-12559 | 19574623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574623&pc=PYX | 05/23/2024 | 05/14/2025 |  | 06/20/2024 |  |  | 05/23/2024 | 06/20/2024 |  |  | 06/30/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024414385 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported when preparing an impella rp for implant to an unknown age patient, the plug pins broke off when inserting the white pump plug into the connection cable (gray end). there was no report of patient harm. | [5 empty] | A.5 ethnicity and race are unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Additional information provided in d10, h3, h6 and h10. the investigation into the reported issue has been completed. pump passed all post sterile inspection checks. clinical states that the pins of the grey connector were damaged during the setup. pump was returned for investigation. the pins of the grey connector were confirmed as damaged. no other abnormalities were observed. the damage was similar to that of improper attempt to connect the grey connectors which resulted in pins being bent and damaged. therefore, the root cause of the pump not detected issue was broken pins of the grey connector due to improper technique.
25213 | 1220648-2024-12560 | 19574677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574677&pc=OZD | 06/11/2024 | 07/31/2024 |  | 06/20/2024 |  |  | 06/11/2024 | 06/20/2024 |  |  | 09/30/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402796 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient during the impella cp procedure developed a hematoma to the groin. the hematoma was expressed, and the site was resutured and angle matched. the patient is stable. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by resuturing and angle matching. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12560 as it is unknown if the initial report also sent the report to the fda. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12560. a new code was added to component codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report 1220648-2024-12560 in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25214 | 1220648-2024-12561 | 19574787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574787&pc=OZD | 05/24/2024 | 07/16/2024 |  | 06/20/2024 |  |  | 05/24/2024 | 06/20/2024 |  |  | 03/31/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327151 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient suffered bilateral limb ischemia and the team had the purge changed from heparin to bicarb. the team lowered the pressors due to loss of pulses. the patient was supported until care was withdrawn per the family's decision and the patient expired. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device from the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report.
25215 | 1220648-2024-12562 | 19574877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574877&pc=OZD | 05/23/2024 | 07/11/2024 |  | 06/20/2024 |  |  | 05/24/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post percutaneous coronary intervention. during support, the patient developed limb ischemia. a femoral-femoral bypass was performed. the family withdrew care and the patient expired. upon review by abiomed's medical safety officer, there is not have enough information to associate the use of the device from the outcome. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added-h6 component code 925 f6/f8 should have been blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿
25216 | 1220648-2024-12563 | 19574936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19574936&pc=OZD | 06/03/2024 | 08/12/2024 |  | 06/20/2024 |  |  | 06/03/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025473394 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (race, age and gender unknown) was implanted with an impella 5.5 device for mechanical circulatory support. the primary indication is unknown. during support, a high purge pressure alarm was displayed. the pump was explanted. a replacement pump was recommended, however the file does not state a replacement was used. no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: high pressure purge impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. the pump was returned for investigation. data logs shows 30-minute gradual rise in purge pressure with noted purge pressure high alarms. flow saturation was noted during purge pressure rise. no placement signal trend or motor current spike was seen during this event. significant amount of biomaterial was observed on the impeller and purge gap. high purge pressure was reproduced during testing. the pump was cut near the motor housing to check for any blockage in the purge line. purge was able to be seen coming out of the catheter after cutting the catheter. as per the clinical report, the doctor troubleshot high purge pressure by changing the purge cassette. purge pressure remained high and was exchanged. the cause of the high purge pressure issue was biomaterial based on returned product investigation.
25217 | 1220648-2024-12564 | 19575009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575009&pc=PYX | 05/24/2024 | 07/17/2024 |  | 06/20/2024 |  |  | 05/24/2024 | 06/20/2024 |  |  | 10/31/2024 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274751 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support post percutaneous coronary intervention (pci). the patient was also placed on an intra-aortic balloon pump. the impella rp access site was bleeding and to treat, the patient had sutures and blood products delivered. the team additionally placed a fem-stop. the tamponade from the coronary pci had caused the patient deterioration and the patient had a ventricular fibrillation arrest and care was withdrawn. the patient expired. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device from the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since additional 4 different mattress and figure 8 sutures were placed to try to stop the bleed. added-h6 impact code 925 f6/f8 should have been left blank in the initial report.
25218 | 1220648-2024-12565 | 19575056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575056&pc=OZD | 06/12/2024 | 08/06/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483244 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella 5.5 implant and once support was initiated, motor current was still flat and lower than expected flows on impella. the patient is on va ecmo (venoarterial extracorporeal membrane oxygenation) currently. it did not look like the 5.5 ingested a clot. visualized impella on tee (transesophageal echocardiography) and nothing was noted to be blocking or interfering with inlet or outlet. the p-level was increased but got suction with anything above p4. a decision was made to optimize patient by balancing ecmo and impella flows while giving volume. impella is in a good position. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. the impella was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely patient condition. added- corrected model number.
25219 | 1220648-2024-12566 | 19575065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575065&pc=OZD | 06/12/2024 | 08/05/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448874 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a leak was noted on the yellow to clear, yellow purge connection to the reservoir, filter on impella cp. it was unsure of when the leak began. it was not noticed in the cath lab or on initial transfer to the or. there were no issues until after the cabg (coronary artery bypass grafting) was started and the purge low alarm sounded. the purge cassette was changed, and alarm continued. a successful purge reservoir/filter bypass was completed. there was no interruption in impella support during the entire cabg case. no further alarms noted. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak has been completed. the root cause of the pump purge leak was a leak from the sidearm but the exact location of the leak was unknown since the product was not returned. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12566: d4 model number. e4 should have been left blank.
25220 | 1220648-2024-12567 | 19575142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575142&pc=OZD | 06/12/2024 | 08/07/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025471208 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella 5.5 was defective upon connecting the pump. there was an alarm, "disconnect and replace impella". a new pump was used successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the pump not detected has been completed. the pump was returned for investigation. the data log was not provided for the case run. the pump was connected to the console without issue, and the case status screen was seen. the pump successfully passed the electrical test and optical signal test. no issue was found in reading the eeprom on the console. no past usage was reported upon extracting eeprom data. there was no issue found during the case. the device functioned/operated to specification. there was no issue found during the case. the device functioned/operated to specification. added-d9 device return date d4-corrected model number.
25221 | 1220648-2024-12568 | 19575163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575163&pc=OZD | 06/12/2024 | 08/05/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  | 11/30/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432769 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella cp implant and the impella came across aortic valve while attempting to recross impella folded in on itself. the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. data logs confirmed pump was in aortic area (about 45 minutes into the case). pump then was stopped manually and disconnected from the console. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely due to attempt to re-cross the valve wirelessly per instructions for use the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12568.
25222 | 1220648-2024-12569 | 19575191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575191&pc=OZD | 06/12/2024 | 07/25/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481074 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and they had access site bleeding. anticoagulation was withheld, no blood products were given, pressure was applied and new dressing was applied. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12569. d4 model number e4 should have been left blank.
25223 | 1220648-2024-12570 | 19575208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575208&pc=OZD | 06/12/2024 | 07/31/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and they had access site bleeding. manual compression was held, angle matching was appropriate, and dressing changed, then femostop placed. femostop adjusted, then removed and manual compression held with decreased bleeding, but still continued to ooze. the impella was then explanted. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to anticoagulation management since the activated clotting time was 258 in the intensive care unit which is above the range limit. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as information was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-12570. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12570 in accordance with updated procedures. e.4 should have been left blank on the initial of manufacturer device report number 1220648-2024-12570 as it is unknown if the initial reporter also sent the report to the fda. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer device report number 1220648-2024-12570. g.2 foreign was selected incorrectly on manufacturer device report number 1220648-2024-12570. h.3 revised selection for was the device evaluated by the manufacturer since the initial submission of manufacturer device report number 1220648-2024-12570 as the device was discarded. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12570. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12570 in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25224 | 1220648-2024-12572 | 19575338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575338&pc=OZD | 06/12/2024 | 08/15/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469487 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 29-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp would not secure and kept saying purge blocked during insertion. the cp was removed and replaced with another. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. clinical reported the pump would not secure and kept saying purge blocked. the data logs confirmed 'purge pressure - high' and 'purge system blocked' alarms. about five minutes after implant, there was a motor current spike followed by a gradual rise in purge pressure which did not resolve until the pump was explanted. the root cause of the high purge pressure was most likely biomaterial ingestion. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12572 in accordance with updated procedures. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12572 as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12572 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
25225 | 1220648-2024-12574 | 19575494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575494&pc=OZD | 05/25/2024 | 07/25/2024 |  | 06/20/2024 |  |  | 05/25/2024 | 06/20/2024 |  |  | 09/30/2025 |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408357 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The user facility reported a 24-year-old female patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp was implanted. shortly after the patient coded and cardiopulmonary resuscitation (cpr) was performed. the impella cp was implanted a little deep because when they pulled it back there seemed to be an unusual bend in the cannula. the impella cp was left at p-2 as cpr was being performed for over an hour. extracorporeal membrane oxygenation was placed. impella flows were run at p-1 because there was suction when running at p-2. the patient was also bleeding from multiple sites and endo-tracheal tube. multiple units of blood were given and hemoglobin (hgb) never came up. also, the physicians were unable to get an arterial line, so it was decided to remove the impella cp and use that arterial access for the arterial line. despite all best efforts and multiple hours of cpr, the patient expired. per abiomed's medical safety officer review, there is not enough information to associate the use of the device with outcome. | [5 empty] | A.5 ethnicity and race are unknown. the impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the low pump flow, positioning issues, and vascular damage/bleeding has been completed. the pump was returned for investigation. there was a cannula kink observed. low flow could not be reproduced. clinical mentions that kink was straightened out, patient was bleeding from multiple sites and cpr was given. ecmo device was also used along with impella. data logs were reviewed and there were constant suction alarms and position alarms were also seen in the case. there was upward shift in the placement signal observed. the patient had a very deteriorating condition. per the data log review and clinical information provided, the root cause of the low pump flow issue was patient condition related. the patient was given cpr for multiple hours. therefore, as per the clinical details provided, the root cause of the positioning issue was use related due to cpr given to the patient. the patient was bleeding from multiple sites, but no other information was provided. therefore, the root cause of the vascular damage - peripheral vessel issue/bleeding was not determined since insufficient clinical information was provided and unknown relation to impella. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ f6/f8 should have been left blank in the initial report.
25226 | 1220648-2024-12575 | 19575507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575507&pc=PYX | 06/12/2024 | 08/06/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025480532 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that hemolysis was noted in the patient on impella rp support. rp reposition sheath pulled back approximately 2 inches from implant. it was further reported that the rp was removed due to hemolysis. | [5 empty] | The investigation for the hemolysis has been completed. according to clinical details, hemolysis was noted on the third day of support during 5.5 placement. the impella was returned for evaluation. upon visual inspection, no biomaterial was observed on the returned product. the pump passed hemolysis testing. the clinical states that the repo sheath had been pulled back 2 cm since implant and kub showed the inflow to be low. the repo sheath was not properly secured. data logs show no motor current (mc) spikes outside of p-level changes. prior to the reported hemolysis, placement signal (ps) and central venous pressure (cvp) are both noisy with negative cvp indicating possible suction conditions. the root cause of the hemolysis was most likely improper positioning of the device. added- d8 updated to yes d4-corrected model number. | The impella device was received from the customer on (b)(6) 2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25227 | 1220648-2024-12573 | 19575564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19575564&pc=OZD | 06/11/2024 | 07/29/2024 |  | 06/20/2024 |  |  | 06/12/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471212 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 24-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 sterile sleeve was torn away from the t-lock on both ends, distal to the patient and at the blue suture hub. the sleeve was torn during a repositioning attempt. the sleeve was covered with sterile dressing and will be monitored. there was no harm to the patient. | [5 empty] | The investigation for sterile sleeve damage has been completed. the impella device was not returned for evaluation. the root cause of the sterile sleeve damage issue was most likely use related as it occurred during a repositioning event. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12573: d4 model number: e4 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25228 | 1220648-2024-12580 | 19578463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19578463&pc=OZD | 05/26/2024 | 07/29/2024 |  | 06/20/2024 |  |  | 05/26/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1172896 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | The complainant reported a console alarm/failure occur after removal. the console displayed a controller failure. there was no noted harm or injury. | [5 empty] | A.1 and a.5 are unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for automated impella controller (aic) - controller error has been completed. the console was received for investigation. the cause of the issue was not determined since issue could not be reproduced. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12580: f6/f8-should have been left blank.
25229 | 1220648-2024-12581 | 19579276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19579276&pc=OZD | 06/04/2024 | 08/13/2024 |  | 06/20/2024 |  |  | 06/04/2024 | 06/20/2024 |  |  |  |  | 06/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457971 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a 58-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient experienced malperfusion of leg and needed escalation of care. the pump was explanted, and an additional impella device was needed. | [5 empty] | A.5 ethnicity and race are unknown. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added-h6 component code 925 f6/f8 should have been left blank in the initial report.
25230 | 1220648-2024-12517 | 19565915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19565915&pc=OZD | 06/08/2024 | 08/01/2024 |  | 06/19/2024 |  |  | 06/08/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504356 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Tear, Rip or Hole in Device Packaging | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during repositioning of the impella cp the sterile sleeve was torn. it appeared possible that the device was not removed from the anchoring system prior to repositioning. the site is intact with angle-matching gauze and leg immobilizer. the cp has been successfully weaned and removed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage (sterile sleeve damage) has been completed. no product was returned for investigation. clinical information stated that the pump was likely repositioned without detaching from the anchor. the patient also had a knee embolizing on the impella leg. the root cause of the sterile sleeve damage was most likely use related. b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12517. e.4 should have been left blank on the initial submitted manufacturer device report number 1220648-2024-12517 as it is unknown if the initial reporter also sent the report to the fda. h.6 code 2199 and 2385 were entered incorrectly on the initial manufacturer device report number 1220648-2024-12517. new codes were added.
25231 | 1220648-2024-12518 | 19565922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19565922&pc=OZD | 06/09/2024 | 08/05/2024 |  | 06/19/2024 |  |  | 06/09/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484487 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 66-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported immediately after the implant, the cp showed saturation with motor currents ranging from 1000-1100, flows of 4.3/4.3, and high purge pressure. the automated impella controller showed flipped white left ventricle waveforms and red aorta waveforms. activated clotting time was only 200 seconds upon pump insertion. medical staff removed the cp and replaced it with another cp without issue. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of saturated flow has been completed. the impella device was not returned for evaluation. the cause of the saturated flow issue was most likely biomaterial per log analysis and clinical information. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12518: e4 should have been left blank in the initial report.
25232 | 1220648-2024-12519 | 19565932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19565932&pc=OZD | 06/10/2024 | 08/05/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479309 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Detachment of Device or Device Component | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient had hemolysis. additionally, there were also low pump flows noted and the patient received blood products. upon explanting, the doctor noted a small amount of biomaterial in the inlet cage. the doctor wondered if the biomaterial in the inlet could have caused the patient's hemolysis. furthermore, upon explanting the pump was in the graft, and the doctor had the impella clamped around the motor housing near the cannula. this caused the pump to separate motor housing from cannula. the patient was taken to the cardiovascular or and the physician was able to manipulate the catheter which allowed the pump to withdraw into the graft. the impella was explanted in normal surgical fashion. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions. including catheter position, pre-existing medical conditions, and small left ventricular volumes. may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis, product damage, and low pump flow has been completed. pump met resistance when attempting to pull back across the valve during explant. the physician clamped down on the cannula causing a detached inflow cage when attempting to remove pump. patient also experienced slight hemolysis. additionally, there was low pump flow noted and patient was given blood products. no data logs were returned. only the cannula with detached inflow cage was returned. evidence of clamping was confirmed on the pump aligning with clinical details. the root cause of the hemolysis was not determined as insufficient product was returned and insufficient clinical information was provided. the root cause of the product damage (detached inflow/outflow - great vessel) was determined to be use-related as evidenced by clamping. the cause of the low pump flow was most likely biomaterial since biomaterial was observed on the inflow upon explant. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12519: e4 should have been left blank.
25233 | 1220648-2024-12520 | 19565937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19565937&pc=OZD | 06/10/2024 | 08/21/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491041 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during support, the impella cp had a pump stop. the device was unable to be restarted. decision was made to explant the pump after the patient being on support for 10 days. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the pump stop has been completed. impella cp pump was returned. visual inspection was performed finding the pump and large purge gap was observed. no other abnormalities were found. bearing wear was confirmed with large purge gap. pump stop occurred on 10th day of support. data log analysis indicated high motor current pump stop with several #13 alarms at the end of the case on 10th day of support. the cause of the pump stop was bearing wear beyond design life period. the impella cp optical device is not indicated for support more than 8 days of design life period. added- d9 device return date, h6 impact code 4627 e1-corrected initial reporter first/last name and email.
25234 | 1220648-2024-12521 | 19566345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566345&pc=OZD | 05/29/2024 | 07/29/2024 |  | 06/19/2024 |  |  | 05/29/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488223 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during preparation for implant of an impella cp device for mechanical circulatory support to a patient experiencing acute myocardial infarction/cardiogenic shock, it was reported that the impella cp had a failure upon prepping. the impella cp was exchanged. there was no report of harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump not detected issue has been completed. pump was returned for evaluation. upon visual inspection, no issues were found with the pump. electrical testing was done and all motor cable lines were found to be in range. pump was plugged into 4 consoles (aic) and 4 out of 4 aics prompted case start for the pump. data logs show that aic had repeated error with eeprom unable to read calibration data resulting in a "impella defective" alarm. there was no problem found with the impella cp pump. upon review of the data logs and pump, it is apparent that the console is the potential cause of the issue. the investigation for the console is found in regulatory report #1220648-2024-12501.
25235 | 1220648-2024-12522 | 19566415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566415&pc=OZD | 06/02/2024 | 07/23/2024 |  | 06/19/2024 |  |  | 06/02/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461390 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed a hematoma at the access site. bivalirudin was discontinued and the patient remained on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added f6/f8-should have been left blank in the initial report. g1 reporting contact fax number.
25236 | 1220648-2024-12523 | 19566543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566543&pc=OZD | 05/30/2024 | 08/07/2024 |  | 06/19/2024 |  |  | 05/30/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488783 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown age, gender and race) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed "system error optical sensor", leading to console and pump exchange. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the logs confirmed the occurrence of "controller error (optical bench)" alarms. an optical fiber break was found inside the red handle near the patient cable connection. the other wires coming out of the patient cable were significantly twisted where the fiber break was found. there were no abnormalities with the bonds. no other abnormalities were found. the root cause of the optical signal issue was a damaged optical fiber line stemming from the internal manufacturing process. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12523: e4 should have been left blank. g1 manufacturing site name and address.
25237 | 1220648-2024-12524 | 19566546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566546&pc=OZD | 06/07/2024 | 07/31/2024 |  | 06/19/2024 |  |  | 06/07/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447773 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed a hematoma at the impella access site, and one unit of packed red blood cells was given. the decision was made to explant the pump due to the hematoma and due to the patient¿s large body habitus, the physicians decided it would be safer to have vascular surgery, do a cutdown and remove the pump surgically. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12524. e4 should have been blank. h.6 code 4755 was reported incorrectly. new code was added to component code.
25238 | 1220648-2024-12525 | 19566590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566590&pc=OZD | 06/03/2024 | 07/30/2024 |  | 06/19/2024 |  |  | 06/03/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025480206 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a 76-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the attempted insertion lasted two hours without success. the medical staff removed the pump and a new pump was successfully used. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since they were successful with implanting the other device. f6/f8 should have been left blank in the initial report.
25239 | 1220648-2024-12526 | 19566821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566821&pc=OZD | 05/30/2024 | 07/24/2024 |  | 06/19/2024 |  |  | 05/30/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025504327 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old female patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the patient turned to the side for a skin assessment and a post positioning hematoma was noted. the physician angel matched the sheath and re-sutured in place and manual pressure was applied. it was also reported that there was a possible soft tissue infection in the access site. oral antibiotics were given for five days as treatment. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the root cause of the access site bleeding was most likely due to use issue since the sheath was angle matched and re-suture was placed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings and cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25240 | 1220648-2024-12527 | 19566859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19566859&pc=OZD | 06/01/2024 | 11/19/2024 |  | 06/19/2024 |  |  | 06/01/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1440851 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the purge disk was not recognized by the automated impella controller (aic). while cleaning, the nurse noticed that the blue purge disk holder on the aic had broken off. before the aic could be replaced, the pump stopped. the pump was ultimately, successfully explanted, and the patient outcome as a result of the event was noted as stable. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. note: this is one of two related events involving the automated impella controller and impella cp pump. | The investigation for the reported console purge disc/flag issues has been completed. the console was not returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the root cause for purge pressure alarms was a broken purge retainer based off the clinical staff reporting and logs. corrections to the initial mfr report:
25241 | 1220648-2024-12528 | 19567055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567055&pc=OZD | 06/02/2024 | 07/16/2024 |  | 06/19/2024 |  |  | 06/02/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1647627 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | 505 | Foreign | Company representation | Distributor |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received an automated impella controller (aic) with a note stating there was a red alarm-battery failure. the batt test showed a cell imbalance in battery two. there was no report of patient involvement or harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) battery failure alarm has been completed. data logs were reviewed finding that the reported red battery failure issue was confirmed. there were numerous instances of battery failure alarms (transport connection blown/missing) (event set #360), battery failure alarm (low voltage) [v < 12v] (event set #350), and battery level low (50%) (event set #23) found in the imc logs. the console was returned for evaluation. during testing, the battery failure issue was able to be reproduced. results of running batttest on telnet revealed that vcell4 in battery2 had a voltage that was significantly less than the rest of the cells in the battery. reported battery failure issue was determined to be caused by internal battery cell imbalance (found in vcell4 of battery2). added-d8 selected device returned to manufacturer d4-corrected model number. g1-corrected mfr site name and address.
25242 | 1220648-2024-12529 | 19567091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567091&pc=OZD | 06/03/2024 | 07/24/2024 |  | 06/19/2024 |  |  | 06/03/2024 | 06/19/2024 |  |  | 01/31/2025 |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314879 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an impella cp device was used for mechanical circulatory support in a patient following an acute myocardial infarction/cardiogenic shock. during impella support, bleeding occurred at the impella femoral access site due to a loose suture on the repositioning sheath. the repositioning sheath slipped out three millimeters. the bleeding was stopped by re-suturing, and no further bleeding occurred. the day after the incident, the impella cp was removed without further problems. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to use issue since suture loosened and was managed by re-sutured in upright position. instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12529: e4 should have been left blank.
25243 | 1220648-2024-12530 | 19567135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567135&pc=OZD | 06/04/2024 | 08/01/2024 |  | 06/19/2024 |  |  | 06/04/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025511373 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp pump became dislocated, leading to pump being explanted, and a new pump placed. there was no patient harm reported. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12530: d4 model number. e4 should have been left blank. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25244 | 1220648-2024-12531 | 19567184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567184&pc=OZD | 06/05/2024 | 08/12/2024 |  | 06/19/2024 |  |  | 06/05/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025488322 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old female patient of unknown race and ethnicity noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that during implant procedure due to the 75-year-old male patient's anatomy, it was not possible to place the 0.018. the insertion of the impella cp was therefore aborted. the patient had a percutaneous coronary intervention (pci) without support successfully performed and there was no direct issue with the pump or wire. the patient was stable. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, due to patient's anatomy, it was not possible to place the 0.018, the insertion of the impella cp was therefore aborted. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ d4-corrected udi number. g1-corrected mfr site address line 1, city, and zip code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: warnings to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25245 | 1220648-2024-12532 | 19567223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567223&pc=OZD | 06/06/2024 | 07/23/2024 |  | 06/19/2024 |  |  | 06/06/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1763601 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had electrical issues with buttons. the patient remained on support with this aic. it was further reported that airplane button was pulsating without a cloud light. turning the console off and back on, medical staff tried to put the aic in flight mode, but the aic continued to do the pulsating airplane light. medical staff ensured the usb connection was secure and tried to perform a 60 second factory reset press without success. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Console logs from the reported day of event did not show any entries on that day, and entries on (before) and (after) indicate that rlm was powered on, and data was transferred from automated impella controller (aic) to rlm. rlm journals obtained after troubleshooting indicated that rl10620 lost hospital wifi configuration and prior to that experienced multiple reboots for no apparent reason. when the aic and rlm were tested in danvers, the reported issue was reproduced, and rlm was initially unable to boot up, with the airplane button flashing rapidly. serial terminal connection to rl10620 showed no activity (not booting up) at that time. inspection of the unit revealed improper routing of vga wires that was causing strain (perpendicular to the pca plane) on miniusb connector, resulting in intermittent power delivery (via usb) to the rlm. once the cables were rerouted, the rlm was able to boot-up. serial terminal data from consecutive boot-ups showed evidence of partition switching due to corrupt microsd card. rlm journal was then downloaded, and also showed that rl10620 had obsolete som sw and firewall fw (unrelated to this complaint). repair: replace backstrap cable as a precaution (assuring proper routing of usb and vga wires to minimize strain), perform functional check of the console, including fcn 105 (note that for fcn 105 the firewall firmware needs to be upgraded as well), and test rlm as per 0042-7309 sec. 24. if the customer wifi configuration cannot be restored, inform impellaconnect team that rl10620 will need to be reconfigured at the site. the cause of the aic issue was a work instruction / technician error. d9 date device returned to manufacturer added e4 should have been left blank in the initial report. g1 reporting contact fax number.
25246 | 1220648-2024-12533 | 19567747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567747&pc=OZD | 06/07/2024 | 08/15/2024 |  | 06/19/2024 |  |  | 06/07/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025490860 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a male patient of unknown age, race, and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped and attempt to re-start was unsuccessful. the pump was switched off and pulled into the aorta. the patient is stable. | [5 empty] | The investigation for the pump stop has been completed. the impella was returned for evaluation. upon visual inspection, the pump had a large purge gap. field data log inspection confirmed that alarm 115 was triggered, however, motor current spikes only reached 1481 ma. additionally, there were two distinct motor current spikes in the logs leading up to the pump stop. the root cause of the pump stop was determined to be bearing wear due to operation of the pump outside of the specified design life. added- d9 device return date g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: impella stopped. ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
25247 | 1220648-2024-12534 | 19567784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567784&pc=OZD | 06/07/2024 | 08/16/2024 |  | 06/19/2024 |  |  | 06/07/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453323 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 failed to prime. the automated impella controller (aic) skipped the screen instructing to ¿connect yellow to yellow¿ and would go straight to the placement screen "ready for insertion" without priming. a new pump was used successfully. there was no patient harm. | [5 empty] | D.9 updated as the pump and purge cassette were returned for investigation. the investigation for the priming issue has been completed. during preparation to insert the impella, the medical team did not see the instruction to connect the yellow luer components on the purge line. it was not resolved by restarting the case. the product was returned for investigation. no data logs were returned, so the team could not investigate whether any alarms went off during the priming steps. in the lab, the failure mode did not reproduce. the purge was primed successfully, the prompt to connect yellow to yellow did come up on the automated impella controller, and the pump was able to run with normal purge pressures and flow rates for several minutes. the root cause of the priming issue was unable to be determined since the issue did not reproduce and no logs were returned for investigation. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12534 as it is unknown the initial reporter also sent the report to the fda. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12534 was submitted. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2024-12534. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25248 | 1220648-2024-12535 | 19567847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567847&pc=OZD | 06/04/2024 | 03/10/2025 |  | 06/19/2024 |  |  | 06/07/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482843 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support post heart transplant rejection. during support, the patient developed a hematoma at the axillary site. the patient was taken to the operation room for a washout. the patient is stable. | Additional information was received that the patient had consistent ventricular fibrillation. the patient was shocked multiple times and was underwent cardioversion to restore to normal sinus rhythm. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B3 event date. b5 updated with the additional information received. h.6 health effect - clinical code and health effect - impact code updated to reflect new information provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12535. g1 reporting contact fax number missing. e4 should have been left blank. | The investigation of access site bleeding and vt/vf/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-12535 in accordance with updated procedures.
25249 | 1220648-2024-12536 | 19567908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567908&pc=OZD | 06/08/2024 | 07/26/2024 |  | 06/19/2024 |  |  | 06/08/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485605 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was sent to the icu after placing the impella successfully but shortly after, developed hypovolemic shock due to bleeding complications secondary to coagulopathy. the patient developed a large hematoma around the site of the repo sheath. patient was ultimately brought to the or for additional perclose devices and weaned off impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely patient condition since the patient was extremely coagulopathic.
25250 | 1220648-2024-12537 | 19567924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567924&pc=OZD | 06/10/2024 | 08/05/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  | 05/31/2025 |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366620 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Hemolysis | Ventricular Fibrillation | Hematuria | Thrombosis/Thrombus | Hemodynamic instability | [8 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 54-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that while on the support, the patient developed red urine and no pulsatility upon arrival at a new, different facility. the complainant reported the cp was deep into the mitral apparatus, which correlated with a placement signal alarm. an apical thrombus was noted on echocardiography in the left ventricle. medical staff then exchanged the cp with an impella 5.5. the patient survived the event. | [5 empty] | The investigation of hemolysis, thrombus, and positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue most likely improper positioning related due to improper technique. as the necessary information was not provided, the root cause of the thrombus was not determined. the cause of the positioning issue most likely use related due to improper technique as the echo showed the pump was deep and in mitral apparatus-which correlates with psl alarm. the pump was unable to repositioned. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12537: e4 should have been left blank. g1 reporting contact fax number missing. h.6 code 2558, 4868, 2130, 4642, 4624, 4641, 4627, and 4614 were reported incorrectly. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25251 | 1220648-2024-12538 | 19567984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567984&pc=OZD | 06/10/2024 | 08/05/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481093 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Low Oxygen Saturation | Hemodynamic instability | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the pump came out of position, deep with slack in the aorta. out of position alarms occurred. the patient experienced respiratory distress with low oxygen, requiring intervention with bilevel positive airway pressure. medical staff eventually successfully weaned the patient off impella support, and the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issue has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12538: e4 should have been left blank. g1 reporting contact fax number missing. h.6 code 4868, 4625, and 4629 were reported incorrectly. new code was added to health effect - clinical code, health effect - impact code.
25252 | 1220648-2024-12539 | 19567991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567991&pc=OZD | 06/10/2024 | 08/06/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471098 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after setting up and priming pump, no purge noted on the outflow. no alarms present no obvious reason. realized an atraumatic towel clip was crushing white cord to the drape. corrected, immediately got a purge system open alarm followed by purge pressure low alarm. | [5 empty] | The investigation for the reported purge leak has been completed. the pump was returned for investigation. the data log shows high purge pressure from the start of the case, followed by a low purge flow alarm. pump insertion was aborted without completing the case start. the pump was never inserted into the body. the pump was tested for high purge pressure. no physical damage was observed on the returned pump. high purge pressure was reproduced at the test start, but it was resolved after 30 minutes of purging and soaking. the pump then ran as expected. according to the clinical report, pump fluid was not coming out of the outflow, and the doctor found that an atraumatic towel clip was crushing the white cord to the drape. the purge system open alarm appeared after the clip was removed. the doctor removed the pump and replaced it. crushing the white cord can create high purge pressure without leaving any physical mark. the cause of the high purge pressure issue was most likely use related, based on given clinical and returned product investigation. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- d6b, d9 device return date d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
25253 | 1220648-2024-12540 | 19567995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19567995&pc=OZD | 06/10/2024 | 08/13/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441969 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | 67the user facility reported a 67- year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella 5.5 insertion the patient¿s artery was very small and impella wouldn't advance past the medial axillary. unsuccessful attempts were made and after 40 minutes of attempting the 5.5 case was aborted for a cp. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. artery was very small and the impella would not advance past the medial axillary as per the clinical description provided. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12540 submitted. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12540 in accordance with updated procedures. e.4 should have been left blank on the initial submitted manufacturer device report number 1220648-2024-12540 as it is unknown if the initial reporter also sent the report to the fda. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2024-12540 submitted. a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12540 in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
25254 | 1220648-2024-12541 | 19568004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568004&pc=OZD | 06/11/2024 | 07/31/2024 |  | 06/19/2024 |  |  | 06/11/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492110 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 19-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed access site bleeding. medical staff performed a blood transfusion. the patient was successfully weaned off impella support and survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12541: e4 should have been blank. g1 mdr reporting address missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25255 | 1220648-2024-12542 | 19568010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568010&pc=OZD | 06/11/2024 | 08/16/2024 |  | 06/19/2024 |  |  | 06/11/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025467051 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old female patient with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that it was unsuccessful placing impella 5.5 due to difficulty getting inlet of the 5.5 past the subclavian artery. the 5.5 insertion was aborted, and cp was placed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. axillary artery was pretty small per the clinical description provided. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use : section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as information was omitted from the initial manufacturer device report number 1220648-2024-12542. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12542 as it is unknown the initial reporter also sent the report to the fda. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12542 in accordance with updated procedures.
25256 | 1220648-2024-12543 | 19568019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568019&pc=OZD | 06/11/2024 | 07/31/2024 |  | 06/19/2024 |  |  | 06/11/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509622 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Failure to Deliver | Adverse Event Without Identified Device or Use Problem | [5 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 46-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was placed successfully, and support was initiated. peel away sheath removed and upon advancement of repositioning sheath, perforation of the artery was noted. manual pressure was held, and fem stop was placed. repositioning sheath was unable to advance or suture the hub down. hematoma was noted. vascular surgery repaired the artery. the cp was removed and 5.5 placed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for vascular damage - peripheral vessel) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12543: d4 model number. e4 should have been left blank . h.6 code 2338 was reported incorrectly. new code was added to health effect - impact code and medical device problem code.
25257 | 1220648-2024-12544 | 19568023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568023&pc=OZD | 06/11/2024 | 07/22/2024 |  | 06/19/2024 |  |  | 06/11/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489321 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post percutaneous coronary intervention. while on support, the patient experienced bleeding at the access site. the impella was manipulated several times causing the repo sheath to pull back and there was bleeding. the patient developed a hematoma and it was unable to be managed. the pump was removed and a 16fr sheath was placed. blood products were provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since the physician manipulated catheter several times each time causing the repo sheath to pull back causing bleeding. e4 should have been left blank in the initial report.g1 reporting contact fax number missing
25258 | 1220648-2024-12545 | 19568030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568030&pc=OZD | 06/12/2024 | 07/18/2024 |  | 06/19/2024 |  |  | 06/12/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1207023 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility biomedical department reported that the automated impella controller (aic) had a battery failure. the aic was removed from service. there was no patient harm or involvement reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The console was received for investigation. after reviewing the logs, the reported battery failure issue was confirmed. there was evidence of battery failure was tested after arriving in field service and the battery failure issue was able to be reproduced. results of running battest on telnet revealed that vcell2 in battery2 had a voltage that was significantly less than the rest of the cells in the battery conclusion: the root cause of the reported battery failure issue was determined to be caused by internal battery cell imbalance. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12545: e4 should have been left blank.
25259 | 1220648-2024-12546 | 19568049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568049&pc=OZD | 06/12/2024 | 07/19/2024 |  | 06/19/2024 |  |  | 06/12/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025506607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 56-year-old female patient with unknown ethnicity with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient on impella cp support was transferred to another facility and upon arrival the patient was found to be ischemic on the same side as the cp placement. the impella was removed and was replaced with intra-aortic balloon pump. the patient was reported as critical | [5 empty] | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication vascular size or variations there of detailed description of the symptoms experienced by the patient physical examination findings, including pulse assessment and visual examination of the affected limb diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography laboratory test results, including blood tests for markers of inflammation or clotting disorders any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. e4 should have been left blank in the initial report. h.6 code 925 was reported incorrectly on manufacturer device report 1220648-2024-12546. new code was added to component code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25260 | 1220648-2024-12547 | 19568363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568363&pc=OZD | 06/03/2024 | 08/12/2024 |  | 06/19/2024 |  |  | 06/03/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489831 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | From the st-elevation myocardial infarction -door-to-unload study, it was reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, there was volume shortage reported. the impella experienced suctioning. the device was repositioned, and fluid was given. the device was ultimately explanted. | [5 empty] | Ethnicity and race are unknown. the impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device was not returned for investigation. data logs not available to review. insufficient clinical data were not sufficient to determine the root cause. therefore, the cause of the low pump flow issue cannot be determined since no complaint device returned and insufficient data provided.
25261 | 1220648-2024-12548 | 19568470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568470&pc=OZD | 06/10/2024 | 08/09/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  | 10/31/2025 |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420447 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The patient continued to decompensate and expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed a purge system blocked alarm. upon changing purge cassette, new cassette prompted "cassette failure" and unable to advance the purge disc into place and clear the notification. a decision was made to place the previous purge cassette and perform aic to aic transfer. after the aic was exchanged, tissue plasminogen activator was initiated successfully. the patient remains on impella 5.5 support on the date of this report. |  |  |  |  | The investigation of high purge pressure has been completed. a purge cassette was returned for evaluation, no pump was returned for evaluation. no abnormalities noted on visual inspection of the cassette. cassette was connected to aic and "defective purge cassette" alarm was triggered. no data logs were returned for analysis. clinical details noted that a high purge pressure alarms were resolved with tissue plasminogen activator (tpa) and when attempting to resolve high purge pressure with cassette change, new cassette received "cassette defective" alarm and was not used. the root cause of the high purge pressure was most likely due to biomaterial as it was resolved with tpa. b5 updated with additional information d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12548: e4 should have been left blank h.6 code 4644 was reported incorrectly. new code was added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follo impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25262 | 1220648-2024-12549 | 19568493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568493&pc=OZD | 06/04/2024 | 07/29/2024 |  | 06/19/2024 |  |  | 06/04/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486422 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support during pci. during support, the patient experienced oozing from the insertion site. clotting powder was applied and pressure was held. the impella was explanted. the pci was performed successfully and the patient was doing well. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped as per the clinical description. added- h6 component code 925 f6/f8 should have been left blank in the initial report.
25263 | 1220648-2024-12550 | 19568585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568585&pc=OZD | 06/04/2024 | 07/18/2024 |  | 06/19/2024 |  |  | 06/04/2024 | 06/19/2024 |  |  | 06/30/2025 |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359968 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella could not pass into the patient's artery due to patient's vessel calcification. the team aborted the impella 5.5 procedure and implanted an impella cp. there was no patient harm. | [5 empty] | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, could not pass the 5.5 device into the axillary artery due to calcification as per the clinical description. f6/f8-should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25264 | 1220648-2024-12551 | 19568862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19568862&pc=OZD | 06/10/2024 | 08/05/2024 |  | 06/19/2024 |  |  | 06/10/2024 | 06/19/2024 |  |  |  |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025483533 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Pericardial Effusion | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during placement of impella cp after cp was started on auto approximately 4 minutes later the patient lost all pulsatility. tte was performed, and it was noted that the patient had a large pericardial effusion. pericardiocentesis was initially performed which did result in an initial recovery of the patient¿s pulsatility. however, the volume reaccumulated and a decision was made with the cardiac surgery team to perform a sternotomy. patient was intubated, sternotomy performed, bypass initiated with circulatory arrest. the impella was removed, and lv was repaired. iabp placed and patient was sent to the icu open-chested, requiring further repair. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the ventricle perforation has been completed. no product or logs were returned for evaluation. clinical details note that additional guidewire was left in lv for use in tavr procedure after planned pci. additionally, it was noted that patient's lv was diseased/damaged before impella placement occurred. it was noted that large pericardial effusion was noted in patient after loss in pulsatility. the root cause of the ventricle perforation cannot be determined as no product or logs were returned for analysis. h6 component code 4755 changed to 925.
25265 | 1220648-2024-12552 | 19570046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19570046&pc=OZD | 06/07/2024 | 07/30/2024 |  | 06/19/2024 |  |  | 06/11/2024 | 06/19/2024 |  |  | 05/31/2025 |  | 06/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352541 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention (pci) was to be implanted with an impella cp device for mechanical circulatory support during pci. the impella cp was unable to cross the aortic valve due to the patient having stenosis. the impella placement was aborted. the patient is stable. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unbale to pass through aortic valve after multiple attempts and patient is having vessel stenosis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
25266 | 1220648-2024-12499 | 19558028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558028&pc=OZD | 05/28/2024 | 07/23/2024 |  | 06/18/2024 |  |  | 05/29/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1434576 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the purge cassette and disc was installed as always, however the automated impella controller (aic) did not recognize the purge system. the staff used a backup purge cassette with the same result. therefore, they changed to another aic and the system worked well after the change of the aic. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The console was received for investigation. after testing in the lab, the reported issue could not be reproduced. console was able to prime the purge cassette and detect purge fluid as normal. there were no visible issues observed with the purge disc retainer and flag. the root cause of this issue was most likely a temporary buildup of ingress on the purge motor shaft. e4 should have been left blank in the initial report. g1 manufacturing site name and address
25267 | 1220648-2024-12500 | 19558234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558234&pc=OZD | 06/03/2024 | 01/30/2025 |  | 06/18/2024 |  |  | 06/03/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that on arrival at the hospital to transfer loan monitors back into the transfer case, the representative switched on the hospital¿s returned automated impella controllers (aic) as one of the consoles had an issue saying system self-check failed. the aic was switching on and off. there was no patient involvement. | [5 empty] | Catalog number, serial number, lot number and unique identifier (udi) # are unknown. device manufacture date is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) power issues has been completed. console imc logs confirm that the controller failure (pud comm loss) - alarm was issued. the aic was returned for evaluation. the failure was reproduced in the lab. a test pc was used to perform a firm ware check on the console revealing the pud not being recognized. the failure was resolved when a known good pud was swapped into the console. the root cause for the system self check failure was a defective pud. added d9 device return date corrections to the initial mfr report: d4-model, catalog, serial, and udi numbers have been added. f6/f8 should have been left blank. g1 mdr reporting contact email. h4-added device mfr date.
25268 | 1220648-2024-12501 | 19558267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558267&pc=OZD | 05/29/2024 | 07/09/2024 |  | 06/18/2024 |  |  | 05/29/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1815943 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic) displayed an alarm stating "impella device defective". the aic was exchanged. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This supplemental report is generated to retract the initial mdr. this was a duplicate file and was reported under our case (b)(4) and mfr report number 1220648-2024-11988.
25269 | 1220648-2024-12502 | 19558403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558403&pc=OZD | 06/01/2024 | 08/08/2024 |  | 06/18/2024 |  |  | 06/01/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488330 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during implantation of an impella cp device to a patient experiencing acute myocardial infarction/cardiogenic shock, the automated impella controller displayed and sounded a "purge disc not detected" alarm. the physician noticed a leak at the purge disk membrane. however, before the purge cassette could be replaced, the impella cp pump stopped. additionally, it was noticed that the pump had pulled into the aorta. the pump was ultimately, successfully explanted, and the patient's outcome as a result of the event was noted as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop and the positioning issues has been completed. an impella cp was successfully delivered through the right femoral artery of the patient. during the case, it was noted that the pump was pulled into the aorta. this was part of the explant. about 5 hours into support, the "purge line click-on not detected" alarm sounded. a leak at the purge disc membrane was discovered. the purge cassette was not replaced because there were none available. the pump stopped within the next 30 minutes and could not be restarted. the root cause of the pump stop was most likely a purge leak at the purge disc membrane as reported in the clinical details. it was later reported that the blue holder was broken off of the aic, which most likely occurred while repositioning the holder. the root cause of the purge leak was most likely a damaged purge disc as reported in the clinical details.
25270 | 1220648-2024-12503 | 19558500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558500&pc=OZD | 05/25/2024 | 07/30/2024 |  | 06/18/2024 |  |  | 05/27/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493670 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and during support, the patient experienced limb ischemia. there was limb ischemia on the right side and compartment syndrome in lower leg. surgical splitting was consequently performed. the patient's outcome from the event was noted as stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). .¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
25271 | 1220648-2024-12504 | 19558976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558976&pc=OZD | 05/27/2024 | 08/16/2024 |  | 06/18/2024 |  |  | 05/28/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450533 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support for 73-year-old male patient, the distal part of the sterile sleeve (close to the red plug) was broken. the device was not removed as a result. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage (sterile sleeve damage) has been completed. the pump was not returned for investigation. according to the clinical report, the distal part of the sterile sleeve (close to the red plug) was found broken. the cause of the product damage (sterile sleeve damage) issue was not determined since the product was not returned and sufficient clinical information was not provided. b.3 revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12504. d.4 unique identifier (udi) # was revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12504. e.1 last name should have been left blank on the initial manufacturer device report number 1220648-2024-12504 as it is unknown what the physician's last name is. e.1 address line 1 and 2 were revised as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12504. e.1 international phone number was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12504 and should have been left blank. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12504 as it is unknown the initial reporter also sent the report to the fda. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12504.
25272 | 1220648-2024-12505 | 19558998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19558998&pc=OZD | 05/29/2024 | 08/05/2024 |  | 06/18/2024 |  |  | 05/29/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480796 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention (pci). during pci procedure, the rotablator was stuck in the circumflex coronary artery and the team was unable to retract it. an emergency coronary artery bypass grafting (cabg) procedure was performed. the impella cp was retracted during the cabg due to risk of damaging the impella with the rotablator. afterwards, the team was unable to replace the impella over the aortic valve. therefore, pump was explanted and a new impella was placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the positioning issue has been completed. product and data logs were not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
25273 | 1220648-2024-12498 | 19559026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19559026&pc=OZD | 05/29/2024 | 07/15/2024 |  | 06/18/2024 |  |  | 05/29/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1384051 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) did not detect the purge cassette. several purge cassettes were tested without success. the console had to be exchanged and the issue resolved. there was no patient involvement. | [5 empty] | The aic device was received for evaluation. upon completion of the investigation a supplemental will be filed. | The investigation for the aic-purge disc/flag issues has been completed. the console was returned for evaluation. purge disc flag was found to be broken (missing at the blue purge disc retainer). clinical staff reported that the aic was unable to detect the purge cassette but after reviewing the logs it was determined that they likely meant to report that the purge disc could not be detected. data imc logs showed no signs of the aic having trouble detecting the purge cassette. however, the data imc logs did have evidence of the console having trouble detecting the purge disc (numerous purge disc not detected alarms). the root cause of the observed issue was a broken purge disc flag. added- h6 impact code 4629 g1-corrected mfr site name and address.
25274 | 1220648-2024-12506 | 19559890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19559890&pc=OZD | 05/30/2024 | 08/06/2024 |  | 06/18/2024 |  |  | 05/30/2024 | 06/18/2024 |  |  | 11/30/2025 |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427349 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient experienced high purge pressure (< 1050 mmhg) during support. medical staff noted no kinks on the purge line and added tissue plasminogen activator to resolve the issue. no patient harm has been reported. | [5 empty] | The device has not been received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. the team on site reported to the field rep that the purge pressures had surpassed 1050 mmhg, so the field rep recommended changing the device and administering tpa. replacing the purge cassette and tpa are said to have resolved the failure mode. neither product nor data logs were returned for investigation. the root cause of the high purge pressure was most likely biomaterial since tpa infusions resolved the complication.
25275 | 1220648-2024-12507 | 19560116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19560116&pc=OZD | 05/31/2024 | 08/16/2024 |  | 06/18/2024 |  |  | 05/31/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025466191 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 54-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that during support, the automated impella controller (aic) displayed high purge pressure (>1075 mmhg) alarm. medical staff replaced the purge cassette and added tissue plasminogen activator to the purge. medical staff noted no kinks on the purge line, and no patient harm has been reported. the patient remains on support. | [5 empty] | The impella device has not been returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. no product or data logs were returned for evaluation. clinical details noted heparin was used in intravenous drip throughout support. the purge cassette was changed and did not resolve issue. tissue plasminogen activator therapy was given. it was not noted if high purge pressure alarms resolved or if the pump was removed. the root cause of the high purge pressure could not be determined as no product or data logs were returned. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ e.1 revised facility name and address line 1 as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12507. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12507 as it is unknown the initial reporter also sent the report to the fda. g.1 revised multiple fields as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12507. h.10 additional manufacturer narrative instructions for use were revised from the initial submission manufacturer device report number 1220648-2024-12507 in accordance with updated procedures.
25276 | 1220648-2024-12508 | 19560192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19560192&pc=OZD | 06/03/2024 | 07/30/2024 |  | 06/18/2024 |  |  | 06/03/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482122 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. the heparin was off due to initial small hematoma at the cutdown site which is resolving. it was further reported that the sterile sleeve became detached from the tuohy borst and was torn. the physician did not want to replace the impella 5.5, so the sleeve was covered with tegaderm, and the patient was monitored for infection. the patient was stable. the patient went for transplant and pump was discarded. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the product damage and the access site bleed has been completed. the impella pump was not returned. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. the cause of the product damage (sterile sleeve damage) issue was not determined since product was not returned.
25277 | 1220648-2024-12509 | 19560204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19560204&pc=OZD | 05/31/2024 | 08/09/2024 |  | 06/18/2024 |  |  | 05/31/2024 | 06/18/2024 |  |  | 10/31/2025 |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024416322 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  |  | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support for 72-year-old female patient, the automated impella controller (aic) displayed purge pressure high ((pp>1044mmhg/ pf < 1ml) error message. the purge cassette was replaced but issue did not resolve. the patient acts were between 160/18, and the dextrose concentration above 5%. there was no patient harm reported. |  |  |  |  | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information the following have been reported incorrectly or missing from manufacturer device report. | The device has not been returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
25278 | 1220648-2024-12510 | 19560409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19560409&pc=OZD | 06/03/2024 | 08/09/2024 |  | 06/18/2024 |  |  | 06/05/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025470451 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The japan user facility reported a patient of an unknown age and in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp came out of the aortic valve and was removed. there was no patient harm. | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12510. e4 should have been left blank.
25279 | 1220648-2024-12511 | 19561269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19561269&pc=OZD | 06/05/2024 | 07/12/2024 |  | 06/18/2024 |  |  | 06/05/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025493946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported the impella cp was attempted via the right axillary approach. the cp was able to be placed in the ventricle, however once the wire was removed, it migrated to the descending aorta. a decision was made to monitor the patient off impella support. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump could not be delivered due to resistance met by patient anatomy. e4 should have been blank in initial report.
25280 | 1220648-2024-12513 | 19561423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19561423&pc=OZD | 06/07/2024 | 07/31/2024 |  | 06/18/2024 |  |  | 06/07/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482967 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed a hematoma. systemic heparin was withheld to control the hematoma. in addition, the staff attempted to reposition the impella 5.5 but was unsuccessful. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the positioning issue has been completed. data logs were reviewed which confirmed pump was in aortic area when pump started then resolved. product was not returned for review. the root cause of positioning issue was unable to be determined due to lack of clinical details. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
25281 | 1220648-2024-12512 | 19561457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19561457&pc=OZD | 06/06/2024 | 07/24/2024 |  | 06/18/2024 |  |  | 06/06/2024 | 06/18/2024 |  |  | 07/31/2025 |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442795 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported an 82-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported suction alarms occurred while the cp was on auto. medical staff unsuccessfully tried to resolve the issue with repositioning under fluoroscopy. medical staff then elected to explant the cp for a possible thrombus and implanted another cp. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for low pump flow has been completed. the product was not returned for review. data review: data logs confirmed the failure mode and clinical narrative. logs showed after pump started, mc spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12512: e4 should have been left blank. h.6 code 4641, 4628, and 1158 were reported incorrectly.
25282 | 1220648-2024-12515 | 19561551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19561551&pc=OZD | 06/08/2024 | 08/30/2024 |  | 06/18/2024 |  |  | 06/08/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482378 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a red alarm and then the pump stopped. the team attempted to restart the pump several times without success. the aic was not exchanged. the decision was made to explant the pump as the patient was stable. there was no reported harm. | [5 empty] | The investigation for the reported pump stop has been completed. on the third day of impella cp support, the pump stopped with red alarms on the automated impella controller. an "impella defective" alarm was noted. they could not restart the pump after multiple attempts. the pump was explanted as a result. the root cause of the pump stop was not determined because the product and logs were not returned. d.4 revised unique identifier (udi) # as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12515. e.1 revised first name, last name and address line 1 as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12515. e.4 should have been left blank in initial manufacturer device report number 1220648-2024-12515 as it is unknown the initial reporter also sent the report to the fda. g.1 revised manufacturing site address line 1 and 2, city and zip code as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12515. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25283 | 1220648-2024-12516 | 19561552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19561552&pc=OZD | 06/08/2024 | 08/01/2024 |  | 06/18/2024 |  |  | 06/08/2024 | 06/18/2024 |  |  | 09/30/2025 |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444886 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the impella would not pass through the sheath. it was then removed and angiogram was performed. there was a kink noticed on angiogram due to patient's torturous vessel. a new impella cp was implanted successfully. there was no patient harm. | [5 empty] | The investigation of difficulty inserting/removing pump through introducer and delivery issues has been completed. clinical states that the introducer was seen to have been kinked due to patient's tortuosity. the introducer was not returned for investigation. the pump was found to have a kinked cannula, likely from the issues being inserted through the kinked introducer. the root cause of the difficulty inserting pump through introducer issue was likely an introducer sheath kink due to the patient's torturous vessel. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ section: warnings & cautions: warnings section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
25284 | 1220648-2024-12514 | 19561554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19561554&pc=OZD | 06/06/2024 | 07/31/2024 |  | 06/18/2024 |  |  | 06/06/2024 | 06/18/2024 |  |  |  |  | 06/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480345 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 66-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed access site bleeding. medical staff administered multiple blood products, and the patient remained stable. the cp was eventually explanted. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. g4-added missing pma number h6 component code 4755 changed to 925.
25285 | 1220648-2024-12461 | 19546770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19546770&pc=OZD | 01/18/2024 | 07/11/2024 |  | 06/17/2024 |  |  | 05/23/2024 | 06/17/2024 |  |  | 06/30/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359705 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Stenosis | [10 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute ischemic stroke with a "possible" relationship to the impella device used: ¿pt with new onset of left facial droop and rue/rle weakness. stroke team called. head ct revealed acute on chronic right subdural hematoma and severe stenosis involving the mid left common carotid artery. noted to have experienced an acute ischemic stroke with possibility of the cause being the stenosis in the mid left common carotid artery. vascular surgery and neurosurgery consults obtained with no acute intervention recommended. symptoms resolved 19jan2024.¿ the patient recovered with no sequelae. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25286 | 1220648-2024-12462 | 19546812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19546812&pc=OZD | 04/05/2024 | 07/22/2024 |  | 06/17/2024 |  |  | 05/26/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448034 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | Abiomed received a notification via abiomed's long-term outcome and quality indicator (loqi) impella registry reporting the patient endured fever and leukocytosis with a "possible" relation to the impella cp. a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post percutaneous coronary intervention. the patient was also supported with venoarterial extracorporeal membrane oxygenation. ¿4/5: febrile overnight with uptrending white blood cell count and blood cultures growing strep (not group a, group b, pneumo, or enterococcus). started on vanc/cefepime. uptrending levophed requirements, septic dose vasopressin added. responsive to 500cc ivf x2. last cardiac index up to 5.0 with svo2 70. 4/6: de-escalated to rocephin per infectious disease (id) doctor. transthoracic echocardiogram (tte) limited negative for infective endocarditis (ie) weened from pressors. off dobutamine. -bacteremia: febrile, hypotensive requiring levo and vaso on 4/4 blood cultures 4/4: group g strep (sensitivities pending) started vanc/zosyn, de-escalated to rocephin 4/5 per id repeat tte 4/5 limited negative for ie swan ganz removed 4/5 plan: repeat cx 48h on 4/6 @9pm continue rocephin follow-up sensitivities id following- pending outpatient parenteral antimicrobial therapy (opat) and possible pic if iv antibiotics needed long term pull venous sheath and arterial lines started vanc/zosyn and septic dose vasopressin added to levophed. line infection favored, swan ganz removed. responded well to intravenous fluids, weened from pressors on 4/5. blood cultures 4/4 growing group g strep (sensitivities pending). de-escalated to rocephin 4/5 per id. repeat cultures to be drawn 4/6 @ 9pm. seen by id, recommended bacteremia was likely related to multiple lines, ecmo and impella and venous sheath have now all been removed. repeat blood cultures have returned negative. recommend total 7 days of rocephin, stop date 04/13/2024¿. the patient outcome is unknown at this time. | [5 empty] | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issue could not be determined. as noted in the impella instructions for use: cp with smartassist: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12462 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12462 was submitted. h5- revised was inadvertently left off the previously submitted manufacture device report 1220648-2024-12462. h.6 code 4755 was reported incorrectly on manufacturer device report 1220648-2024-12462. a new code was added to component code. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25287 | 1220648-2024-12463 | 19546839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19546839&pc=OZD | 01/09/2024 | 07/18/2024 |  | 06/17/2024 |  |  | 05/27/2024 | 06/17/2024 |  |  | 10/31/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412206 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured pericardial effusion with tamponade with a "probable" relationship to the impella device used: ¿course complicated by pericardial effusion w tamponade s/p pericardiocentesis. - pericardial tamponade 2.5l removed with pericardial centesis (b)(6) 2024. unclear if iatrogenic from mcs. pericardial drain was placed. plan: - resolved - pericardial drain removed (b)(6). pericardial effusion with echo evidence of tamponade, drained in cath lab for 2.5 serosang fluid. successful pericardiocentesis with 1 attempt using ultrasound and flouroscopic guidance. approximately 2.5l of sanguineous pericardial fluid was drained; contrast injection via the pericardial drain confirmed placement within the pericardium. the pericardial drain was sutured in at the conclusion of the procedure. 3. given large volume of blood aspirated from the pericardial space, there was concern of myocardial injury during placement of previous mcs devices¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. f6/f8 should have been left blank in the initial report. h5-should have been yes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25288 | 1220648-2024-12464 | 19546878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19546878&pc=OZD | 06/02/2023 | 07/12/2024 |  | 06/17/2024 |  |  | 05/29/2024 | 06/17/2024 |  |  | 12/31/2024 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292423 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | Abiomed received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding (hematoma) with a "possible" relationship to the impella device used: a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. ¿(b)(6): ct: impella cardiac device is in place with a large right axillary hematoma extending inferiorly from the access site. (b)(6): hematoma to right chest wall is soft with no overlying skin changes and no active bleeding from skin tunnel (b)(6) 2020: right chest hematoma stable with no need for surgical intervention (b)(6) 2020: ct: decrease in size of the right chest wall hematoma. (b)(6) 2001: ct scan concerning for hematoma (b)(6): left axillary pocket hematoma evacuation, washout subxiphoid pericardial drainage¿. the patient recovered with no sequelae." | [5 empty] | The investigation for access site bleeding/hematoma has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 added health effect - impact code 4624 f6/f8-should have been left blank in the initial report. h6 component code corrected to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25289 | 1220648-2024-12465 | 19547416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19547416&pc=PYX | 05/17/2024 | 06/28/2024 |  | 06/17/2024 |  |  | 05/17/2024 | 06/17/2024 |  |  | 01/31/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024325058 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support and developed hemolysis. in the unit, it was noted the patient¿s urine was red, and labs were hemolyzing. approximately, two days later care was withdrawn. the patient was noted to have subsequently expired. upon abiomed's medical safety officer review, there is not enough information to associate the use of device with the patient's outcome. | [5 empty] | H.6 code was inadvertently omitted from the previously submitted report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the product was not returned. the data logs were reviewed and there was no motor current spike, positioning issues or a trend in placement signal. there were suction alarms throughout the whole case. due to lack of product returned and insufficient clinical information, the root cause of the hemolysis issue could not be determined.
25290 | 1220648-2024-12466 | 19547528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19547528&pc=PYX | 05/17/2024 | 07/24/2024 |  | 06/17/2024 |  |  | 05/18/2024 | 06/17/2024 |  |  | 06/30/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024424269 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 73-year-old male patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that following the implant, the automated impella controller indicated high purge pressure. tissue plasminogen activator was added to the purge and the motor current was monitored. it was reported that the patient expired during the night while one support. the post market surveillance team made the decision at the discretion of abiomed's medical safety officer to file as death. the death event occurred and no alternative cause for the death event was identified as the likely cause for the patient outcome. | [5 empty] | The investigation for high purge pressure has been completed. no product was returned. the data logs confirm 'purge pressure high' alarms. the data logs show that about 30 minutes after implant, there was a motor current spike followed by a gradual rise in purge pressure. there was high purge pressure for about 13 minutes and tissue plasminogen activator (tpa) was added to the purge after which the high purge pressure resolved and started decreasing towards normal values. the root cause of the high purge pressure was most likely biomaterial ingestion. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12466: f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smartassist system: section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
25291 | 1220648-2024-12469 | 19547945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19547945&pc=OZD | 05/20/2024 | 07/17/2024 |  | 06/17/2024 |  |  | 05/20/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The complainant reported a 53-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that while on support, the impella cp¿s tuohy-borst valve broke while the medical staff was placing the repositioning sheath. the sterile sleeve was advanced and locked to the sheath and the lock was twisting in place instead of advancing forward to hold the catheter. the medical staff secured the impella cp using tegaderm and transferred the patient to the intensive care unit. the patient eventually expired. upon review by abiomed's medical safety officer review, there is not enough information to associate the use of the device with the outcome. | [5 empty] | The investigation for product damage (sterile sleeve damage) has been completed. the root cause of the product damage (sterile sleeve damage) was not determined because no product was returned. f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25292 | 1220648-2024-12470 | 19548078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548078&pc=OZD | 06/04/2024 | 08/20/2024 |  | 06/17/2024 |  |  | 06/04/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314950 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Component Incompatible | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) indicates purge disk not recognized and it was noted that the purge disk holder was loose. therefore, the console was exchanged, and the issue resolved. the patient is stable. |  |  |  |  | The investigation for the console (aic) purge disc/flags issues has been completed. logs are irrelevant to this complaint as the clinical staff reported that the purge disc retainer was loose. however the logs were reviewed, purge disk not detected and purge pressure low alarm was confirmed. the alarms were due to the cause and effect of the broken purge retainer. the aic was returned for evaluation. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. purge disk not detected alarm observed by the reporter was likely due to the broken purge disc retainer. | The impella device was received from the customer on (b)(6) 2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
25293 | 1220648-2024-12471 | 19548094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548094&pc=OZD | 06/05/2024 | 07/11/2024 |  | 06/17/2024 |  |  | 06/05/2024 | 06/17/2024 |  |  | 09/30/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410031 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The complainant reported a 50-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported medical staff had to explant the cp due to the patient developing distal limb ischemia. the patient survived the event. | [5 empty] | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. f6/f8-should have been left blank in the initial report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25294 | 1220648-2024-12472 | 19548130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548130&pc=OZD | 05/23/2024 | 07/31/2024 |  | 06/17/2024 |  |  | 05/23/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481519 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a male patient of unknown age, race, and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after five days of support, the pump stopped. several attempts were made to restart the pump. the patient¿s pulsatile flow eight milliliters per hour and the pulse pressure was 505 mmhg. the mechanical circulatory trend quickly increased and then stopped. the decision was made to withdraw care and the patient expired. the post market surveillance team made the decision at the discretion of abiomed's medical safety officer to file as death. the death event occurred and no alternative cause for the death event was identified as the likely cause for the patient outcome. | [5 empty] | The investigation for pump stop has been completed. impella cp pump was returned. visually inspected and no biomaterial observed in the pump. large purge gap was observed during inspection due to bearing wear. no other abnormalities were found. pump stop was observed on the 6th day of support. the cause of the pump stop was bearing wear within design life. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12472: f6/f8-should have been left blank. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25295 | 1220648-2024-12473 | 19548204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548204&pc=OZD | 05/23/2024 | 07/16/2024 |  | 06/17/2024 |  |  | 05/23/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478495 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 57-year-old male patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported the patient suffered a stroke during impella 5.5 support. the patient's left side was immobile as a result of the stroke. support with the implanted pump continued following the diagnosis. the decision was made to withdraw care and the patient expired. per abiomed's medical safety officer review, there is not enough information to associate the use of the device with the outcome. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. f6/f8 should have been blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25296 | 1220648-2024-12474 | 19548456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548456&pc=OZD | 05/25/2024 | 07/30/2024 |  | 06/17/2024 |  |  | 05/25/2024 | 06/17/2024 |  |  | 08/31/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386793 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a male patient with acute myocardial infarction/cardiogenic shock had an attempted implant with an impella cp for mechanical circulatory support. it was reported that the insertion was aborted due to the patient¿s anatomy. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for delivery issues has been completed. the root cause of the delivery issue was determined to be patient condition because the implant was aborted due to small vessels. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12474 and should not have been. h.6 code 2199 was entered incorrectly on the initial manufacturer device report number 1220648-2024-12474. a new code has been added.
25297 | 1220648-2024-12475 | 19548486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548486&pc=OZD | 06/05/2024 | 07/24/2024 |  | 06/17/2024 |  |  | 06/05/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the patient had access site bleeding. after the implant, the patient had copious bleeding from the impella site. manual pressure was held and forward tension sutures were replaced to maintain a better angle. the bleeding subsided. while on support, a hematoma developed on the patient left groin (impella side) and blood product and platelets were given. the patient had bleeding from the ecmo (extracorporeal membrane oxygenation) site and the impella was removed. upon removal, a large clot was noted in the impella cannula. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the root cause of the access site bleeding was most use issue likely due to use issue since hemostasis was achieved by replacing forward sutures. the root cause of the thrombus could not be determined without additional clinical details.
25298 | 1220648-2024-12476 | 19548877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548877&pc=OZD | 06/06/2024 | 07/30/2024 |  | 06/17/2024 |  |  | 06/06/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466376 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella cp was inserted and started but it was unable to get flows above 1l/min. placement was confirmed under fluoro. the cp was pulled out and a new cp was placed successfully. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella was not returned for evaluation. data logw were reviewed finding that a suction alarm occurred. no positioning issue occurred. there were no issues found to confirm the cause of reported issue. the cause of the low pump flow issue cannot be determined since the complaint device not retuned for investigation. d4-corrected model number.
25299 | 1220648-2024-12477 | 19548884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19548884&pc=OZD | 05/31/2024 | 07/24/2024 |  | 06/17/2024 |  |  | 05/31/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506771 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 37-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support upon arrival at intensive care unit. there was oozing present at the groin site. light pressure femostop was applied. oozing continued so the femostop was replaced with an added stitch around the impella site. the bleeding was resolved. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most use issue likely due to use issue since hemostasis was achieved by adding additional stitch. added-h6 component code 925 f6/f8 should have been left blank in the initial report.
25300 | 1220648-2024-12480 | 19553709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19553709&pc=OZD | 06/06/2024 | 07/30/2024 |  | 06/17/2024 |  |  | 06/06/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484133 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 70-year-old male patient with pre-operation placement had an unsuccessful implant with an impella 5.5 for mechanical circulatory support. the complainant reported due to tortuous vessel anatomy, the implant failed. when medical staff would get close to dropping the 5.5 in the left ventricle, the wire would buckle and lose positioning. medical staff considered placing an intra-aortic balloon pump and attempting another 5.5 implant, though via the left axillary artery instead. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, due to torturous vessel anatomy pump insertion was unsuccessful as per the clinical description. added- h6 impact code 2199.
25301 | 1220648-2024-12482 | 19553758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19553758&pc=OZD | 06/06/2024 | 08/06/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  | 08/31/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024387485 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 85-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the patient on impella cp support had lower limb ischemia. a bypass was done using a reverse sheath. it was further reported that at the time of surgical removal of the cp, the pulsation was weak, so a ct (computed tomography) scan was performed. a thrombus was found in the iliac artery, which was retrieved using a fogarty catheter. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. d4-corrected model number.
25302 | 1220648-2024-12478 | 19554078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554078&pc=OZD | 06/06/2024 | 07/22/2024 |  | 06/17/2024 |  |  | 06/06/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479606 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, the impella 5.5 had a change in the left ventricle waveforms followed by a high purge pressure alarm. the team will convert the purge solution to 2mg/50ml of tissue plasminogen activator in attempt to resolve the issue. the impella 5.5 was explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for high purge pressure has been completed. the impella device was not returned for evaluation. the data logs confirm numerous continuous high purge pressure and purge system blocked alarms. the data logs show that there was a motor current spike the morning of 6/624 after which there was a gradual rise in purge pressure. tissue plasminogen activator (tpa) was added to the purge but the high purge pressure did not resolve for the rest of the case. the root cause of the high purge pressure was most likely due to biomaterial. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
25303 | 1220648-2024-12479 | 19554220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554220&pc=OZD | 06/06/2024 | 07/24/2024 |  | 06/17/2024 |  |  | 06/06/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025458661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an 85-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was sitting very shallow and while advancing pump, the cannula kinked and prolapsed below outflow in ascending aorta with acute reduction of flows to around 1 l/min. attempted repositioning was unsuccessful and there was no improvement in flows and restriction of collapsed cannula remained. the cp was explanted and replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely was patient condition due to patient anatomy as the pump was shallow and advanced and led to cannula kink below outflow. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12479: e4 should have been left blank.
25304 | 1220648-2024-12481 | 19554241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554241&pc=OZD | 06/06/2024 | 07/30/2024 |  | 06/17/2024 |  |  | 06/06/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461413 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing coronary artery bypass and aortic valve replacement was implanted with an impella 5.5 device for mechanical circulatory support. during the implant, the subclavian vein was nicked and began to bleed into the chest cavity. three units of blood were provided to the patient and sutures were applied. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves .¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage has been completed. no product was returned. clinical details states that a vein was nicked during tunneling and was bleeding into the access location and chest cavity. the root cause of the vascular damage - peripheral vessel was not determined due to lack of sufficient clinical information. added- d6b.
25305 | 1220648-2024-12483 | 19554731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554731&pc=OZD | 06/07/2024 | 08/14/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025466086 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 76-year-old male patient with post-cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported high purge pressure occurred. after the high pressure occurred, fluid was found to be leaking from the air filter. medical staff chose to monitor the situation while continuing support. medical staff increased the amount of heparin in the purge and removed air from the purge set several times. however, over time, the purge flow rate decreased, and the purge pressure increased again. the pump also stopped during support. the 5.5 was eventually replaced, and the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak, high purge pressure, and pump stop has been completed. the logs and product were returned and evaluated. the logs showed spikes in purge pressure 15 hours into support. then, near the end of support, there was an increase then spike in motor current followed by a pump stop. biomaterial was found in the purge gap of the returned pump in addition to blood in the purge line. the sidearm filter was sealed on one side with bone wax and taped shut in the field after a leak was discovered. while running the pump in the lab, a crack down the length of the sidearm luer that was leaking was found. the purge leak from the sidearm luer prevented the purge pressure from climbing above 200 mmhg. the pump stop was reproduced in the lab. the root cause of the pump stop was the damaged sidearm filter, which led to blood ingress in the motor. added- d9 device return date.
25306 | 1220648-2024-12484 | 19554735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554735&pc=PYX | 06/07/2024 | 07/31/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  | 10/31/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024423993 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted in for pulmonary embolism was implanted with an impella rp device for mechanical circulatory support and experienced access site bleeding and hemolysis during support. the impella rp pump was successfully placed. however, day two of support, oozing from the impella access site was observed, and it was noted the patient had marked signs of hemolysis in the urine. the groin site was cleaned, re-sutured to stop the bleeding, and pressure applied. additionally, the patient was given (b)(4) units of red blood cells. furthermore, in an effort to stop the hemolysis, the p-level was decreased; however, as thought to be the potential cause of the hemolysis, the decision was made to remove the impella rp pump. it was noted the patient tolerated the explant and was stated to be in stable condition. placement of a swan to further monitor hemodynamic pressures was being considered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia. | The investigation for access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. data logs reviewed: no motor current trend observed. no placement signal trend observed. flow was below required range thorough out the case. no clear indications confirm the reason for reported issue. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. the cause of the access site bleeding was not determined since neither the product nor sufficient clinical information was provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12484. d4 model number. e4 should have been left blank.
25307 | 1220648-2024-12486 | 19554738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554738&pc=PYX | 06/07/2024 | 07/23/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1261094 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the automated impella controller (aic) malfunctioned on day two of the support which prompted the aic to be replaced. the purge disc was not detected. there was no harm or injury to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller failure to read input signal has been completed. the console was not returned and data logs were not provided. the cause of the issue was not determined since product was not returned. d6a and d6b should have been blank. the device is not implanted.
25308 | 1220648-2024-12485 | 19554739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554739&pc=PYX | 06/07/2024 | 07/25/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  | 03/31/2025 |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024394949 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient who underwent cardiac surgery was implanted with an impella rp device for mechanical circulatory support for bipella support to allow for extracorporeal membrane oxygenation (ecmo) decannulation. while the patient was on impella rp support, there were low pump flows. the rp device was explanted. the patient continued with ecmo support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was received for investigation. the data logs confirmed the failure mode and clinical narrative. logs showed after pump started and ran for 8.5 hours several suction alarms were observed. upward shift in motor current trend and low flows were seen. product was returned in damaged condition (no pump plug). visual inspection found biomaterial inside the pump housing. no other issues were found on the rest of the pump. as per given clinical information, patient was maintained with low acts (sub-therapeutic) during insertion of rp flex. the root cause of the low flow was biomaterial ingestion. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12485: e4 should have been left blank. g1 reporting contact fax number missing.
25309 | 1220648-2024-12487 | 19554741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554741&pc=OZD | 06/07/2024 | 07/23/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450133 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed to support the patient with a history of aaa (abdominal aortic aneurysm) and other cardiac comorbidities. the pump was placed femorally and the placement caused limb ischemia. the patient had a known history of pad/pvd (peripheral vascular disease). there was no flow to the foot. the team had the limb intervened upon- angioplasty and thrombectomy performed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the patient's limb ischemia has been completed. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition; any concomitant medication; vascular size or variations thereof; detailed description of the symptoms experienced by the patient; physical examination findings, including pulse assessment and visual examination of the affected limb diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography laboratory test results, including blood tests for markers of inflammation or clotting disorders any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. d4. corrected model number h6 patient impact code was missing from initial mdr
25310 | 1220648-2024-12488 | 19554753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554753&pc=OZD | 06/07/2024 | 07/23/2024 |  | 06/17/2024 |  |  | 06/07/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025484347 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed limb ischemia while on support. there was no distal flow detected. the impella cp was surgically removed and thrombectomy performed of lower left extremity. the patient was successfully transitioned the impella 5.5 and reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 health effect - impact code added 4624 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12488. e4 should have been left blank.
25311 | 1220648-2024-12489 | 19554754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554754&pc=PYX | 06/08/2024 | 07/22/2024 |  | 06/17/2024 |  |  | 06/08/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025478656 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male patient with enlarged right atrium and right ventricle was implanted with the impella rp device for mechanical circulatory support. it was reported that overnight the pump had lower flows due to possible ingested clot and there were signs of hemolysis. the hemolysis was observed with hematuria and the pump was explanted. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. clinical details stated that flow rates decreased on from 3.2 to 1.8 with possible clot ingestion. the data logs reveal low pump flow for matching p-levels. there was a motor current spike and a high purge pressure spike followed by low flow which occurred prior to when hemolysis was reported. the product was returned for evaluation. upon visual inspection, biomaterial was found on and beneath the impeller. a hemolysis test was carried out and the mih values was 10.48. the root cause of the hemolysis issue was determined to be ingested biomaterial. added- d9 device return date.
25312 | 1220648-2024-12490 | 19554763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554763&pc=OZD | 06/08/2024 | 07/31/2024 |  | 06/17/2024 |  |  | 06/08/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025483558 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported that on day 4 of support the patient develop an access site hematoma that was treated by sandbag and a hold was placed on the hearing. the patient remains on support and is reported to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12490: e4 should have been left blank.
25313 | 1220648-2024-12491 | 19554775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554775&pc=OZD | 06/08/2024 | 03/31/2025 |  | 06/17/2024 |  |  | 06/08/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692221 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | Component Incompatible | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support for 64-year-old male patient, the automated impella controller (aic) displayed 'disc not detected' alarm following purge fluid change. a new disc was inserted, and the issue seemed to be resolved but then 'purge system open' alarm occurred. the purge system was checked, and no leakage was noted. the system was deaired and the alarms once again resolved. however, it was noted that fluid now came out the purge cassette area and something black was broken. the console was not exchanged as a result. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the data logs from the complaint date revealed that the console issued purge disc not detected. during the inspection, a broken purge disc retainer was confirmed. the purge retainer was completely broken off. the root cause of this issue was a broken purge disc retainer. d.9 date device returned/information was added. d.2 revised common device name as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-12491. d.4 revised model number and catalog number as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12491. d.6a should have been left blank on the initial manufacturer device report number 1220648-2024-12491. e.1 revised last name, facility name and city since the initial manufacturer device report number 1220648-2024-12491 was submitted in accordance with updated procedures. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12491 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-12491 was submitted in accordance with updated procedures. g1 reporting contact address was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12491. g.1 revised manufacturing site name and address section as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12491. h.8 revised selection as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12491.
25314 | 1220648-2024-12492 | 19554776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554776&pc=OZD | 06/09/2024 | 08/01/2024 |  | 06/17/2024 |  |  | 06/09/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468421 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post cardiac surgery a few days prior. shortly after placing the impella cp, purge flows decreased and purge pressure increased. the cassette changed but ultimately the pump stopped. the impella cp was explanted and replaced with a new impella cp. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Data log shows that purge pressure begins to rise with purge flow decreasing immediately. 11 minutes into support there are high pressure puge (hpp) and purge system blocked alarms. motor current is trending upwards. 15 minutes after the first hpp alarm there is alarm 13 hmc pump stop. the pump is unable to be restarted. the hpp is eventually resolved, though remains at the higher end. hpp was reproduced in the lab. hmc pump stop was also reproduced, and the pump was unable to be run. upon cutting a window in the catheter blood was found in the purge line. the root cause of the pump stop was blood ingress. the cause of the blood ingress was unknown. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12492: a5 should have been blank ethnicity was unknown. e4 should have been left blank. g1 reporting contact fax number missing.
25315 | 1220648-2024-12493 | 19554782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554782&pc=OZD | 06/09/2024 | 07/23/2024 |  | 06/17/2024 |  |  | 06/09/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468583 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 30-year-old male patient with non ischemic cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient¿s right axillary graft and pocket became infected during support. medical staff then decided to remove the 5.5, clean the site, and implant a new 5.5 via left axillary access. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 added health effect - impact code 4641 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12493: g1 reporting contact fax number missing e4 should have been left blank.
25316 | 1220648-2024-12494 | 19554790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554790&pc=OZD | 06/09/2024 | 08/19/2024 |  | 06/17/2024 |  |  | 06/09/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491507 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 39-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient experienced oozing at the groin. medical staff performed a computed tomography (ct) scan, which revealed a retroperitoneal hematoma. the repositioning sheath was protruding approximately 2.5 cm from the puncture site and was not sutured, which contributed to further oozing. medical staff administered two units of red blood cells to the patient. medical staff resolved the oozing by repositioning the repositioning sheath and by placing sutures. the patient remained on support until the eventual explant. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was returned for evaluation. the returned product was visually inspected and no abnormalities were found. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, after the repo-sheath was repositioned and sutured in place, the bleeding problem has been resolved as per the clinical description provided. g1-corrected mfr site name and address.
25317 | 1220648-2024-12495 | 19554793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554793&pc=OZD | 06/09/2024 | 07/31/2024 |  | 06/17/2024 |  |  | 06/09/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469299 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 66-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that, after the implant, the cp had saturated flows. medical staff noted high purge pressures, high motor currents, and activated clotting time of 200 seconds. medical staff then decided to remove and replace the cp. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for saturated flow has been completed. impella pump was not returned for investigation. data logs were reviewed and second motor current spike during ramp up was observed at 0.02 second after pump start. saturated flows found immediately post the event. act was not maintained per recommendation during impella insertion. the cause of the saturated flow issue was most likely biomaterial per log analysis and clinical information.
25318 | 1220648-2024-12496 | 19554800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554800&pc=OZD | 06/10/2024 | 08/13/2024 |  | 06/17/2024 |  |  | 06/10/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025481009 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a male patient of unknown age post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 was unable to pass the anastomosis. the 5.5 case was aborted. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the pump did not advance due to frictional resistance of the artificial blood vessel, and insertion was discontinued as per the clinical description provided. the instructions for use referenced in the initial report is for the impella 5.5 for use during cardiogenic shock in the following sections: section: use of impella with transcatheter aortic valves.
25319 | 1220648-2024-12497 | 19554811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554811&pc=PYX | 06/03/2024 | 07/18/2024 |  | 06/17/2024 |  |  | 06/10/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025471857 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the impella flows diminished over time. staff suspected a blood clot, however a clot was never confirmed. the heparin had been stopped for over nine hours post impella placement to perform a washout. blood products were given but not due to low flow issue. the impella rp was replaced with a new impella rp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for low pump flow has been completed. ow pump flow reported by clinical on day 2 of support. the data logs confirm low pump flows for matching p-levels with continuous suction aligning with the clinical detail of the stopped heparin administration. the product was inspected and biomaterial was found around the impeller. the low pump flow reproduced in the lab. the root cause of the low pump flow was determined to be biomaterial due to ingestion. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12497: f6/f8-should have been left blank. g1 reporting contact fax number missing.
25320 | 1220648-2024-12467 | 19554879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554879&pc=OZD | 05/19/2024 | 07/22/2024 |  | 06/17/2024 |  |  | 05/19/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441231 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support as a bailout for tachycardia arrest status post left anterior descending artery stent. the patient developed hemolysis following impella cp implant. the patient's urine became bloody and the labs came back hemolyzed. repositioning was performed, but the hemolysis persisted. the plans to start dobutamine were made and additional intervention was not performed. the patient's family decided to withdraw care and the impella cp and intravenous medications were shut off. the patient expired. upon abiomed's medical safety officer review, there is not have enough information to associate the use of the device with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. neither the device nor data logs were returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12467 and should not have been. h.6 added code 925 in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12467 was submitted,
25321 | 1220648-2024-12468 | 19554882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19554882&pc=OZD | 05/19/2024 | 07/12/2024 |  | 06/17/2024 |  |  | 05/19/2024 | 06/17/2024 |  |  |  |  | 06/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489201 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post right coronary artery utilizing 5 drug eluting stents. while on support, the patient had no flow in the right lower extremity. the impella was explanted. care was withdrawn and the patient expired. upon review by abiomed's medical safety officer, there is not enough information to associate the use of the device with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. data logs were not provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determine f6 and f8 should have been blank. a5 checked not hispanic/latino (inadvertently omitted in the initial mdr.
25322 | 1220648-2024-12455 | 19544120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19544120&pc=OZD | 04/04/2024 | 10/25/2024 |  | 06/15/2024 |  |  | 06/05/2024 | 06/15/2024 |  |  |  |  | 06/15/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1581781 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility reported a 74-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that on day 7 the pump stopped. the pump stop prompted pump explant. upon explant thrombosis was seen at the pump. the pump was not replaced as the weaning was ongoing at the time of the pump stop. | [5 empty] | The investigation for the console (aic) pump stop has been completed. data logs from the console confirmed pump stop and pump restart attempts from the date of complaint. real time logs confirm that the motor current drops to zero without over current load reflecting a pump stop driven by the aic. the console was returned for evaluation and upon evaluation, was run at all p-levels with a similar pump without replicating the failure mode. power supply from the pbm to the impellatronics board was tested to be within specification. the root cause for the pump stop could not be determined due to not being reproduced. added- d9 device return date corrections to the initial mfr report: f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25323 | 1220648-2024-12425 | 19536826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19536826&pc=OZD | 05/17/2024 | 07/26/2024 |  | 06/14/2024 |  |  | 05/17/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487913 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a male patient of unknown age in st (sinus tachycardia) elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the patient had relevant bleeding at the puncture site, red blood cells (rbc) were administered, and the problem was resolved. the groin was dry and there was no more bleeding. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
25324 | 1220648-2024-12426 | 19537175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537175&pc=OZD | 05/17/2024 | 07/09/2024 |  | 06/14/2024 |  |  | 05/17/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025452703 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 76-year-old male of unknown race and ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that a left ventricular thrombus was identified on echocardiogram. the decision was made to withdraw care and the patient expired. per abiomed's medical safety officer review, there is not enough information to associate use of device with outcome. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25325 | 1220648-2024-12427 | 19537357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537357&pc=OZD | 05/17/2024 | 08/16/2024 |  | 06/14/2024 |  |  | 05/17/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1461907 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that following the implant, it was unable to start due to an automated impella controller (aic) error. a new console was obtained, and the pump was transferred over successfully. the patient condition declined, and the decision was made to withdraw care and the patient expired. per abiomed's medical safety officer review, there is not enough information to associate use of device with outcome. | [5 empty] | The investigation for the console (aic) yellow alarm issue has been completed. the data logs were reviewed which align with a transient loss of optical data. the aic logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. the aic was returned for evaluation. upon functional testing, the failure mode was not reproduced while running a known good pump. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue. f.6 and f.8 dates were reported on manufacturer device report number and should not have been. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
25326 | 1220648-2024-12428 | 19537360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537360&pc=OZD | 05/17/2024 | 07/09/2024 |  | 06/14/2024 |  |  | 05/17/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489427 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed ischemia in the left leg following placement. the pulse in the left leg became absent, and the flow had stopped in the extremity. there was no dopplerable pulse beyond the popliteal artery, and the leg became postured and stiff. the patient was escalated to impella 5.5 support, and a vascular surgeon was consulted for potential completion of a fasciotomy. days following, the patient's family made the decision to withdraw care, medical safety review outcome noted there is not enough information to associate use of the device with the patient's demise. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25327 | 1220648-2024-12430 | 19537406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537406&pc=OZD | 05/20/2024 | 07/24/2024 |  | 06/14/2024 |  |  | 05/20/2024 | 06/14/2024 |  |  | 10/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412490 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient facing an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding which was successfully managed with access suture stitches. the patient received one blood bag. the patient's hemoglobin on the following day was 6.2. and was noted to be hemodynamically stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most use issue likely due to use issue since hemostasis was achieved by adding sutures on the access site. g1-corrected mfr site name and addressh6 component code 4755 changed to 925 | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ifu statements: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25328 | 1220648-2024-12429 | 19537411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537411&pc=OZD | 05/17/2024 | 07/29/2024 |  | 06/14/2024 |  |  | 05/17/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a patient began to have red urine during impella cp support. the physician arrived to reposition the unit, but the decision was ultimately made to explant the pump. the patient was considered stable following the event. two days following the explant, the decision was made to withdraw care and the patient expired. per abiomed's medical safety officer review, there is not enough information to associate use of device with outcome. | [5 empty] | The investigation hemolysis has been completed. based on the clinical details, it is unknown if the pump was in the correct position. the data logs were reviewed and found that suction alarms occurred. no motor current spike was observed. no positioning issue indicated the cause of the reported issue. it is unclear if the hemolysis issue was related to positioning issue as the pump was explanted before repositioning. the cause of the hemolysis issue could not be determined since the complaint device was not returned for investigation and insufficient clinical data was provided. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12429 and should not have been. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2024-12429 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25329 | 1220648-2024-12431 | 19537463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537463&pc=OZD | 05/20/2024 | 07/29/2024 |  | 06/14/2024 |  |  | 05/21/2024 | 06/14/2024 |  |  | 10/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024415828 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient experiencing postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient has gastrointestinal (gi) bleeding due to a known disease. the bleeding persisted, and the patient has already been given several blood products. the team is aware of the too low act levels, but due to the patient's condition it could not be increased. the clinical staff discussed with a planned to remove the device as the clinical staff would like to decrease anticoagulation further and considers this to be too "dangerous" with the impella pump in place. it was noted there is no bleeding around impella access site. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient condition because of severe gastrointestinal bleeding and act values are too low of 140 than the therapeutic range. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ g1-corrected mfr site name and address. h6-component code 4755 changed to 925. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25330 | 1220648-2024-12432 | 19537522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537522&pc=OZD | 05/21/2024 | 07/23/2024 |  | 06/14/2024 |  |  | 05/21/2024 | 06/14/2024 |  |  | 08/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394999 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during an implant procedure for a patient experiencing acute myocardial infarction/cardiogenic shock, red alarms (impella stopped/retrograde flow/high current) appeared. four attempts were made to fix the issue, including changing the automated impella controller (aic) but the issue did not resolve. therefore, the pump was explanted and another device was used. there was no report of harm to the patient. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for pump did not start has been completed. the pump was returned for evaluation. upon inspection, red lumen was still present in the pump. electrical testing was performed and all lines were within specification. pump was plugged into automated impella controller (aic) and able to purge without issue. data logs were reviewed and rt logs show that pump was completing purge procedure while pump was outside of body. pump was attempted to be started while outside of the body and "impella stopped - motor current high" alarm was triggered. clinical details noted that pump was never inserted in the patient. the root cause of the pump did not start was due to use as the pump was attempted to be run outside of the body with the red lumen still inserted.
25331 | 1220648-2024-12433 | 19537600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537600&pc=OZD | 05/22/2024 | 07/23/2024 |  | 06/14/2024 |  |  | 05/22/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473146 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high-risk cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. ten minutes post implant, there was a sudden alarm due to high purge pressure. there were no observed kinks in the purge lines. staff contacted customer support and mentioned that the catheter had been clamped with a metal clamp shortly before the alarm occurred. customer support advised replacing the pump. a new impella 5.5 was successfully implanted. there was no report of harm to the patient. | [5 empty] | The investigation for high purge pressure has been completed. the impella 5.5 pump was returned. visually inspected and no biomaterial blockage in purge gap or catheter kink was observed. pump and cassette were connected to the console and reproduction testing was done. high purge pressure did not reproduce. pump was running at p-9 and no purge pressure issues were observed during testing. no other abnormalities were found. data logs were reviewed and gradual rise in purge pressure with saturated flows at start and low pump flows later. several high purge pressure alarms were recorded during support. the root cause of the high purge pressure issue was biomaterial per log analysis and since issue did not reproduce with return product. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
25332 | 1220648-2024-12434 | 19537741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19537741&pc=OZD | 05/22/2024 | 07/25/2024 |  | 06/14/2024 |  |  | 05/22/2024 | 06/14/2024 |  |  | 09/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024401392 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient experienced ischemia; therefore, the impella cp was exchanged to an impella 5.5. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. h6-component code 4755 changed to 925.
25333 | 1220648-2024-12388 | 19538865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19538865&pc=OZD | 05/13/2024 | 07/10/2024 |  | 06/14/2024 |  |  | 05/13/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient exhibited no distal pulses and cold lower extremities without pulses and hemolysis. the impella was repositioned and sent to have a computed tomography scan and was being treated with dialysis. due to the poor prognosis of the patient, the family decided to withdraw care and the patient expired. per abiomed's medical safety officer review. there was no association between impella use and outcome. | [5 empty] | The investigation for the reported limb ischemia and the hemolysis has been completed. product was not returned for investigation. partial data logs were returned and investigated. there were no motor current spikes seen in the partial logs. the root cause of the hemolysis issue was not determined since product was not returned, partial data logs were missing, and insufficient clinical information was provided. the cause of the limb ischemia could not be determined without additional clinical details. added- h6 component code 925 f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25334 | 1220648-2024-12387 | 19538866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19538866&pc=OZD | 05/29/2024 | 07/29/2024 |  | 06/14/2024 |  |  | 05/29/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465184 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a 61-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that, after the implant, the impella would not position correctly ¿ the pigtail appeared to be in the papillary muscle. medical staff attempted to reposition the impella several times and the flows would not go past 0.5-1 liters per minute without suction. the cp was then removed and replaced. the patient survived the event. | [5 empty] | The investigation for the low pump flow and the positioning issues has been completed. product was not returned for evaluation. data logs were reviewed which showed after pump started, motor current (mc) spiked occurred followed by low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. no positioning alarm was triggered during the support. this confirmed the reported failure. based on clinical description, the root cause of positioning issue was most likely use related. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
25335 | 1220648-2024-12389 | 19538869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19538869&pc=OZD | 05/29/2024 | 08/08/2024 |  | 06/14/2024 |  |  | 05/29/2024 | 06/14/2024 |  |  | 11/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422142 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Detachment of Device or Device Component |  |  |  |  | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a 54-year-old male patient with post cardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that the pump stopped on day 40. the stop occurred after the patient stood up to ambulate and damaged the pump by pulling and separating the plug. the pump was removed after pulling the remaining portion of the pump into the aorta. the patient survived the event. | [5 empty] | No product was returned. data logs were reviewed and 115 pump stop alarms were observed but no motor current spike greater than 30000 was seen. alarm with motor current spike was likely not reflecting due to sudden or immediate breakage of the motor cable lines and stoppage of the pump. gradual rise in purge pressure with no motor current spike was seen. no high purge pressure alarms were found to be triggered. tissue plasminogen activator (tpa) was added to resolve the issue and the purge trend gradually reduced to normal. the purge event was observed 16 hours prior to the pump stop and not relevant to it. as per clinical description, the patient was getting ready to ambulate and stood up prior to the nurse being ready with the impella cord and likely caught under something and pulled apart. per the doctor comments, the impella catheter was pulled from the red plug and separated at the gray connector resulting in failure. there was no likely product damage bond failure with excessive force being applied due to patient movement in support leading to pump stop with open electrical lines. the cause of the pump stop issue was most likely an open electrical line per clinical information and log analysis. the cause of the rise in purge pressure issue was most likely due to biomaterial. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12389 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12389 was submitted. h.6 codes 1762 and 4628 were reported incorrectly on the initial manufacturer device report number 1220648-2024-12389. a new code has been added to health effect - clinical code. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of initial manufacturer device report 1220648-2024-12389 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25336 | 1220648-2024-12435 | 19539260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539260&pc=OZD | 05/16/2024 | 07/24/2024 |  | 06/14/2024 |  |  | 05/16/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025491190 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The complainant reported a patient (race unknown) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support secondary to worsening mental status due to cardiogenic shock. the patient's mental status had been waxing and waning with severe hyponatremia over prior six days. during support, the patient developed hematoma at the axillary impella site. bival temporarily was held and one unit of packed red blood cells was transfused. it was further reported that the patient went in to sustain ventricular tachycardia and was unable to be resuscitated. the patient expired. abiomed does not have enough information to associate use of device with outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the ventricular tachycardia (vt) and the access site bleed has been completed. the impella device was not returned for evaluation. clinical details were provided which were sufficient to determine the root cause of the access site bleed. the root cause of the access site bleeding was most likely anticoagulation as hemostasis was achieved by stopping bival temporarily. as the necessary information was not provided, the root cause of the vt was not determined. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12435 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12435 was submitted. h.6 code 4641 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12435 and a new code has been added to health effect - impact codes. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12435 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
25337 | 1220648-2024-12436 | 19539305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539305&pc=OZD | 05/20/2024 | 07/18/2024 |  | 06/14/2024 |  |  | 05/20/2024 | 06/14/2024 |  |  | 11/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425172 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that complications were encountered with the introducer following the successful implant of an impella cp pump. it was stated that the peel away sheath failed to crack / peel away from around the catheter and had to be removed with scissors. there was no patient harm resulting from the failure. following the implant, it was also noted that a hematoma developed around the access site. femostop was placed to manage the condition and support with the implanted pump was maintained. (not enough information to associate use of device with outcome.) | [5 empty] | The investigation for the difficult insertion issue has been completed. the introducer was returned and insufficient score lines were observed. the introducer was returned in 2 uneven pieces (as stated by the clinical to have been cut with scissors) separated in a different location than the score lines. the root cause of the introducer damage issue was likely due to insufficient score lines leading to sheath peeling issue. added- d9 device return date d4-corrected model number. f6/f8 should have been left blank in the initial report. | The impella device was received from the customer on 29-may-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
25338 | 1220648-2024-12437 | 19539353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539353&pc=OZD | 05/18/2024 | 07/09/2024 |  | 06/14/2024 |  |  | 05/18/2024 | 06/14/2024 |  |  | 03/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024321779 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed to support a patient with cgs/ami (cardiogenic shock / acute myocardial infarction¿) and cardiomyopathy. the pump was placed but the patient suffered multiple vt (ventricular tachycardia) arrests and was shocked. they attributed the vt to a malpositioned impella pump. the patient expired on support the following day. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. added-h6 impact code 1802, d4-corrected model number f6/f8 should have been left blank in the initial report. g1 reporting contact fax number missing. h5 changed to yes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25339 | 1220648-2024-12438 | 19539362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539362&pc=OZD | 05/23/2024 | 07/31/2024 |  | 06/14/2024 |  |  | 05/23/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466023 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support and low pump flow was encountered. after implant a thrombus was seen in the left atrium; however, this could later not be found with transthoracic echocardiogram/transesophageal echocardiogram. the impella cp pump showed low flow with continuous suction. the pump positioning was noted as good and preload sufficient. the working diagnosis noted thrombus in the cannula. the physician's plan was noted as remove the pump possibly with a sentinel/triguard device, no new pump would be placed, the patient continued with extra-corporeal membrane oxygenation (ecmo) support, and low dose inotropes based on hemodynamics. the pump was successfully explanted, and upon explant, there no visible thrombus in the cannula. the patient was continued with ecmo support as per the plan and noted to be hemodynamically stable. | [5 empty] | Data was not returned for review. pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue under bench test. the root cause of the low flow was most likely a patient condition as no issue was found on the pump and position was reported to be good. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
25340 | 1220648-2024-12439 | 19539381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539381&pc=OZD | 06/02/2024 | 09/25/2025 |  | 06/14/2024 |  |  | 06/02/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/02/2024 |  |  |  |  | The user facility reported a 48-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and they did not pulses on their left foot. vascular surgery was consulted, but they opted to monitor. additionally, the placement signal monitoring was disabled. staff confirmed placement. furthermore, the patient had oozing at the access site. the site was re-sutured. | [5 empty] | The investigation for access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was not determined as insufficient clinical data was provided and the product was not returned. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12439. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12439 in accordance with updated procedures. d.6b revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12439. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12439 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12439 was submitted. h.6 codes 3038 and 925 were added in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-12439 was submitted. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12439 in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. placement monitoring was disabled on the first or second day of support due to alarms about placement in aorta that were incorrect according to position verification by echocardiography. an optical signal issue was suspected, but the root cause was not determined since the product was not returned and data logs were not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. g1 reporting contact email updated.
25341 | 1220648-2024-12297 | 19539398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539398&pc=OZD | 06/11/2021 | 03/25/2025 |  | 06/14/2024 |  |  | 06/11/2021 | 06/14/2024 |  |  | 02/28/2023 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022048712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the perclose proglide suture failed. angioseal was used and hemostasis was achieved. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25342 | 1220648-2024-12440 | 19539677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539677&pc=OZD | 06/03/2024 | 07/22/2024 |  | 06/14/2024 |  |  | 06/03/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471240 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. following the sheath insertion for impella 5.5 implant, the right radial arterial line waveform was lost. a 6 fr sheath was inserted to antegrade and an angiogram was performed. attempts were made to remove potential clot/plaque but flow remained absent on the radial line. with flow to the ulnar artery present, the decision was made to proceed with the impella 5.5 implant for patient support. | [5 empty] | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 added health effect - impact code 4624 f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25343 | 1220648-2024-12441 | 19539771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539771&pc=OZD | 06/04/2024 | 07/18/2024 |  | 06/14/2024 |  |  | 06/04/2024 | 06/14/2024 |  |  | 09/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024403595 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion of impella cp was unable to cross the aortic arch. the decision was made to remove the impella and perform the percutaneous coronary intervention without impella support. the patient still required mechanical circulatory support, and a new impella cp was successfully impanated. the patient was stable. | [5 empty] | The pump was returned for investigation. the root cause of the delivery issue is determined to be use issue because new pump inserted successfully. patient sent to icu on impella support and the patient was stable. d9 date device returned to manufacturer added f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25344 | 1220648-2024-12442 | 19539968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19539968&pc=OZD | 05/23/2024 | 07/31/2024 |  | 06/14/2024 |  |  | 05/23/2024 | 06/14/2024 |  |  | 03/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024331877 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a male of an unknown age, race, and ethnicity was put on impella cp support for high-risk percutaneous coronary intervention (pci). the pump was placed without issue however, the pump experienced a sudden decrease in flow. the patient became hemodynamically unstable and required cardiopulmonary resuscitation (cpr). the suction alarms persisted. the patient underwent pericardial effusion and pericardiocentesis and required cpr three times. the pump was ultimately explanted, and the patient was transferred to another hospital. the patient status is unknown. | [5 empty] | Returned complaint cp pump visually inspected. no biomaterial observed. no cannula kink observed. dried dextrose resolved around impeller gap. no broken blades found. pump connected to aic to reproduce low flow testing. pump run for more than 20 minutes at p-9. no pump low flow reproduced. pump flow was above specification required limit. the cause of the low pump flow issue most likely was patient condition related. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12442: f6/f8-should have been left blank . g1 manufacturing site address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25345 | 1220648-2024-12444 | 19540253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19540253&pc=OZD | 06/24/2023 | 07/12/2024 |  | 06/14/2024 |  |  | 06/04/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1537012 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a 60-year-old female patient with heart failure / cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the associated automated impella controller had purge alarms. medical staff remedied the situation by replacing the aic ¿ the computer replacement alone was not enough to resolve the issue. the patient has been on support for three weeks and remains on support. | [5 empty] | The investigation into the aic - purge issue ¿ other has been completed since the initial report was submitted. the console was returned and tested. upon visual inspection, found the dextrose deposits on the purge motor gasket. data logs were reviewed and the purge system blocked alarms were confirmed. the root cause for the purge system block was fluid ingress into the aic and the dextrose deposits on the purge motor gasket. d9 device return and date added f6 and f8 should have been blank. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
25346 | 1220648-2024-12447 | 19540689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19540689&pc=OZD | 05/27/2024 | 08/05/2024 |  | 06/14/2024 |  |  | 06/03/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025478929 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported that despite optimal impella cp insertion, the 75-year-old female patient developed a moderate hematoma in the right groin. the hematoma was managed conservatively with 1 unit of red blood cells. the pump was successfully explanted, and patient is continued with support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 component code 925 d4-corrected model number. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25347 | 1220648-2024-12446 | 19540708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19540708&pc=OZD | 05/23/2024 | 09/23/2024 |  | 06/14/2024 |  |  | 05/23/2024 | 06/14/2024 |  |  | 06/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427760 |  | 1000080 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump that was placed for support of a patient with a known corevalve. the pump had low flows and the pump support was discussed with md and abiomed rep. the pump support continued for a total of 45hours until family made choice to withdraw care. the pump was not replaced due to the plan to withdraw care. | [5 empty] | Correction to mdr mfr report #1220648-2024-12446-01: added-d10 added tavr as a concomitant medical product as the investigation found the fracture was determined to be characteristic of interaction with transcatheter aortic valve replacement (tavr). | The impella device was received from the customer on 30-may-2024 therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow. pump was returned and investigated. significant biomaterial was found in the cannula and around impeller blades. also, the impeller blades were found to be chipped and damaged after removal of the biomaterial. the root cause of the low pump flow issue was fractured impeller blade(s). the fracture was characteristic of an interaction with a transcatheter aortic valve replacement (tavr). added- b1 selected product problem. d4-corrected model number. d8-device available for evaluation changed to yes. f6/f8 should have been left blank in the initial report.
25348 | 1220648-2024-12451 | 19541427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19541427&pc=OZD | 05/28/2024 | 07/25/2024 |  | 06/14/2024 |  |  | 05/28/2024 | 06/14/2024 |  |  | 10/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025452923 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The complainant reported a 53-year-old male patient with a mitraclip was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the purge pressure was 1030 mmhg and the pump flow dropped to 2 ml. medical staff exchanged the purge cassette and added tissue plasminogen activator. medical staff noted there were no kinks in the purge line, and no patient harm has been reported. | [5 empty] | The investigation for high pressure purge has been completed. pump and data logs were not returned for investigation. high purge pressure was reported during case. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12451: f6/f8-should have been left blank. g1 manufacturing site name and address. g2 report source: foreign missing. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25349 | 1220648-2024-12453 | 19542046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542046&pc=OZD | 06/04/2024 | 08/01/2024 |  | 06/14/2024 |  |  | 06/04/2024 | 06/14/2024 |  |  | 11/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024425923 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported a 44-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump accessed limb became ischemic 2 days after insertion. the limb was dusky and so the team performed bypass via the extracorporeal membrane oxygenation (ecmo) loop. the patient remains on support and was reported as stable. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added-b5 mso review, d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25350 | 1220648-2024-12458 | 19542182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542182&pc=OZD | 06/06/2024 | 08/05/2024 |  | 06/14/2024 |  |  | 06/06/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471840 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hemorrhagic Stroke | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and they had a cerebral hemorrhage. the impella was removed and an intra-aortic balloon pump was placed. the relationship between the impella and the cerebral hemorrhage is unknown. reportable due to patient's cerebral hemorrhage and unknown relationship to the impella. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12458: b2 outcomes attributed to adverse event: required intervention to prevent permanent impairment/damage was inadvertently not selected. d4model number. f6/f8-should have been left blank. g1 reporting contact facility name missing. g1 reporting contact fax number missing. h.6 codes 1942 and 4641 were reported incorrectly. new codes have been added to health effect - clinical code and health effect - impact code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25351 | 1220648-2024-12459 | 19542240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542240&pc=OZD | 06/05/2024 | 07/23/2024 |  | 06/14/2024 |  |  | 06/05/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506480 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The complainant reported a 57-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the intensive care unit medical staff attempted to reposition the cp due to the pigtail and inlet having been caught in the mitral valve apparatus. medical staff then decided to go to the cardiac catheterization lab to attempt to manipulate the cp using fluoroscopy and echocardiography. the cp got pulled back into the aorta, and medical staff had to replace the cp. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for positioning issues has been completed. the impella device was not returned for evaluation. based on clinical description, the root cause of positioning issue was most likely use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12459: f6/f8-should have been left blank; g1 reporting contact fax number missing; h.6 code 2199 was reported incorrectly. new code was added to health effect - impact code.
25352 | 1220648-2024-12454 | 19542493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542493&pc=OZD | 06/04/2024 | 07/23/2024 |  | 06/14/2024 |  |  | 06/04/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442095 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | Tachycardia | Foreign Body Embolism | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | [8 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a 72-year-old male patient with acute myocardial infarction/cardiogenic shock had was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient would experience positional ventricular tachycardia (vt) ¿ when the patient sat up or ambulated, they would experience vt. additionally, tissue was ingested into the pump motor when attempting to deliver cardioplegia. the pump stopped and wouldn¿t restart, so medical staff replaced the pump. the patient survived the event. | [5 empty] | The investigation for vf/vt/af/at, thrombus, and pump stop has been completed. the data logs confirm pump stop alarms. there was a motor current spike that occurred while the pump was at p2 before the pump stop. the product was inspected and biomaterial was found in the cannula at inflow and outflow cages and around the impeller. attempts were made to clear the biomaterial which were unsuccessful. the pump was run in the lab and the pump stop reproduced with high motor current. the root cause of the pump stop was determined to be biomaterial. the necessary information was not provided, the root cause of the vt/vf was not determined. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12454: f6/f8-should have been left blank g1 reporting contact fax number missing h.6 code 4439, 1721, 4445, 4627, and 4644 were reported incorrectly. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25353 | 1220648-2024-12452 | 19542773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542773&pc=OZD | 02/22/2024 | 07/09/2024 |  | 06/14/2024 |  |  | 05/21/2024 | 06/14/2024 |  |  | 07/31/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443416 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured cardiac arrest with a "possible" relationship to the impella device used: ¿after impella removed on (b)(6), pt with hypotension then pea/bradycardia- cpr started. decision to place patient back on femoral va ecmo (venoarterial extracorporeal membrane oxygenation). altogether, appx 15 mins from hypotensive episode to flowing on ecmo (extracorporeal membrane oxygenation). decision not to replace impella as pt with 50 points pulsatility at this point¿. the patient recovered with sequelae. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. f6/f8 should have been left blank in the initial report. g1 reporting contact email corrected. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25354 | 1220648-2024-12448 | 19542775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542775&pc=OZD | 05/24/2024 | 07/29/2024 |  | 06/14/2024 |  |  | 05/24/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1113650 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Select Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown automated impella controller (aic) alarm occurred during impella pump implantation to an unknown aged patient. there was a purge disc not recognized message noted. the aic was successfully exchanged for new aic, and the impella pump was successfully used with no report of harm to the patient. | [5 empty] | The investigation for the console (aic) purge/flag issue has been completed. the aic was returned for evaluation. upon functional testing, the purge disc not detected issue was unable to be reproduced. there was no visible issues with the purge disc flag and retainer. data logs were reviewed finding the reported purge disc not detected issue was confirmed. there was one instance of purge disc not detected alarm (event set #402) found in the logs on the reported occurrence date. the alarm was cleared 12 seconds later once the purge disc was inserted. no problem found after testing. reported purge disc not detected issue with this console is suspected to be use related. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25355 | 1220648-2024-12450 | 19542796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542796&pc=OZD | 05/25/2024 | 08/01/2024 |  | 06/14/2024 |  |  | 05/25/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490889 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The using facility reported an unknown age patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp heart pump for mechanical circulatory support. the attending physician noted the impella flow was zero, and the pump position was checked. due to the suspected pump stop, the patient was taken back to the catheterization lab, where the impella cp pump was restarted and continued to run without problems with no report of harm to the patient. the impella cp pump was successfully explanted four days later. | [5 empty] | The investigation low pump flow has been completed. no abnormality or problem found. no data logs available for review. based on clinical data the root cause of the low pump flow issue most likely was improper positioning related due to improper technique as after repositioning the impella the flow went back to normal flow range at p-9. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
25356 | 1220648-2024-12460 | 19542877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19542877&pc=OZD | 06/05/2024 | 07/30/2024 |  | 06/14/2024 |  |  | 06/05/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504196 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic was implanted with an impella cp device for mechanical circulatory support and while on support experienced limb ischemia. the patient was admitted in with pulmonary embolism and deep vein thrombosis (dvt). the dvt was in the left leg. the patient was placed on both extra-corporeal membrane oxygenation and impella cp pump for support. the limb with the impella cp had to have an intervention for restoration of the flow. the team performed cutdown and aspiration as compartment syndrome was developing. the prognosis was poor to keep the limb. the patient was noted to be in stable condition. | [5 empty] | The investigation for the limb ischemia and purge cassette failure has been completed. the impella has not been returned for evaluation. however, data logs were reviewed which revealed a gradual rise in purge pressure was observed over a period of 5 minutes. the rise in purge pressure was observed to be a logarithmic rise. a motor current spike was also observed before the rise in purge pressure. data logs show the high purge pressure was able to resolve over a period of approximately 3 hours. clinical details noted that "high purge pressure" alarms were resolved after purge cassette was changed. purge solution used contained sodium bicarbonate. no kinks in line were observed by clinical team. pump was able to run for another 5 days before explant without issue. additionally, clinical details noted that the patient was experiencing limb ischemia before impella was inserted and intervention was performed once pump was implanted. the cause of the high purge pressure was biomaterial per data log review. the root cause of the limb ischemia could not be determined without additional information. added- h6 component code 925 d4-corrected model number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25357 | 1220648-2024-12449 | 19543038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19543038&pc=OZD | 06/04/2024 | 07/23/2024 |  | 06/14/2024 |  |  | 06/04/2024 | 06/14/2024 |  |  | 09/30/2025 |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446663 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post coronary intervention on left anterior descending artery. once the patient was in the intensive care unit, the team observed a blood clot in the left ventricle. the patient coded and inappropriate flows were noted. in addition, the impella cp came back across aortic valve after cardiopulmonary resuscitation. the physician removed the impella cp. there is currently no indication that impella support caused or contributed to the patient coding | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for low pump flow and positioning issues has been completed. the complaint device not returned for investigation. data log reviewed: suction alarms occurred. positioning issue alarms occurred. motor current spikes observed and clinical data confirmed clot in left ventricle (lv)noted. the flow didn't improved after repositioning performed. the cause of the low pump flow most likely was biomaterial ingest. the cause of the positioning issue most likely was use related due to improper technique as the pump went back to aorta. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12449: a5 should have been blank, ethnicity was unknown f6/f8-should have been left blank. g1 reporting contact fax number missing. h.6 code 2199 was reported incorrectly and 925 was inadvertently added twice.
25358 | 1220648-2024-12443 | 19543148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19543148&pc=OZD | 05/24/2024 | 08/16/2024 |  | 06/14/2024 |  |  | 05/24/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488830 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a 64-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient experienced sudden pump stop, the staff tried several times to restart the pump without success. the staff had a suspicion of clot; therefore, the pump was removed. the patient did not suffer any consequences and the patient was anticoagulated due to surgical heart intervention the day before. the patient is stable. | [5 empty] | D.6b explant date was received and was added. the investigation for the pump stop has been completed. the impella was returned for evaluation. no data logs were returned. the product was inspected and after clearing the dextrose buildup, a large impeller purge gap was observed. the pump was run in the lab and the pump stop reproduced. according to the clinical details, the pump stop occurred on day seven of support. the patient was not properly anticoagulated. the root cause of the pump stop was a bearing wear product malfunction as the pump stop occurred within design life. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.4 unique identifier (udi) # was revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12443. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12443 and should not have been. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer device report number 1220648-2024-12443. h.3 if the device was not evaluated should of reflected device evaluation anticipated, but not yet begun on the initial manufacturer device report number 1220648-2024-12443. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of initial manufacturer device report number 1220648-2024-12443 in accordance with updated procedures. | The impella device was received from the customer on (b)(6) 2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25359 | 1220648-2024-12457 | 19543166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19543166&pc=OZD | 06/05/2024 | 08/12/2024 |  | 06/14/2024 |  |  | 06/05/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509532 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella cp pump began to experience persistent suction alarms during support. troubleshooting was attempted, but the pump was unable to obtain flows above .2 lpm. as a result of the noted issue, the decision was made to remove the pump. there was no patient harm. | [5 empty] | D.9 updated as the pump was returned for investigation. the investigation low pump flow has been completed. the impella cp pump was visually inspected. no cannula kink/no biomaterial was observed. broken blades were found in the impeller. the root cause of low pump flow issue was a damaged impeller due to the interaction with other devices. a.5 ethnicity should have been left blank on the initial manufacturer device report number as it is unknown. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number. d.4 revised model number from the initial submission of initial manufacturer device report number in accordance with updated procedures. d.4 serial number was revised as it was previously entered incorrectly on the initial submission of initial manufacturer device report number. f.6 and f.8 dates were reported on manufacturer device report number and should not have been. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial submission of initial manufacturer device report number. g.2 foreign was inadvertently selected on the initial submission of initial manufacturer device report number. h.6 codes 2199 and 2954 were entered incorrectly on the initial manufacturer device report number. a new code was added to the health effect - impact codes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25360 | 1220648-2024-12445 | 19543338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19543338&pc=OZD | 05/24/2024 | 07/23/2024 |  | 06/14/2024 |  |  | 05/24/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488278 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a 61-year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported, after the removal of the sheath post-procedure, medical staff assumed the patient¿s puncture site to be closed. the next day, medical staff had to intervene due to bleeding problems. medical staff implanted a graft in the perforated groin. medical staff also noted an infection. the patient survived the event. | [5 empty] | The investigation for access site bleeding and infection/sepsis has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to use issue since the puncture was assumed to be closed after removal of pump causing the bleed. because this evidence has not been provided, the root cause of the infection cannot be determined the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12445: f6/f8-should have been left blank . g1 manufacturing site name and address. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿ acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25361 | 1220648-2024-12456 | 19543642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19543642&pc=OZD | 06/05/2024 | 07/24/2024 |  | 06/14/2024 |  |  | 06/05/2024 | 06/14/2024 |  |  |  |  | 06/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474230 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant was placing a 5.5 to support a patient admitted in and awaiting transplant. the 5.5 pump was to bridge to transplant. the pump had issues with delivery and stopped 3x in delivery/pump start. the pump was explanted and replaced. no lasting harm from delay in initiation. | [5 empty] | Pump was returned for evaluation. electrical testing was performed on the pump and all motor cable lines were found to be within a normal range. a small kink was observed just distal of the kink protector on the catheter side of the red handle. upon visual inspection of the inside of the pump cannula and outflow cage, the edge of the outflow cage was bent inwards in a small section. the impeller blade was not able to pass the deformity on the inside of the outflow cage. pump was purged in a bucket of water without issue and attempted to be started multiple times. each time a "impella stopped - motor current high" and "impella stopped (rpm deviations)" alarm were triggered. data logs were reviewed and rt logs from case show the pump was attempted to be started 5 times, each with a "impella stopped - motor current high" alarm. clinical details noted that the pump had difficulty passing through the chest wall due to patient's anatomy. the root cause of the delivery issues was most likely due to patient anatomy. the root cause of the pump did not start is due to a damaged outflow cage. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12456: b2 outcomes attributed to adverse event should have been left blank. d4 model number. f6/f8-should have been left blank. h.6 code 2920 and 2338 were reported incorrectly. new code was added to medical device problem code. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25362 | 1220648-2024-12421 | 19527779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19527779&pc=OZD | 05/17/2024 | 07/24/2024 |  | 06/13/2024 |  |  | 05/17/2024 | 06/13/2024 |  |  | 10/31/2025 |  | 06/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421039 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following successful high risk percutaneous coronary intervention for 67-year-old female patient on impella cp support, the two angioseal methods did not work. it was noted that compressing the puncture site was difficult due to adipose groin. the patient was given blood products. hemostasis was ultimately achieved with manual compression and loose application of a femstop. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12421: f6/f8-should have been left blank. g1 manufacturing site name and address. h.6 code 4755 was reported incorrectly. new code was added to component code.
25363 | 1220648-2024-12422 | 19527875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19527875&pc=OZD | 05/17/2024 | 07/23/2024 |  | 06/13/2024 |  |  | 05/17/2024 | 06/13/2024 |  |  | 11/30/2024 |  | 06/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023280581 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported an unspecified undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after 25 days of impella 5.0 support for female patient, decision was made to escalate to impella 5.5 as device had persistent low purge alarms. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for low pump flow has been completed. the impella device and data were not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12422: a5 race and ethnicity missing. f6/f8-should have been left blank . g1 manufacturing site name and address.
25364 | 1220648-2024-12423 | 19529592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19529592&pc=OZD | 05/17/2024 | 07/23/2024 |  | 06/13/2024 |  |  | 05/17/2024 | 06/13/2024 |  |  |  |  | 06/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488826 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 30-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the physician reported a suction alarm. the female patient was small (160 centimeters (cm)) with a small heart. the impella cp is three cm in the middle of the left ventricle, but according to him it is not convincing because the outlet basket is very far in the aorta, the right heart and the vena cava are not necessarily filled optimally, and the patient also has a pericardial effusion. the patient has already received two liters of volume. the placement signal (ps) curve showed breathing-dependent artifacts since no left ventricle (lv) value is displayed and the ps curve has breathing-dependent artifacts. the physician asked whether the impella might be defective, and this could not be confirmed but also not denied, since the lv value was only a calculated value. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for positioning issues has been completed. the returned complaint cp pump visually inspected. cannula kink observed near outlet cage. no other abnormality observed. the cause of the positioning issue was patient condition related since the patient had a small left ventricle resulted in cannula kink. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12423: f6/f8-should have been left blank.
25365 | 1220648-2024-12424 | 19530180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19530180&pc=OZD | 05/14/2024 | 07/12/2024 |  | 06/13/2024 |  |  | 05/14/2024 | 06/13/2024 |  |  | 09/30/2024 |  | 06/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280180 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The complainant reported that after connecting and prepping an impella cp for patient support, red alarms immediately occurred when the startup was initiated. the complainant reported that the impella cp stopped following a high motor current. a second pump was prepped and implanted without issue, and no patient harm has been reported. | [5 empty] | The investigation for the reported pump did not start issue has been completed. according to the clinical, immediately after beginning impella cp support, red alarms were recorded. a second pump was then opened, prepped, and ramped up with no issues. no product was returned for a visual inspection. data log analysis confirmed a high motor current pump stop immediately after the pump was started. the cause of the pump start issue was not determined because no product was returned and not enough clinical information was given. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: impella stopped. ¿if the impella catheter as stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ added-h6 impact code 4629 f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25366 | 1220648-2024-12352 | 19518247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518247&pc=OZD | 07/13/2021 | 03/20/2025 |  | 06/12/2024 |  |  | 07/13/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088035 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient exhibited hematuria and hemolyzed labs. it is unknown if any intervention was done for the hemolysis. the patient remained on impella support and was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
25367 | 1220648-2024-12353 | 19518272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518272&pc=OZD | 06/21/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 06/21/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072128 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/21/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ectopy when the pump was turned down to a pressure of less than nine. an echocardiogram performed at bedside revealed the pump was jumping back and forth and was worse at lower levels causing the ectopy. to resolve the issue, the physician advanced the device into the ventricle which stopped the ectopy, and then turned the patient¿s pressure to two. the physician monitored the patient¿s heart for a few minutes and noted the patient¿s condition was stable. the patient was successfully weaned from the impella and the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25368 | 1220648-2024-12354 | 19518417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518417&pc=OZD | 07/09/2021 | 03/20/2025 |  | 06/12/2024 |  |  | 07/09/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085736 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/09/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. prior to the insertion of the impella, a perclose device was attempted to be placed by snapped due do calcium in the patient's vessel. the impella was placed, the peel=away sheath was removed and an additional perclose device was inserted. the perclose devices were tamped, and bleeding was not resolved. an 8 french angioseal was inserted and an angiogram was performed that revealed no distal flow to the access site. vascular surgery was called and a femoral to femoral bypass was performed to restore distal flow. the impella was successfully explanted at the end of the percutaneous coronary intervention. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25369 | 1220648-2024-12355 | 19518633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518633&pc=OZD | 05/08/2024 | 07/09/2024 |  | 06/12/2024 |  |  | 05/16/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464707 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured sustained ventricular tachycardia (recurring) with a "possible" relationship to the impella device used: ¿the patient has a remote history of ventricular tachycardia (prior to index admission). however, the patient developed sustained ventricular tachycardia on 8 may 2024, for which they were cardioverted from vt to sinus rhythm with av (atrioventricular¿) pacing. on several days after this event, the patient developed recurring vt requiring antiarrhythmics, electrolytes, additional cardioversions, as well as adjustments to their existing aicd (automatic implantable cardioverter defibrillator¿) settings for overdrive pacing.¿ it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
25370 | 1220648-2024-12356 | 19518697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518697&pc=OZD | 04/30/2024 | 06/28/2024 |  | 06/12/2024 |  |  | 05/18/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494468 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured loss of flow in the left superficial femoral artery with a "definite" relationship to the impella device used: ¿successful peripheral interventional revascularization of the proximal to mid left profunda femoris artery and left superficial femoral artery. after impella sheath removal and balloon tamponade there was now flow in the sfa (superficial femoral artery) requiring pta with 8mm armada balloon then flow was lost in the profunda requiring pta with 4.0mm coronary balloon. with good flow min both vessels at the end of procedure impella placed in the left common femoral artery, however distal flow in the lsfa was lost and impella/sheath were removed, and hemostasis achieved with balloon tamponade. pta of the lsfa and left profunda were required to restore flow left sided impella due to hypotension post-pci that was removed due to vascular compromise and replaced with iabp on the right side¿. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. removed f6 and f8.
25371 | 1220648-2024-12357 | 19518852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518852&pc=OZD | 07/01/2021 | 03/20/2025 |  | 06/12/2024 |  |  | 07/01/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082652 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported the patient¿s pedal pulses were not palpable. doppler was going to be performed to confirm pulses. it is unknown how the limb ischemia was resolved. the patient survived the procedure. | [5 empty] | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25372 | 1220648-2024-12358 | 19518926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518926&pc=OZD | 07/03/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/03/2021 | 06/12/2024 |  |  | 03/31/2022 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1474716 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis secondary to pump positioning issues. hemolysis resolved with pump removal. additionally, the patient experienced a large clot which was observed upon pump removal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
25373 | 1220648-2024-12359 | 19518970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19518970&pc=OZD | 07/11/2021 | 03/20/2025 |  | 06/12/2024 |  |  | 07/11/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080400 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the device was found to be out of position in the aorta. the physician was able to advance the device back into a better position. however, when advancing the device, the patient had a short run of ventricular tachycardia (vt). no intervention was given for the vt, and it resolved spontaneously. the patient remained on impella support and was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25374 | 1220648-2024-12360 | 19519012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519012&pc=OZD | 07/03/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/03/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022062640 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction alarms that were resolved by administering blood products and volume to the patient. additionally, the patient experienced access site bleeding that was resolved by withholding heparin. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25375 | 1220648-2024-12361 | 19519033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519033&pc=OZD | 06/04/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 06/04/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067428 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding with heparin being removed from the purge and systemic. the patient was successfully weaned off of support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25376 | 1220648-2024-12362 | 19519178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519178&pc=OZD | 07/06/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 07/06/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022053011 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, groin bleeding and atrial fibrillation were noted. heparin was put on hold temporarily and levophed was restarted. the patient continued on support until the device was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25377 | 1220648-2024-12363 | 19519242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519242&pc=OZD | 07/06/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 07/06/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043582 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support bleeding and possible continuous renal replacement therapy was noted. heparin was withheld. the patient was successfully weaned and the device was explanted. the patient survived the procedure and remained on extracorporeal membrane oxygenation. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25378 | 1220648-2024-12364 | 19519370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519370&pc=OZD | 06/04/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 06/04/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069838 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon insertion of pigtail, the patient experienced ventricular tachycardia, also, oozing from the impella site was noted. no further resolution provided for ventricular tachycardia or oozing at impella site. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | D6: explant date unknown the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25379 | 1220648-2024-12365 | 19519431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519431&pc=OZD | 07/06/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 07/06/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced a drop in platelets, bleeding, and deep pump positioning was noted. a heparin-induced thrombocytopenia panel was sent and the physician decided to replace heparin with sodium bicarbonate in purge system. two units of blood were administered. the impella cp device was also repositioned. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25380 | 1220648-2024-12366 | 19519613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519613&pc=OZD | 03/22/2024 | 06/12/2024 |  | 06/12/2024 |  |  | 05/20/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447945 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/20/2024 |  |  |  |  | The user facility in japan reported a 61-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the purge pressure high alarm would not go off. it was confirmed over the phone that there were no kinks in the purge line. as a precaution, the purge set was replaced, and the situation was monitored. after replacing the purge set, the flow rate was about 2.5ml/hr, and the purge pressure high alarm disappeared. there was no patient harm. investigation results indicates that the severity of this complaint is s3. | [5 empty] | Data logs and the pump were not returned. the purge cassette had no visual abnormalities. when run in the lab with 2 separate substitute pumps, the high-pressure purge (hpp) was not reproduced, and the purge cassette purged as expected. there were no purge cassette line kinks found. the root cause of the high purge pressure issue was not determined since the pump and data logs were not returned.
25381 | 1220648-2024-12367 | 19519651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519651&pc=OZD | 06/04/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 06/04/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055294 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support, the patient experienced several episodes of ventricular tachycardia (vtach) initially, along with a loss of distal flow and the observation of a large left ventricle thrombus. to address the loss of distal flow, a fem-fem bypass was performed, and the vtach episodes were promptly corrected through defibrillation. the patient was successfully weaned off of support. | [5 empty] | The cause of the limb ischemia, vf/vt/af/at, and thrombus was not determined since product was not returned and all the necessary information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: limb ischemia ventricular tachycardia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25382 | 1220648-2024-12368 | 19519666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519666&pc=OZD | 05/25/2024 | 07/18/2024 |  | 06/12/2024 |  |  | 05/25/2024 | 06/12/2024 |  |  | 06/30/2025 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373237 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a patient (race unknown) was to be implanted with an impella cp device for mechanical circulatory. . when the pump was advanced into the left ventricle, the pump was kinking. the impella was explanted and a new pump was used. there was no patient harm. | [5 empty] | Product was returned and investigated. there was a kink noticed near the outlet of the pump. per the clinical information, cannula kinking was observed after successful insertion once the pump was past the aortic valve. therefore, the root cause of the product damage (cannula kink) was most likely use-related due to improper technique. d9 date device returned to manufacturer added f6/f8-should have been left blank in the initial report. g1 reporting contact fax number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25383 | 1220648-2024-12369 | 19519667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519667&pc=OZD | 07/06/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 07/06/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022051770 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, there were noted positioning issues that were resolved by explanting and re-implanting the impella cp device. it was also noted that the patient was escalated to the impella 5.5 device. the patient survived the procedure. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25384 | 1220648-2024-12370 | 19519733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519733&pc=OZD | 05/28/2024 | 07/16/2024 |  | 06/12/2024 |  |  | 05/28/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450442 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was removed due to hematoma at the impella site. three units of prbc (packed red blood cells) were given. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25385 | 1220648-2024-12371 | 19519749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519749&pc=OZD | 07/07/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 07/07/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047303 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hematuria as the patient¿s lab results were positive for hemolysis. the device was repositioned as the inlet was stated to be in the middle of the ventricle. the patient¿s urinary output remained dark, leading to another device reposition. the patient continued on support until the device was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25386 | 1220648-2024-12372 | 19519777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519777&pc=OZD | 05/17/2024 | 07/17/2024 |  | 06/12/2024 |  |  | 05/17/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025481236 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 77-year-old female in st elevated myocardial infarction with ventricular tachycardia was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was being moved and the pump stopped. the team troubleshot by replacement of the automated impella controller, which worked. the patient still expired despite the pump support and with unknown neurological status the decision was made to withdraw care by the family. pms team made the decision at the discretion of management to file as death. the death event occurred and no alternative cause for the death event was identified as the likely cause for the patient outcome. | [5 empty] | The investigation for the pump stop has been completed. no product was returned for evaluation. data logs were reviewed and logs from first console used show a motor current (mc) spike and immediate pump stop. pump was attempted to be restarted multiple times without success. pump was switched to a second aic and pump was able to run for approximately an hour while patient coded multiple times. data logs from second aic show that motor current, pump flow, and placement signal were all variable throughout logs as patient was decompensating and coded multiple times. the root cause of the temporary pump stop cannot be determined as no product was returned. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: impella stopped: if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings and cautions: to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25387 | 1220648-2024-12373 | 19519800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519800&pc=OZD | 05/17/2024 | 03/19/2025 |  | 06/12/2024 |  |  | 05/17/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1087349 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 77-year-old female in st elevated myocardial infarction with ventricular tachycardia was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was being moved and the pump stopped. the pump also had an automated impella controller alarm. the team troubleshot by replacement of the automated impella controller, which worked. the patient still expired despite the pump support and with unknown neurological status the decision was made to withdraw care by the family. pms team made the decision at the discretion of management to file as death. the death event occurred and no alternative cause for the death event was identified as the likely cause for the patient outcome. | [5 empty] | The investigation of automated impella controller (aic) ¿ pump stop has been completed. the console logs from the reported day of event are consistent with complaint. the console was tested, but the issue was not reproduced. because the console operated within specification during testing, it¿s been concluded that the pump stop was unrelated to the console. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-12373 in accordance with updated procedures. f6/f8-should have been left blank on initial manufacturer device report 1220648-2024-12373. g1 reporting contact email revised on manufacturer device report 1220648-2024-12373 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25388 | 1220648-2024-12374 | 19519865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519865&pc=OZD | 06/06/2021 | 03/24/2025 |  | 06/12/2024 |  |  | 06/06/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022051831 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was noted that the procedure was aborted due to vascular anatomy and that the impella cp pump was discarded during cpr and massive kinking, a new impella was used. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25389 | 1220648-2024-12375 | 19519919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519919&pc=OZD | 07/08/2021 | 03/21/2025 |  | 06/12/2024 |  |  | 07/08/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084658 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced a hematoma in the medial thigh. manual pressure was applied to the site with no success. it was noted that blue suture hub had advanced into tissue tract. the physician stated that sheath exchange and securement went without issue. the decision was made to wean the patient off support and explant the pump. the impella cp device was explanted, manual pressure was applied and femostop was used to achieve hemostasis. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25390 | 1220648-2024-12376 | 19519966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19519966&pc=OZD | 07/19/2021 | 03/21/2025 |  | 06/12/2024 |  |  | 07/19/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089558 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. patient experienced three failed percloses. manual pressure was applied for 45 minutes to successfully achieve hemostasis. in addition, there was access site bleeding at the groin due to patient sitting too early after the procedure. manual pressure was reapplied, and two units of packed red blood cells were given. the impella was explanted in the procedural area. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25391 | 1220648-2024-12377 | 19520028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520028&pc=OZD | 06/10/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 06/10/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071326 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. therefore, the physician made the decision to remove the impella. reportedly after removal of the impella and manual pressure the groin site was clean, dry, and intact and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25392 | 1220648-2024-12378 | 19520214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520214&pc=OZD | 06/17/2021 | 10/29/2025 |  | 06/12/2024 |  |  | 06/17/2021 | 06/12/2024 |  |  | 09/30/2022 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021006164 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/17/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. due to suspicion that patient was experiencing limb ischemia as well as return of pulse pressure with weaning, the physician opted to remove impella and start the patient on levophed. it is unknown if the actions resolved the limb ischemia. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | F6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-12378 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-12378 in accordance with updated procedures. h6 component code and investigation conclusions revised from the initial submission in accordance with updated procedures.
25393 | 1220648-2024-12379 | 19520224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520224&pc=OZD | 06/10/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 06/10/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050244 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2021 |  |  |  |  | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding and hematoma requiring an unknown amount of transfused packed red blood cells. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25394 | 1220648-2024-12380 | 19520393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520393&pc=OZD | 06/14/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 06/14/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063202 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/14/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hematoma to the right groin from the procedure that was demarcated and soft. to resolve the hematoma fem-stop for pressure and two units of packed red blood cells were administered. the impella cp was explanted in the lab due to retroperitoneal bleed. a new impella cp with smart assist was placed in the left groin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ vascular dissection: impella cp with smart assist for use during cardiogenic shock and high risk pci & impella 5.5 with smart assist for use during section: warnings & cautions: cautions as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25395 | 1220648-2024-12381 | 19520599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520599&pc=OZD | 05/14/2024 | 06/12/2024 |  | 06/12/2024 |  |  | 05/14/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490005 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Ischemia | [11 empty] | Death | Life Threatening | Required Intervention | Disability |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient facing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced irreversible limb ischemia and an embolism. the patient subsequently expired. the patient was placed on venous-arterial extra-corporeal membrane oxygenation support one hour prior to the impella cp pump insertion. the patient experienced limb ischemia and a perfusion catheter was placed to restore flow. on day three of the support, the team performed an above-knee amputation. on day seven, there was an air in the purge system alarm and troubleshooting was initiated. however, the decision was made to withdraw care and the patient subsequently expired. upon review by abiomed's medical safety officer, there is not enough information to exclude the impella cp pump as an associated factor in the patient's outcome. | [5 empty] | The investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. the root cause of the embolism could not be determined since the product was not returned and insufficient clinical details were provided. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25396 | 1220648-2024-12382 | 19520815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520815&pc=OZD | 06/16/2021 | 03/21/2025 |  | 06/12/2024 |  |  | 06/16/2021 | 06/12/2024 |  |  | 12/31/2022 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022025665 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hematuria. the physician ordered fluid and stopped integrilin platelet aggregation inhibitor. patient's condition was stable post procedure. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25397 | 1220648-2024-12383 | 19520990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19520990&pc=OZD | 07/20/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/20/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022058099 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2021 |  |  |  |  | The complainant reported that a 81 year old female patient was presented with ventricular tachycardia (vt). the patient taken to the electrophysiology (ep) laboratory for impella supported vt ablation. the impella cp was placed in the left femoral artery (lfa) without issue. bleeding was observed from all access sites including the implantable cardioverter defibrillator (icd) pocked. four (4) units of packed red blood cells (prbc¿s) were transfused to the patient achieving hemostasis. a fem stop was also utilized at the site of impella access. it was reported that there were no doppler pulses at either foot. it is unknown if doppler pulses returned to the lower extremities. the patient experienced several episodes of vt, and the icd was adjusted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25398 | 1220648-2024-12384 | 19521003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521003&pc=OZD | 05/13/2024 | 06/27/2024 |  | 06/12/2024 |  |  | 05/13/2024 | 06/12/2024 |  |  | 09/30/2025 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402412 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that no pulses were present in the impella leg following impella cp implant. an external bypass procedure was discussed, but the physicians opted to try to regain pulses with lower pressors, a warming blanket, and the application of nitro paste to the foot. the decision was made to withdraw care. it is unknown if the actions taken resolved the condition as the pump was subsequently explanted the following day. the patient was expired at time of explant. upon review by abiomed's medical safety officer, there is no association between impella use and the patient's outcome. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25399 | 1220648-2024-12385 | 19521076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521076&pc=OZD | 04/27/2024 | 07/25/2024 |  | 06/12/2024 |  |  | 05/29/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025470046 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a 61-year-old male patient with post cardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported they believed the patient died from a blood clot that had formed during the 5.5 insertion, followed by multiple organ failure. additional details pertaining to that allegation were not provided. | [5 empty] | The investigation for thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12385: f6/f8-should have been left blank g1 reporting contact fax number missing g2 report source: foreign missing | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25400 | 1220648-2024-12386 | 19521127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521127&pc=OZD | 06/16/2021 | 03/17/2025 |  | 06/12/2024 |  |  | 06/16/2021 | 06/12/2024 |  |  | 11/30/2022 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021014763 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the impella was placed easily via the right femoral artery. single access was attempted but the catheter could not be advanced. second access was obtained via left femoral artery and percutaneous coronary intervention of the left main, left anterior descending performed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25401 | 1220648-2024-12390 | 19521303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521303&pc=OZD | 05/29/2024 | 08/01/2024 |  | 06/12/2024 |  |  | 05/29/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461168 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support as escalation of care from an impella cp device. the pump was unable to pass through axillary artery. surgeons felt vessel was borderline in size. angiogram showed calcium and tortuosity. a 7mm balloon was inserted and the pump implant was reattempted to no avail. the procedure was aborted. | [5 empty] | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the patient had calcium and tortuosity and was unable to cross even after multiple attempts. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12390 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12390 was submitted. h.6 code 2199 was entered incorrectly on the previously submitted manufacturer device report 1220648-2024-12390. a new code has been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report 1220648-2024-12390 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
25402 | 1220648-2024-12391 | 19521350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521350&pc=OZD | 05/13/2024 | 07/21/2024 |  | 06/12/2024 |  |  | 05/13/2024 | 06/12/2024 |  |  | 09/30/2025 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444881 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was placed but the waveforms and low flows were not optimal. after 43 minutes of support, the pump was explanted as they felt the low flows were due to a clot. the pump was explanted, and no thrombosis was seen. the second pump was placed but the family then made the patient's care as comfort only and the patient expired. per abiomed's medical safety officer review, there was no association between impella use and outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the pump was not returned for investigation. data logs were returned and investigated. flows at p9 were seen deteriorating followed by suction and purge pressure blocked alarm. the purge alarm resolved immediately. placement signal trend was declining followed by suction and low flows with p-level drop. per the clinical information provided, the root cause of the low flow issue was most likely patient condition related. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12391 and should not have been. h.6 code 4629 was inadvertently omitted from the previous manufacturer device report number 1220648-2024-12391 and has been added.
25403 | 1220648-2024-12392 | 19521431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521431&pc=OZD | 05/30/2024 | 06/19/2024 |  | 06/12/2024 |  |  | 05/30/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a broken purge disk. there was no active patient use at the time. the team will follow the ifu and use a backup console for patients. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added- d8-checked box for device returned to mfr. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
25404 | 1220648-2024-12393 | 19521510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521510&pc=OZD | 07/22/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/22/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059836 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported that a 68 year old male patient with occluded right coronary artery in catheterization lab was suddenly unstable during intervention. patient placed on impella cp for hemodynamic support. while on support, it was reported that there was bleeding from groin, mouth and nose. also left breast hematoma and suction alarms were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25405 | 1220648-2024-12394 | 19521516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521516&pc=OZD | 06/02/2024 | 07/15/2024 |  | 06/12/2024 |  |  | 06/02/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2024 |  |  |  |  | The user facility reported a 84-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient developed diminished flow in their lower right extremity one day into impella cp support. the pump was removed and a thrombotic occlusion was discovered in the right external iliac artery (reia). a percutaneous revascularization of the reia was performed with a return to flow noted at the end of the procedure. | [5 empty] | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25406 | 1220648-2024-12395 | 19521664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521664&pc=OZD | 06/03/2024 | 07/16/2024 |  | 06/12/2024 |  |  | 06/03/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025459713 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced limb ischemia on the impella implant lower extremity. distal perfusion catheter placed by ic but unsuccessful in restoring flow to lle, impella cp weaned and removed. vascular surgery was consulted to take patient to operating room and have vessels repaired to restore flow to extremity, which was successful. the impella was weaned and removed. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿
25407 | 1220648-2024-12396 | 19521712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19521712&pc=OZD | 07/26/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/26/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089371 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/26/2021 |  |  |  |  | The complainant reported the a 53 year old female patient was implanted with an impella cp for mechanical circulatory support. while on support, loss of blood flow to leg, and suction alarms were noted. the patient was administered albumin for suction alarms. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25408 | 1220648-2024-12397 | 19522042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522042&pc=OZD | 06/04/2024 | 08/06/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466784 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | Additional information received in medwatch report mw5157023, that the patient expired after an hour and 15 minutes after the pump issue. it is unknown if the cause of death is related to the pump. | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for an elderly patient with mvd cad (mitral valve disease, coronary artery disease) and in need of pci (percutaneous coronary intervention¿). the pump was placed for support and the pump stopped after just 20 minutes of support. the stopped pump was replaced, and support initiated on the 2nd pump. |  |  |  |  | The investigation of pump stop has been completed. data logs show that 12 minutes into the case there is a high motor current pump stop alarm 13 following an increase in motor current. purge pressure begins to rise immediately after this and triggers high pressure purge (hpp) and purge system blocked alarms 2 minutes later. the pump is unable to be restarted and the hpp does not resolve. upon investigation of the product there is biomaterial round the impeller and in the purge gap. the pump passed electrical testing. the pump stop alarm 13 was reproduced in the lab and the pump was unable to be run. hpp was also reproduced in the lab after purging for 10 plus minutes. no blood ingress was found. the root cause of the pump stop was most likely biomaterial. the following have been updated on manufacturer device report 1220648-2024-12397: b1 updated to adverse event. b2 outcomes attributed to adverse event updated to death (date of death is unknown). b5 updated with additional information. b7 patient history updated. d9 date device returned to manufacturer added. e4 updated to yes. h1 type of reportable event updated to death. h6 new codes have been added to health effect - clinical code and health effect - impact code. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12397: f6/f8-should have been left blank | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
25409 | 1220648-2024-12398 | 19522059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522059&pc=OZD | 07/27/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/27/2021 | 06/12/2024 |  |  | 06/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093980 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/27/2021 |  |  |  |  | The complainant reported that a 76 year old female patient on impella cp returned from having mitral repair in the operating room. it was reported that the patient received fluid for low flows and was having urine output monitored closely. in addition, hemolysis was noted. it was reported that the device was being removed due to cardiac function and possible limb ischemia distal to impella. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25410 | 1220648-2024-12399 | 19522127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522127&pc=OZD | 06/04/2024 | 08/07/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  | 10/31/2025 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024425783 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 92 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported a 92-year-old male having a aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a leak from around the luer connector. it was further reported that purge pressure increased, and purge flow rate decreased. purge cassette was replaced. there was no patient harm. | [5 empty] | The investigation for the high purge pressure and the purge leak has been completed. clinical details states there was increased purge pressure and decreased purge flow rate, so the purge set was replaced. according to clinical notes, there were no purge line kinks observed, and the purge set replacement resolved the issue. data logs indicate gradual increase in purge pressure with little saturated flows seen at the end of case. the root cause of the high purge pressure was most likely biomaterial due to the gradual increase in purge pressure. the root cause of the purge leak was not determined since the product was not returned. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25411 | 1220648-2024-12400 | 19522173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522173&pc=OZD | 06/04/2024 | 07/17/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351112 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) failed the self-test. the aic was removed from service. there was no patient involvement. | [5 empty] | The investigation for the reported console (aic) system self check issue has been completed. the console was returned for evaluation. during the initial boot, the console rebooted due to the communication issue with powermod_1, resulting in a ¿system self-check failed screen.¿ visually confirmed the pbm contamination. data logs were reviewed finding that during the initial bootup, the console recorded a communication issue with powermod_1. due to the communication issue, the console rebooted automatically. again, during the reboot, there was a communication issue with powermod_1. then the console displayed the ¿system self-check failed¿ screen and recorded communication issues with powermod_1 persistently. this confirms the reported failure mode. the root cause of the system self-check failure was pbm contamination. added- d9 device return date f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25412 | 1220648-2024-12401 | 19522266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522266&pc=OZD | 06/04/2024 | 07/25/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025451697 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported a 67-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma during impella cp implant. as a result of the noted hematoma, the doctor made the decision to remove the pump and a covered stent was placed at the site. hemostasis was achieved and the hematoma resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding/hematoma has been completed. there was no product or data logs returned to further investigate the case. the root cause of the access site bleeding is most likely to be a use issue based on all the clinical notes. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12401.
25413 | 1220648-2024-12402 | 19522354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522354&pc=OZD | 06/04/2024 | 07/16/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  | 06/30/2025 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024383604 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a numb extremity however, their pulse was able to be dopplered. the pump was explanted and a small thrombus was noted. a thrombectomy was performed to the common iliac. | [5 empty] | The investigation into the patient's thrombus and limb ischemia has been completed. the abiomed products were not returned and data logs were not provided. as all the necessary information was not provided, the root cause of the thrombus and limb ischemia was not determined. a5 selected non hispanic/latino (inadvertently omitted on the initial mdr) f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25414 | 1220648-2024-12403 | 19522693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522693&pc=OZD | 07/27/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/27/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/27/2021 |  |  |  |  | The complainant reported that a 58 year old male patient with occluded right coronary artery, and ramus with severe stenosis in the left anterior descending (lad) and left main coronary artery. impella was placed pre-percutaneous coronary intervention (pci) via the right femoral artery without issue. intervention of the right coronary artery completed first, which resulted in the patient going into ventricular fibrillation and requiring intubation and cardiopulmonary resuscitation. return of spontaneous circulation was achieved, and further pci of the lad, and ri performed. in addition, an external fem/fem was done to help limb perfusion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25415 | 1220648-2024-12404 | 19522719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522719&pc=OZD | 06/04/2024 | 08/08/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025459940 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Restricted Flow rate | Increase in Pressure | No Flow |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp smart assist support and on the third day of support, there was a decrease in purge flow rate. there was eventually a purge blockage alarm and the flow rate was below 1ml/hr. since the bleeding occurred on the day of hm introduction, there may have been a problem with the technique used during graft construction. in addition, the bleeding occurred at night, and the doctor on duty was not the patient's regular doctor, so the discovery of the cause and treatment were delayed. due to the increase in bleeding caused by the inability to respond early, it became necessary to stop using heparin in the purging fluid. it is believed that this treatment also led to the purging blockage. the impella was removed. | [5 empty] | The investigation of access site bleeding and high purge pressure has been completed. field data logs were investigated, and the high purge pressure alarm was confirmed. in investigating returned product, initial visual inspection yielded no abnormalities. the pump was primed and soaked, and purge pressure was trending upwards. once the pump was started, the purge pressure quickly dropped to normal levels. subsequent visual inspection revealed biomaterial near the impeller. with the biomaterial out of the cannula, the pump ran normally in a bucket. the biomaterial likely built up over time rather than being ingested, considering there were no motor current spikes. the root cause of the high purge pressure was determined to be biomaterial. the root cause of access site bleeding was not determined since not enough clinical detail was provided and the introducer was not returned. additional treatment information was inadvertently not included in manufacturer device report 1220648-2024-12404: the bleeding caused a decrease in extracorporeal membrane oxygenation (ecmo) flow, which required the administration of 14 units of red blood cells. the patient's blood pressure dropped slightly, but both the ecmo and impella began to function without problems after a blood transfusion. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12404: d6a and d6b incorrect. f6/f8-should have been left blank. h1 type of reportable event. h.6 code 4627, 1248 and 2991 were reported incorrectly. new codes were added to health effect - impact code and medical device problem code. | The impella device was received from the customer on (b)(6)2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25416 | 1220648-2024-12405 | 19522886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522886&pc=OZD | 07/29/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 07/29/2021 | 06/12/2024 |  |  | 05/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088997 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/29/2021 |  |  |  |  | The complainant reported that a 79 year old male patient decompensated after undergoing coronary artery bypass grafting, so the decision to escalate to the impella cp was made. during support, it was reported that heparin was stopped due to bleeding issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." this report is being filed as part of a retrospective review of historical records.
25417 | 1220648-2024-12406 | 19522941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19522941&pc=OZD | 06/04/2024 | 07/19/2024 |  | 06/12/2024 |  |  | 06/04/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463832 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The complainant reported a 52-year-old male patient with post-cardiotomy cardiogenic shock / low cardiac output syndrome had an attempted implant with an impella 5.5 for mechanical circulatory support. the complainant reported the implant was unsuccessful ¿ medical staff noted that a transesophageal echocardiography showed the mitral struts on the new valve protruded into the left ventricular outflow tract obstruction close to the aortic valve. medical staff then aborted the implant. | [5 empty] | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump could not be delivered due interaction with mitral struts on new valve. e1 initial reporter name and phone number missing. f6/f8-should have been left blank in the initial report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25418 | 1220648-2024-12407 | 19523029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19523029&pc=OZD | 06/04/2024 | 07/30/2024 |  | 06/12/2024 |  |  | 06/05/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025489174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2024 |  |  |  |  | The user facility reported a 75-year-old male noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was kinked. the impella was removed and replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation product damage (cannula kink) has been completed. the impella device was not returned for evaluation. based on clinical description, the root cause of the cannula kink was most likely use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12407: f6/f8-should have been left blank. g1 manufacturing site address.
25419 | 1220648-2024-12409 | 19523716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19523716&pc=OZD | 08/15/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 08/15/2021 | 06/12/2024 |  |  | 06/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102832 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2021 |  |  |  |  | The complainant reported a 76 year old male patient was placed on impella cp support. it was reported that while in the intensive care unit (icu), the patient developed a slight tract ooze, and the icu intensivist adjusted the angle and inserted the sheath further into the groin. it was feared at this point that the vessel was sheared, because a giant hematoma formed. after pressure was held and the blood expelled, any time pressure was lifted. the patient was bleeding profusely from his groin. patient transferred for emergent vascular surgery, possible impella removal and new placement, and possible rp placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). limb ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25420 | 1220648-2024-12410 | 19523900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19523900&pc=OZD | 07/18/2021 | 03/21/2025 |  | 06/12/2024 |  |  | 07/18/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080944 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a subdural hematoma but recovered with no deficit and with no indication of additional intervention related thereto. the patient remained on impella support and was successfully weaned. | [5 empty] | The cause of the stroke/cva was not determined since product was not returned and all the necessary information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
25421 | 1220648-2024-12411 | 19523995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19523995&pc=OZD | 08/15/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 08/15/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022058252 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2021 |  |  |  |  | The complainant reported a 26 year old male patient was placed on impella cp support. it was reported that impella was inserted and removed due to leg ischemic. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. explant date is unknown. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25422 | 1220648-2024-12408 | 19524016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524016&pc=OZD | 08/02/2021 | 03/25/2025 |  | 06/12/2024 |  |  | 08/02/2021 | 06/12/2024 |  |  | 06/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP - EU |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2022092420 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, that patient was in atrial fibrillation, was still tachycardic the following day, and normal rate the day after. additionally, it was reported that there was onset of renal failure. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25423 | 1220648-2024-12412 | 19524119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524119&pc=OZD | 08/18/2021 | 06/12/2024 |  | 06/12/2024 |  |  | 08/18/2021 | 06/12/2024 |  |  | 06/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095092 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/18/2021 |  |  |  |  | The complainant reported that a 72 year old female patient was placed on impella cp support. it was reported that doctor forgot to preclose so decided to use balloon tamponade and post close with 2 perclose and 1 angioseal. the patient became hypotensive, hemoglobin dropped 2 points. a perforation in iliac artery where balloon tamponade was done was found. viabahn stent placed with good result. levo weaned off and blood being given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
25424 | 1220648-2024-12413 | 19524160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524160&pc=OZD | 07/11/2021 | 03/25/2025 |  | 06/12/2024 |  |  | 07/11/2021 | 06/12/2024 |  |  | 02/28/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047028 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2021 |  |  |  |  | The user facility reported a 51-year-old male on impella cp for mechanical circulatory support. during support, bloody urine noted to foley catheter, partial thromboplastin time and chemistry samples showed evidence of hemolysis. hemoglobin remained low despite 2 units of red blood cells transfused. inlet measured at 2.5 cm from annulus and pigtail not well positioned in midventricular space. the power level of the impella was lowered, and a decision was made to perform new trans-thoracic echocardiogram and reposition the impella cp at bedside. the device was pulled away from sub annular structures and correctly positioned across the valve with a measurement of 4.04 cm. the impella p-level was then resumed to previous therapy. the patient remained on impella support and was successfully weaned. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
25425 | 1220648-2024-12414 | 19524262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524262&pc=OZD | 08/18/2021 | 03/25/2025 |  | 06/12/2024 |  |  | 08/18/2021 | 06/12/2024 |  |  | 04/30/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed a large hematoma at site and it was pressed out. hematoma resolved. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding, hematoma, impella cp with smart assist system. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25426 | 1220648-2024-12415 | 19524401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524401&pc=OZD | 07/12/2021 | 07/01/2025 |  | 06/12/2024 |  |  | 07/12/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2021 |  |  |  |  | The user facility reported a 45-year-old male on impella cp for mechanical circulatory support. during support it was reported that the echocardiogram showed a hypertrophic left ventricle (lv) that was under filled. the impella cp device was 2 cm in the lv, and the pigtail hindered the catheter from being advanced. the impella cp device was working and the patient was not hemolyzing. while removing the sheath the impella cp device was pulled across the aortic valve and would not re-cross. the impella cp device was wired out and replaced with a new impella cp. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-12415. g1 reporting contact email revised on manufacturer device report 1220648-2024-12415 in accordance with updated procedures.
25427 | 1220648-2024-12416 | 19524665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524665&pc=OZD | 07/13/2021 | 03/25/2025 |  | 06/12/2024 |  |  | 07/13/2021 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient, of unknown race and ethnicity, noted as high-risk percutaneous coronary intervention was attempted to be implanted with an unknown impella device for mechanical circulatory support. the device could not successfully be inserted, and the implant was aborted. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | D4: device lot number, catalog number, serial number, lot number, expiration date, and udi are unknown. h4: date of manufacturer is unknown the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records.
25428 | 1220648-2024-12417 | 19524784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524784&pc=OZD | 07/16/2021 | 10/06/2025 |  | 06/12/2024 |  |  | 07/16/2021 | 06/12/2024 |  |  | 03/31/2023 |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/16/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced red urine, but no intervention was performed. after the patient was weaned and the impella was explanted, the patient's right leg, where the impella had accessed the body, became ischemic. the patient was brought to the or to investigate and treat possible clots that had developed in the right femoral artery. | Complainant reported the patient was on impella cp support due to high-risk percutaneous coronary intervention. while on support a large hematoma, loss of pulse, red urine, and ischemic leg was noted. the large hematoma was noted to be non-impella related; therefore not reportable. it was also noted that there was no bleeding at the site after pump removal. patient was taken back to the operating room to explore the groin for clot. reportable due to limb ischemia, thrombus, and hemolysis. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25429 | 1220648-2024-12418 | 19524935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524935&pc=OZD | 06/05/2024 | 07/23/2024 |  | 06/12/2024 |  |  | 06/05/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485631 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella 5.5 for escalation from an impella cp to a patient for mechanical circulatory support. the impella 5.5 pump went down the descending aorta and pulled the wire out of the left ventricle. the case was aborted and eventually the decision was made to monitor the patient off impella support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance." | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, as difficulty gaining access to lv with multiple wires and catheters. everything was going down the descending aorta per the clinical description provided.
25430 | 1220648-2024-12419 | 19524960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524960&pc=OZD | 06/06/2024 | 08/16/2024 |  | 06/12/2024 |  |  | 06/06/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481909 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Improper Flow or Infusion | Infusion or Flow Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The user facility reported a 21-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a 5.5 pump to support a patient in need of care escalation from the cp while bridging to transplant. the 5.5 pump had high purge pressures and saturated flows. the team troubleshot by tpa (tissue plasminogen activator) for the high pp and the saturated flows are being monitored. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | D.9 updated as the pump and purge cassette were returned for investigation. the investigation has been completed. the pump was returned for evaluation. upon visual inspection, a kink around the 35 centimeter (cm) mark on the catheter and biomaterial was present in outflow cage covering impeller blade and purge gap. pump stop reproduction testing was performed and blood ingress was present in purge line up to around the 50 cm mark. the pump was plugged into automated impella controller and high purge pressure was reproduced. the pump was cut just below the motor housing during testing and high purge pressure was resolved. additionally, a kink deformation in the purge line was found around the 35 cm mark. data logs were reviewed and a sharp increase in purge pressure over less than a two minute period with a simultaneous decrease in purge flow. high purge pressure alarms were triggered and one hour later, an gradual increase in motor current was observed with pump flows becoming saturated. a pump stop then occurred two hours after high purge pressure alarms were triggered. clinical details noted high purge pressure alarms were triggered and the purge cassette was changed and tissue plasminogen activator was added to the purge in an attempt to resolve high purge pressure alarms without success. the pump had a pump failure and emergent explant was needed. the root cause of the pump stop was due to a kinked purge line around the 35 cm mark resulting in blood ingress in the purge line, buildup of biomaterial on the motor and purge gap causing saturated flows and eventually leading to a pump stop. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.3 us zip code revised as this information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12419. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12419 and should not have been. h.6 codes 1248 and 2954 were entered incorrectly on the initial manufacturer device report number 1220648-2024-12419. new codes were added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12419 in accordance with updated procedures.
25431 | 1220648-2024-12420 | 19524979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19524979&pc=OZD | 06/06/2024 | 08/13/2024 |  | 06/12/2024 |  |  | 06/06/2024 | 06/12/2024 |  |  |  |  | 06/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025467796 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2024 |  |  |  |  | The complainant reported a 54-year-old female patient with post-cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant reported the pump stopped unexpectedly during patient support. the pump was able to restart at support level p1, but subsequently stopped again after thirty minutes. as a result of the noted failure, medical staff decided to explant the pump. medical staff then achieve hemostasis with manual pressure and a femoral stop. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. the product and data logs were not returned for investigation. impella cp device was used in support until 24 days when the pump stop occurred. the impella cp optical device was not indicated for support more than eight days of design life period according to design trace matrix per instructions for use. the cause of the pump stop issue was not determined since the product and data logs were not returned for investigation. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12420 and should not have been. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-12420 in accordance with updated procedures.
25432 | 1220648-2024-12293 | 19508681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19508681&pc=OZD | 07/19/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 07/19/2021 | 06/11/2024 |  |  | 01/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022046001 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the patient was bleeding including a hematoma that required the patient to receive a blood transfusion. it was also noted that the patient experienced hemolysis from insertion of impella until it was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related events are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25433 | 1220648-2024-12294 | 19508720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19508720&pc=OZD | 07/14/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 07/14/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084576 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2021 |  |  |  |  | The complainant reported that the patient had been placed on impella cp support after presenting with non-st-elevation myocardial infarction (nstemi). upon insertion to the left ventricle, the patient went into ventricular tachycardia (vt) and started coding. cardiopulmonary resuscitation was performed on the patient and return of spontaneous circulation (rosc) was achieved. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25434 | 1220648-2024-12295 | 19508799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19508799&pc=OZD | 07/22/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 07/22/2021 | 06/11/2024 |  |  | 03/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071443 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a 60 year old male patient was placed on impella cp support. it was reported that there were efforts to reposition impella twice, but pigtail appears to be stuck in the papillary muscle. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
25435 | 1220648-2024-12296 | 19508994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19508994&pc=OZD | 07/21/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 07/21/2021 | 06/11/2024 |  |  | 10/31/2022 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021004219 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/23/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion, the physician noted the patient's vessel size was small for the impella introducer and was concerned with causing limb ischemia when the sheath was inserted. the physician was advised to perform a femoral to femoral bypass, if needed, should the sheath cause limb ischemia. access was obtained to the right groin and dilated up to 14 french oscor sheath, and distal flows to the lower right extremity were found to be totally occlusive. a femoral to femoral bypass was performed and the impella was successfully inserted. it was reported while on support and during weaning of the patient from impella, the patient experienced an episode of ventricular tachycardia and the patient received amiodarone. it was advised for a slow wean process and the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25436 | 1220648-2024-12298 | 19509148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509148&pc=OZD | 06/11/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/11/2021 | 06/11/2024 |  |  | 02/28/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056482 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was shocked four times due to ventricular fibrillation (vf). additionally, the patient developed access site bleeding, a large hematoma and retroperitoneal bleeding. the patient was provided blood products. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25437 | 1220648-2024-12299 | 19509210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509210&pc=OZD | 06/14/2021 | 11/14/2024 |  | 06/11/2024 |  |  | 06/14/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, clot was seen throughout the vessels. percutaneous transluminal angioplasty and intravascular ultrasound were used to obtain optimal revascularization. angiogram showed a complete revascularization. impella was then weaned and explanted without hemodynamic compromise. patient's condition was stable post procedure. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12299. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
25438 | 1220648-2024-12300 | 19509222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509222&pc=OZD | 03/18/2024 | 06/11/2024 |  | 06/11/2024 |  |  | 03/18/2024 | 06/11/2024 |  |  | 11/30/2025 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024434128 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, there was high purge pressure. staff troubleshot by adding tissue plasminogen activator and replacing the purge cassette unsuccessfully. plan was to wean and remove anyway, purge issues expedited the wean. there was no patient harm reported. | [5 empty] | No product was returned for evaluation. data logs were reviewed and the log at time of high purge pressure alarms show that an increase in purge pressure over a period of approximately 5 minutes. no deviations in motor current or motor speed at time of purge increase. approximately 1 minute before purge pressure begins to increase and purge flow starts to drop, a disruption in placement signal is observed, motor current and pump flow become momentarily saturated at that time. the log at the end of the case shows high purge pressure beginning to resolve as pump is weaned down approximately 10 hours after high purge pressure alarms are first triggered. clinical details noted that a cassette change was attempted, and tissue plasminogen activator (tpa) was also administered, both did not resolve high purge pressure alarms. the root cause of the high purge pressure is most likely due to a kink in the purge line.
25439 | 1220648-2024-12301 | 19509260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509260&pc=OZD | 06/30/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/30/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080849 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2020 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, slight bleeding was observed at the impella access site, and the medical staff was concerned that the growing bleed could be a perforated artery. therefore, the patient was weaned, and the impella was explanted to address the site bleed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25440 | 1220648-2024-12302 | 19509280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509280&pc=OZD | 06/16/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/16/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069888 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/16/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported some oozing at the impella site. percloses were tightened to reduce oozing and mattress sutures were placed as well. additionally, it was reported that there was a positioning issue; the impella was placed too deep. the device was repositioned. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25441 | 1220648-2024-12303 | 19509365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509365&pc=OZD | 05/14/2024 | 07/25/2024 |  | 06/11/2024 |  |  | 05/14/2024 | 06/11/2024 |  |  |  |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488961 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a male unknown age in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump immediately after implantation showed "purge-pressure low" alarm. the alarm appeared again and again until explantation. the problem could not be solved and unfortunately, the pump was not removed, even after a rep recommendation. the patient expired while on support. | [5 empty] | The investigation for the purge leak has been completed. low purge pressure alarms reported upon implant. no product was returned. the data logs show low purge pressure alarms characteristic of a purge leak and low pump flow for matching p-levels. there were also "impella position unknown" alarms, suction alarms and the placement signal was trending downwards. the root cause of the purge leak - pump was not determined as no product was returned and insufficient clinical information was provided. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi states the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ b1-selected adverse event b2-selected death and added date of death d4-corrected model number f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address h1-type of reportable event changed to death | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
25442 | 1220648-2024-12304 | 19509398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509398&pc=OZD | 05/13/2024 | 07/18/2024 |  | 06/11/2024 |  |  | 05/13/2024 | 06/11/2024 |  |  |  |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493664 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The patient experienced ischemic signs, however, pulses were still checkable. the patient's abdomen was opened due ischemic abdomen, and patient expired in the night. the patient expired while on support and per the medical safety officer not enough information to exclude use of the device from the outcome. | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support for 47-year-old male patient, the pump stopped, and the automated impella controller (aic) displayed 'impella stop"/"motor current high" and "impella stop/"retrograde flow" alarms. "restarting the pump was not successful. therefore, the pump was explanted. there was no reported patient harm. |  |  |  |  | Pump was returned for investigation cut from the catheter. there was heavy biomaterial found in the out flow cage, around the impeller. the purge gap was noticed to be larger than usual. the pump was ran for around 232 hours when the failure was recorded. this is over the design life of the impella cp pump ie. 192 hours or 8 days. pump use exceeded the design life per design trace matrix. therefore, the root cause of the pump stop was bearing wear beyond indication for use. the following information that was inadvertently omitted from manufacturer device report 1220648-2024-12304: b1 adverse event. b2 death; date of death. b5 revised. h1 type of report death. h6 health impact code 1802. corrections: d4 model number. f6/f8-should have been left blank in the initial report. g1 manufacturing site name and address. | The impella device was received from the customer on (b)(6)2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
25443 | 1220648-2024-12305 | 19509422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509422&pc=OZD | 06/16/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/16/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079228 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/16/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported manual pressure was applied due to a slow, slight ooze. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25444 | 1220648-2024-12306 | 19509577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509577&pc=OZD | 07/23/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 07/23/2021 | 06/11/2024 |  |  | 02/28/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | Hematuria | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced hematuria and blood samples sent to the laboratory for testing were hemolyzed. it was recommended to the medical team to add volume to the patient, lower p level on the pump and reposition the device. the patient received fluids, the physician advanced the pump 1.5cm into the ventricle, and adjusted the p level down to p4. the patient continued to have a low urinary output and was placed on continuous renal replacement therapy. it was also reported during impella support the patient became hyperdynamic with a septic profile and the decision was made to wean and explant the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related events are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25445 | 1220648-2024-12307 | 19509596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509596&pc=OZD | 06/17/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/17/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080391 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma at the impella access site. the patient was provided blood products and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25446 | 1220648-2024-12308 | 19509638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509638&pc=OZD | 06/17/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/17/2021 | 06/11/2024 |  |  | 01/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022041009 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that 4 french oscar sheath buckled and had to be removed. the patient experienced bleeding at the impella access site. a 16 french cook sheath was placed and the impella placed through sheath. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25447 | 1220648-2024-12309 | 19509694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509694&pc=OZD | 06/17/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/17/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a high suspicion that the patient had retroperitoneal bleed. the decision was made to remove the impella due to the risk for bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25448 | 1220648-2024-12310 | 19509696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509696&pc=OZD | 05/18/2021 | 02/21/2025 |  | 06/11/2024 |  |  | 05/18/2021 | 06/11/2024 |  |  | 01/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022040684 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was noted to have ischemia on the impella leg. the limb ischemia was resolved after the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia - reversible has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25449 | 1220648-2024-12311 | 19509753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509753&pc=OZD | 06/17/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/17/2021 | 06/11/2024 |  |  | 01/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022042446 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/17/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the perclosures did not properly provide hemostasis. manual pressure was held. additional perclose was applied and 8 fr angioseal was utilized. hemostasis was obtained. a hematoma was observed and evacuated at that time. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25450 | 1220648-2024-12312 | 19509788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509788&pc=OZD | 06/04/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/04/2021 | 06/11/2024 |  |  | 02/28/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. access was gained for impella placement through the right femoral artery utilizing a fourteen french peel away sheath. single access technique was attempted with a seven french destination sheath however, the medical professional could not advance the sheath to a comfortable position. the sheath was removed, and re-access achieved at the ¿two o¿clock¿ position using a fast cash sheath. the second attempt at implant was effective with no additional complications. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25451 | 1220648-2024-12313 | 19509886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509886&pc=OZD | 06/18/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/18/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/18/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a long run of ventricular tachycardia. the patient coughed out and then experienced atrial fibrillation rapid ventricular response. the patient recovered well. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25452 | 1220648-2024-12314 | 19509894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509894&pc=OZD | 07/16/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 07/16/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088996 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after several days on impella support, the physician was unable to cross the aortic valve despite multiple attempts. the device was replaced with another impella. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25453 | 1220648-2024-12315 | 19509908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509908&pc=OZD | 06/04/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/04/2021 | 06/11/2024 |  |  | 03/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022062661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a good amount of oozing around groin, and the heparin was discontinued to get bleeding under control. heparin was re-administered. impella flows were reduced, and aortic placement signal was not accurate to invasive pressures. impella was repositioned in the ventricle and still same outcome. decision to replace the impella catheter due to decreased flows and continued suction events after repositioning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25454 | 1220648-2024-12316 | 19509941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509941&pc=OZD | 06/18/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/18/2021 | 06/11/2024 |  |  | 03/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022061822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/18/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was recovering from sepsis and the fever was down. medical or surgical intervention was not noted. the patient continued on support until the device was weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25455 | 1220648-2024-12317 | 19509947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509947&pc=OZD | 05/14/2024 | 06/11/2024 |  | 06/11/2024 |  |  | 05/14/2024 | 06/11/2024 |  |  |  |  | 06/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1576134 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The automated impella controller (aic) was returned for evaluation and inspection of the aic revealed a malfunction of the purge disk detect flag. | [5 empty] | The automated impella controller was returned as part of further investigation for a non-reportable complaint which indicated placement signal unreliable alarm occurred while connected to an impella 5.5 pump implanted in a patient for bridge to transplant. the console and data logs received for evaluation/analysis indicated the following: console logs from the day of the user facility complaint are consistent and show ¿placement signal not reliable¿ alarms. alarms occurring due to placement signal out of range were associated with gradual decrease (drift) of placement signal and are unrelated to the console. inspection of the console, however, revealed a new complaint issue/malfunction of the purge disk detect flag due to purge fluid ingress, and compromised backstrap cable explaining frequent reboots due to intermittent power. there was also significant contamination on the optical pump connector that could not be fully cleaned. the cause of the aic issue was contamination.
25456 | 1220648-2024-12318 | 19509993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19509993&pc=OZD | 06/19/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/19/2021 | 06/11/2024 |  |  | 03/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063212 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/19/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced groin access site bleeding. bleeding continued despite use of femstop. the site was redressed and monitored. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25457 | 1220648-2024-12319 | 19510027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510027&pc=OZD | 06/06/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/06/2021 | 06/11/2024 |  |  | 12/31/2022 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022033255 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, limb ischemia was observed at the site where the impella cp was inserted. as a result, the decision was made to remove the device, which was successfully accomplished. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: limb ischemia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25458 | 1220648-2024-12320 | 19510106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510106&pc=OZD | 06/20/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/20/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075338 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/20/2021 |  |  |  |  | The complainant reported a patient, race unknown, in non-st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient was experiencing vascular bleeding at the impella access site. the patient was taken to the operating room for a surgical repair of the femoral artery. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25459 | 1220648-2024-12321 | 19510137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510137&pc=OZD | 06/19/2021 | 11/14/2024 |  | 06/11/2024 |  |  | 06/19/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079713 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/19/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. thrombus was suspected. heparin was on the low side and there was no peripheral heparin hung and minimal coming from purge. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. added-h6 component code 925 corrections to the initial mfr report: f6/f8 should have been left blank.
25460 | 1220648-2024-12322 | 19510166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510166&pc=OZD | 06/06/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/06/2021 | 06/11/2024 |  |  | 03/31/2021 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068969 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/06/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, when the peel-away sheath was removed, and repo sheath advanced a hematoma was noted at site. multiple maneuvers and attempts were made to slow hematoma growth without success, the impella was removed after ecmo was placed | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25461 | 1220648-2024-12323 | 19510193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510193&pc=OZD | 06/20/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/20/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069887 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/20/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site oozing and a hematoma developed after removal of the peel away sheath. the repositioning sheath was successfully inserted and a femstop placed. after a couple hours, bleeding started again and the hematoma size increased. manual pressure was held until the physician determined to remove the impella and continue holding manual pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25462 | 1220648-2024-12324 | 19510210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510210&pc=OZD | 07/17/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 07/17/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078278 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2021 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was then reported that the patient developed a large hematoma approximately four (4) hours post-implant that extended into the scrotum and penis. the device was then explanted with primary closure in surgery. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25463 | 1220648-2024-12325 | 19510248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510248&pc=PYX | 05/10/2024 | 07/26/2024 |  | 06/11/2024 |  |  | 05/10/2024 | 06/11/2024 |  |  | 11/25/2025 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024435267 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. information received on 05/13/2024 via abiomed¿s long-term outcome and quality indicator impella registry regarding the patient that endured pulmonary hemorrhage with a "definite" relationship to the impella device used: the patient was intubated, heparin was reversed w/protamine and pulmonary service was contacted for emergent assistance. patient also received aspirin and ticagrelor. urgent bronchoscopy performed. clot occluding the lower left lobe (lll) noted, no active hemorrhage. nimbex considered to ensure clot remains in lll bronchus and doesn't get coughed out. maintain ventilation and oxygenation to achieve good hemostasis. patient did not tolerate. bronchoscopy on (b)(6) shows no clot in lll. cta showed lll opacities correlated to pulmonary hemorrhage which was caused by device placement during procedure.¿ the patient recovered with sequelae. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. rp flex with smart assist system with automated impella controller & rp with smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ | Additional information was received indicating this is a duplicate report. the investigation for this issue will be reported in mfr 1220648-2024-11189.
25464 | 1220648-2024-12326 | 19510251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510251&pc=OZD | 07/17/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 07/17/2021 | 06/11/2024 |  |  | 02/28/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055673 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2021 |  |  |  |  | The complainant reported a patient with aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hematuria and thrombocytopenia. as a result, heparin was put on hold. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25465 | 1220648-2024-12327 | 19510267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510267&pc=OZD | 06/07/2021 | 03/21/2025 |  | 06/11/2024 |  |  | 06/07/2021 | 06/11/2024 |  |  | 01/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022040304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the urinary output (uop) was positive for hematuria. it is unknown how hemostasis was achieved regarding hemolysis. weaning was successful and the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25466 | 1220648-2024-12328 | 19510285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510285&pc=OZD | 06/21/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/21/2021 | 06/11/2024 |  |  | 12/31/2022 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022034949 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2021 |  |  |  |  | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding when the long peel out sheath was being peeled out. the patient was given a blood transfusion with an unknown amount of packed red blood cells. it was reported the bleeding at the insertion site subsided and hemodynamics were stabilized. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25467 | 1220648-2024-12329 | 19510368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510368&pc=OZD | 07/18/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/18/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089663 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pulse was not palpable in the distal foot. additionally, it was reported that the patient had groin bleeding and small hematoma at insertion site. decision was made to remove impella at bedside. quick wean performed and impella was removed. hemodynamics remained stable and hemostasis was achieved. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25468 | 1220648-2024-12330 | 19510588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510588&pc=OZD | 06/07/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/07/2021 | 06/11/2024 |  |  | 02/28/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050416 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while attempting to place the patient on impella cp support, there was difficulty in sheath insertion during cardiopulmonary resuscitation. as a result, impella implant was aborted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
25469 | 1220648-2024-12331 | 19510727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510727&pc=OZD | 06/09/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/09/2021 | 06/11/2024 |  |  | 03/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported the site was closed with two perclose sutures due to access site bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25470 | 1220648-2024-12332 | 19510852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510852&pc=OZD | 06/09/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/09/2021 | 06/11/2024 |  |  | 02/28/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056178 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had severe hemolysis in the intensive care unit. the pump migrated and was tangled in the papillary muscle. the physician tried to reposition the pump but lost positioning in the left ventricle. multiple attempts were made to properly position the impella cp but ultimately failed due to unique patient anatomy. the physician elected to remove the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25471 | 1220648-2024-12333 | 19510900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510900&pc=OZD | 07/18/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/18/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were concerns about distal ischemia due to ecmo/impella on the same side. as a result, impella was explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25472 | 1220648-2024-12334 | 19510958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19510958&pc=OZD | 07/18/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/18/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091817 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2021 |  |  |  |  | The complainant reported a patient noted as high-risk surgery was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was having ectopy and the doctor decided to remove impella since he could not get optimal position for resolution. the doctor tired several times with echo guidance and ended up pulling back in the end. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25473 | 1220648-2024-12335 | 19511074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19511074&pc=OZD | 06/11/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/11/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022074873 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed slight oozing at the impella cp access site. manual pressure was held for five minutes and the previously deployed perclose sutures were tightened without issue. it was noted that patient had a right bundle branch block prior to start of the case. the physician pulled the impella back closer to aortic valve and complete heart block resolved temporarily but patient became bradycardic. native rhythm did not return. the physician was positive that the patient rhythm would return if given enough time and patient was very stable on temporary pacer, so decision was made to put temporary pacer in place and remove impella. impella was removed without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25474 | 1220648-2024-12336 | 19511198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19511198&pc=OZD | 06/07/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/07/2021 | 06/11/2024 |  |  | 12/31/2022 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022030263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that there was no flow dismally in the left leg. the physician decided to perform external femoral-femoral bypass procedure. additionally, the patient also experienced hemolysis. it is not known how hemostasis was achieved. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related events are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25475 | 1220648-2024-12337 | 19511352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19511352&pc=OZD | 06/11/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/11/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072642 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding at the impella access site was noted. there was no purge or systematic heparin. additionally, the patient developed thrombus. a penumbra catheter was used to remove the thrombus and reestablish left anterior descending artery flow. the patient¿s condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related events are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25476 | 1220648-2024-12338 | 19511539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19511539&pc=OZD | 06/11/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/11/2021 | 06/11/2024 |  |  | 03/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067808 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during explanation, the preclose suture had failed. additional perclose was deployed, and manual pressure was held over groin. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25477 | 1220648-2024-12339 | 19512203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19512203&pc=OZD | 06/22/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 06/22/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2021 |  |  |  |  | The complainant reported a patient, ethnicity unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the doctor wanted to leave the impella implanted. however, when the peel away was taken out and repo sheath inserted the patient developed a hematoma. the doctor decided to explant the impella as a result. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25478 | 1220648-2024-12340 | 19512611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19512611&pc=OZD | 04/03/2022 | 03/20/2025 |  | 06/11/2024 |  |  | 04/03/2022 | 06/11/2024 |  |  | 09/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022142120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was found that the device contained a clot, and the pump level could not rise above setting p3. the impella was removed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25479 | 1220648-2024-12341 | 19512808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19512808&pc=OZD | 07/25/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/25/2021 | 06/11/2024 |  |  | 06/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093373 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2021 |  |  |  |  | The user facility reported a 60-year-old male patient implanted with an impella cp for mechanical circulatory support. while on support hemolysis, a clot was noted in the proximal left anterior descending artery. in addition, there were spikes on the placement signal, and positioning issues were noted. the device was repositioned; however, hemolysis did not resolve. the patient was also placed on sedation. the physician decided to remove the impella cp device. manual compression was held until hemostasis was achieved. doppler pulses present post removal. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25480 | 1220648-2024-12342 | 19512861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19512861&pc=OZD | 07/25/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 07/25/2021 | 06/11/2024 |  |  | 06/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094466 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/25/2021 |  |  |  |  | The complainant reported the patient was implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock. while on support, the patient experienced limb ischemia with no reference to any intervention required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25481 | 1220648-2024-12343 | 19512953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19512953&pc=OZD | 06/24/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 06/24/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022077299 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. the reporter indicated the right leg was mottled with no palpable pulses. to resolve the limb ischemia, angioplasty of the right superficial femoral artery was performed. the patient remained on impella support and was sent to the intensive care unit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25482 | 1220648-2024-12344 | 19512978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19512978&pc=OZD | 07/25/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/25/2021 | 06/11/2024 |  |  | 04/30/2022 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1486612 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2021 |  |  |  |  | The user facility reported a 50-year-old male paient implanted with an impella cp for mechanical circulatory support. while on support, the patient developed oozing from the groin site and the central catheter (picc) site. the impella dressing was changed with angle matched and the site was still oozing slightly. it was recommended to decrease heparin concentration in purge to maintain act 160-180. the patient received blood products. hemostasis was achieved with a femstop and manual compression. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25483 | 1220648-2024-12345 | 19513068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19513068&pc=OZD | 07/26/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/26/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079799 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2021 |  |  |  |  | The user facility reported a 65-year-old male patient implanted with an impella cp for mechanical circulatory support. it was reported that while the patient was on support, the patient experienced a hematoma that required surgery to remove. the impella cp device was successfully weaned and explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25484 | 1220648-2024-12346 | 19513202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19513202&pc=OZD | 07/14/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/14/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078934 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed oozing at the access site. the dressing was changed to maintain entry angle and expose/tighten the touchy. in addition, heparin in the purge bag was lowered from 50u/ml to 25u/ml. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25485 | 1220648-2024-12347 | 19513617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19513617&pc=OZD | 07/14/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/15/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing and hematoma at the groin site. in response, heparin was withheld from the purge cassette, and manual pressure was held for 25 minutes, which resolved the issue. in addition, staff found that the patient's feet were cold and there was no doppler response for the left dorsalis pedis and posterior tibial. the limb ischemia was present until the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25486 | 1220648-2024-12348 | 19516693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19516693&pc=OZD | 07/18/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/18/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079716 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was found to be septic, which was treated with antibiotics. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the infection/sepsis could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
25487 | 1220648-2024-12349 | 19516703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19516703&pc=OZD | 07/19/2021 | 03/20/2025 |  | 06/11/2024 |  |  | 07/19/2021 | 06/11/2024 |  |  | 01/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022042903 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, pump moved into aorta. it was decided to explant the impella. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25488 | 1220648-2024-12350 | 19516727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19516727&pc=OZD | 07/19/2021 | 06/11/2024 |  | 06/11/2024 |  |  | 07/19/2021 | 06/11/2024 |  |  | 04/30/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070535 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported that a 82 year old male patient was placed on impella cp support. it was reported that patient received blood throughout the night and bleeding from multiple sites. urine was noted as bloody therefore heparin was restarted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
25489 | 1220648-2024-12351 | 19516739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19516739&pc=OZD | 07/19/2021 | 05/20/2025 |  | 06/11/2024 |  |  | 07/19/2021 | 06/11/2024 |  |  | 05/31/2023 |  | 06/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089365 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported an attempt was made to insert an impella cp via the same access (as previous impella 5.0) in a 65 year old male patient but it did not go through. so, an impella cp was inserted via the right femoral artery. the pump turned on without any problems. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25490 | 1220648-2024-12216 | 19497326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497326&pc=OZD | 07/19/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/19/2021 | 06/10/2024 |  |  | 05/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084682 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported that a 73 year old male patient began to decompensate and blood pressure went from 130/68 mmhg to 70/60 mmhg. decision was made to place impella cp. it was reported that while the pump was going over the aortic arch, the 0.018 wire kicked out of the left ventricle. the pump was removed and flushed to attempt reinsertion. reinsertion attempted, but unable to pass impella beyond the iliac and the impella cp was removed. the patient was stable and impella cp insertion was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25491 | 1220648-2024-12217 | 19497371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497371&pc=OZD | 07/16/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/16/2021 | 06/10/2024 |  |  | 06/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089667 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was found that the patient's right lower extremity was mottled and showed no distal pulses on doppler. the physician chose to explant the impella. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25492 | 1220648-2024-12218 | 19497401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497401&pc=OZD | 07/16/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/16/2021 | 06/10/2024 |  |  | 12/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022032863 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion of the impella cp, resistance was met when entering the left ventricle. the device appeared to be hugging the posterior wall. the decision was made to remove the impella and re-access the left ventricle with the pigtail. the impella cp was flushed before re-insertion and successfully advanced, though the patient was noted to have a short, calcified aortic root and small left ventricle. the patient then remained on impella support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25493 | 1220648-2024-12219 | 19497559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497559&pc=OZD | 07/05/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/05/2021 | 06/10/2024 |  |  | 05/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087387 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported there were some suction alarms in which was the patient was given blood products. the patient continued on support until the device was successfully weaned. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25494 | 1220648-2024-12220 | 19497577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497577&pc=OZD | 02/19/2024 | 06/10/2024 |  | 06/10/2024 |  |  | 02/19/2024 | 06/10/2024 |  |  | 01/31/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024314151 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 82 -year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hematoma and oozing at femoral site following peel away placement. manual pressure was held for 30min, and hematoma expressed. hematoma and oozing resolved, two sutures placed and the site was dressed using best practices. additionally, at start of procedure suction alarms were occurring. volume was given, alarms unresolved and repositioning of the device was attempted. patient starting lvedp -1.5. it was stated that the patient may be hypovolemic. physician decided to change out the cp to another cp device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
25495 | 1220648-2024-12221 | 19497604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497604&pc=OZD | 07/06/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/06/2021 | 06/10/2024 |  |  | 07/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1494063 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. there was oozing observed around the access site, near the sheath, when the patient was transferred from the table to the hospital bed. the device was explanted several hours later. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25496 | 1220648-2024-12222 | 19497613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497613&pc=OZD | 07/19/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/19/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the patient was bleeding from the impella access site. the site was redressed and the sheath was angle matched to stop the bleeding, however later it was noted that there was more bleeding, and systemic heparin was held until bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25497 | 1220648-2024-12223 | 19497620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497620&pc=OZD | 05/12/2024 | 07/17/2024 |  | 06/10/2024 |  |  | 05/12/2024 | 06/10/2024 |  |  | 09/30/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444901 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The user facility reported a 54-year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia during support. the medical staff performed one round of cardiopulmonary resuscitation and epinephrine was given to achieve a return of spontaneous circulation. coinciding with the event, a pump positioning alarm sounded. the medical staff noted the pump was too deep and successfully repositioned the pump. the patient eventually expired. upon review by abiomed's medical safety officer, there is no association between impella use and the patient¿s outcome. the patient's peri-procedural condition was poor. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for the impella cp with smart assist system lists ventricular arrhythmia an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." f6/f8 should have been left blank in the initial report.
25498 | 1220648-2024-12224 | 19497642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497642&pc=OZD | 07/07/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 01/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1469675 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support it was reported that heparin was removed due to bleeding concerns. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25499 | 1220648-2024-12226 | 19497644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497644&pc=OZD | 07/07/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 05/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085181 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during insertion a large, softball sized hematoma was noted at the access site. forward tension sutures were placed, manual pressure was applied and angle matching was completed. hemostasis was achieved. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25500 | 1220648-2024-12225 | 19497645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497645&pc=OZD | 06/27/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 06/27/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, multiple suction alarms and low flows occurred. additionally, the patient was bleeding from everywhere. the patient was given one unit of packed red blood cells and it was recommended to turn off heparin drip off. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25501 | 1220648-2024-12227 | 19497688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497688&pc=OZD | 07/07/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083202 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was attempted to be implanted with an impella device for mechanical circulatory support. it was reported that insertion of the impella cp was unsuccessful due to the patient's anatomy. the case was aborted. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25502 | 1220648-2024-12229 | 19497713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497713&pc=OZD | 07/19/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/19/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022074252 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had an increased white blood count and the physician had concern the patient was septic. the plan of care had been to remove the impella the same afternoon that sepsis was suspected as the patient was hemodynamically stable. intervention for the suspected sepsis is unknown. the patient was weaned and the impella cp was successfully removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The cause of the infection /sepsis was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25503 | 1220648-2024-12228 | 19497718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497718&pc=OZD | 06/28/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 06/28/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022074337 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, concerns for hemolysis due to hematuria was noted. no resolution provided for hematuria. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25504 | 1220648-2024-12230 | 19497727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497727&pc=PYX | 04/10/2023 | 07/05/2024 |  | 06/10/2024 |  |  | 04/10/2023 | 06/10/2024 |  |  | 07/31/2024 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2023281621 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 9 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | LVAD | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported a 9-year-old female admitted with unexpected right heart failure and left ventricular assist device (lvad) placement and impella rp was implanted for circulatory support. the pump was placed in (b)(6) 2023, and the support was for a total of 3 days, the team later in (b)(6) 2024 relayed there was hemolysis that occurred during the support that prompted dialysis. additionally, there was fibrous material within the inlet cage at the time of explant. the patient was reported as stable and fully recovered. | [5 empty] | The investigation for the reported hemolysis has been completed. data logs show numerous ingestion-like instances. the clinical states that a clot was observed in the inflow cage on the pump upon implant. the impella was returned for evaluation. a small amount of biomaterial was observed on the impeller on the returned device. the pump passed hemolysis testing with an modified index of hemolysis (mih) of 2.94. the root cause of the hemolysis was most likely due to biomaterial. added h6 component code 925 f6/f8 should have been left blank. g1 reporting contact fax number missing. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
25505 | 1220648-2024-12231 | 19497839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497839&pc=OZD | 07/07/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 07/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1490017 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the clinical team experienced a difficult insertion of the impella due to hematoma at insertion site. at the end the test wean proved achievable, so device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
25506 | 1220648-2024-12232 | 19497862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497862&pc=OZD | 07/20/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/20/2021 | 06/10/2024 |  |  | 03/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022065732 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. heparin was removed from the purge solution until bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25507 | 1220648-2024-12233 | 19497897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497897&pc=OZD | 06/29/2021 | 09/25/2025 |  | 06/10/2024 |  |  | 06/29/2021 | 06/10/2024 |  |  | 10/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021008196 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia and the impella pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation for the reported limb ischemia - reversible has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25508 | 1220648-2024-12235 | 19497955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497955&pc=OZD | 06/04/2024 | 07/24/2024 |  | 06/10/2024 |  |  | 06/04/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504025 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Death | Life Threatening | Disability | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported a 88-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and after the explant, the patient had a perforation and extravasation from the arteriotomy/sheath site. a covered stent was placed, but the extravasation worsened. a balloon tamponade was placed. the patient coded and was resuscitated. an echocardiogram was performed, and the ejection fraction had dropped to 20%. the patient coded again and expired in the operating room. | 2 units of blood products were administered as treatment. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The pump and introducer set were returned and evaluated. no damage was found on the introducer hub, valve, sheath, or tip. there were no other abnormalities found. the repositioning sheath was not used in the field and therefore not evaluated. data logs were not applicable for this investigation. the root cause of the vascular damage was use related because the complaint occurred after pump removal during closure. d9 date device returned to manufacturer added. instructions for use: section: warnings and cautions: cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12235: b2 date of patient death missing. b5 additional treatment of blood transfusion was missing from narrative. f6/f8-should have been left blank . g1 reporting contact fax number missing. h6 health effect - impact code codes 1802 and 4643 missing.
25509 | 1220648-2024-12234 | 19497997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19497997&pc=OZD | 07/07/2021 | 05/20/2025 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080347 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that there was a distal left anterior descending perforation from the guidewire. the issue was treated with a thrombin patch. the impella was weaned and explanted without issue. bedside echo revealed no effusion from perforation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25510 | 1220648-2024-12236 | 19498032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498032&pc=DWA | 03/22/2024 | 06/10/2024 |  | 06/10/2024 |  |  | 03/22/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | DWA |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1360581 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the backup automated impella controller (aic) displayed 'red battery error' after starting. the console was used to implant the impella cp. after successful pump insertion, the patient was unable to be transferred to the icu due to the battery problem. the issue was unable to be resolved with reset attempt. the console was exchanged without any patient harm reported. the patient expired on support. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | (h3) the investigation into the reported "aic - battery failure (red alarm)" has been completed since the original report was filed. product was returned for investigation. the console was received back for investigation, the batteries were confirmed to be unable to power the console and could not charge. replacing the batteries, with batteries that were not locked restored charging capabilities and provided full power to the system. during data analysis, the console ic6235 was turned off on (b)(6) 2023 with the batteries registering at cell voltages of 4v which was the last time the console was used before the complaint date (ic6235 ltpower 4_18_23.png). 11 months later on (b)(6) 2024, the console immediately has a battery failure (alarm #360). at this point, the console registered cell voltages around 2.3-2.5v. the cause of the automated impella controller (aic) issue was a depleted battery due to most likely the console not being charged in the year between uses.
25511 | 1220648-2024-12237 | 19498047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498047&pc=OZD | 05/31/2024 | 07/22/2024 |  | 06/10/2024 |  |  | 05/31/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462131 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Component Incompatible | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the automated impella controller (aic) had a purge disc not detected alarm when priming the cassette. staff switched out the cassette and still received purge disc not detected alarm. controller was switched out and cassette primed as expected. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | After reviewing the data logs, the issues with the automated impella controller (aic) not being able to detect the purge disc was confirmed. device analysis: purge disc flag was found to be broken (missing at the blue purge disc retainer). the root cause of the observed issue was a broken purge disc flag. the investigation for aic - purge disc/flag issues has been completed. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12237: d4 model number. d6a/d6b should have been left blank. f6/f8-should have been left blank. g1 manufacturing site facility name and manufacturing site fax number.
25512 | 1220648-2024-12238 | 19498077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498077&pc=OZD | 06/03/2024 | 08/09/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482119 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a 79-year-old female patient with pre-operation placement was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient¿s access site started bleeding. medical staff administered two units of red blood cells to the patient. the pump was explanted after six days, and the patient survived the event. | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement, the cause of the bleeding was due to the patient's movement as per the clinical description provided. b5 updated with additional information the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12238: b2 outcomes attributed to adverse event inadvertently left blank. f6/f8-should have been left blank. g1 reporting contact fax number missing. g2 report source: foreign was inadvertently not selected.
25513 | 1220648-2024-12239 | 19498110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498110&pc=OZD | 06/03/2024 | 07/25/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1762153 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received determined the automated impella controller (aic) had inadvertently been dropped in the user facility leading to the aic replacement. | The user facility reported that the automated impella controller (aic) display flickered intermittently during patient support. it was noted there were no changes in flow or support and there was no report of patient harm. |  |  |  |  | Additional information added to b5. | Console was received for investigation. there was no explicit evidence in the logs to confirm the reported failure mode ¿ flickering screen. this console was tested after arriving in fs. immediately after the initial boot-up, the console started flickering, and the flickering pattern was horizontal (flickering_screen_ic10971.mov). after power cycling the console, there was no issue. after some time, a gain observed the flickering screen. this issue occurs occasionally, which confirms the reported failure mode (fm-32022). no damage was found on the console hardware. the root cause of the intermittent flickering screen was not determined. the root cause for dropping the automated impella controller (aic) was the use issue. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12239: g1 reporting contact fax number missing. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the aic was not possible. upon investigation closure, a supplemental mdr will be filed.
25514 | 1220648-2024-12240 | 19498111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498111&pc=OZD | 06/03/2024 | 08/01/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486646 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a 47-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient¿s right lower extremity became mottled below the access site. medical staff were unable to obtain pulses and monitored the patient¿s condition. the patient remains on support. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- b5 mso review f6/f8 should have been left blank in the initial report.
25515 | 1220648-2024-12241 | 19498123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498123&pc=OZD | 06/03/2024 | 07/23/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471119 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella 5.5 device for mechanical circulatory support to a patient experiencing postcardiotomy cardiogenic shock. the impella 5.5 pump was inserted; however, due to the patient¿s anatomical positioning of the subclavian, it was unsuccessful to traverse the cannula through the vessel. manipulation of vessel, abduction of arm and shoulder bump attempted. stiff buddy wire was also utilized without success. the impella 5.5 pump was pulled and attempted again without short lateral tunnel and stiffer .018 guidewire without success. the impella insertion procedure was aborted. there was no report of adverse effects to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, device inserted but due to anatomical positioning of subclavian, unable to traverse cannula through vessel as per the clinical description. d4 model number
25516 | 1220648-2024-12242 | 19498165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498165&pc=PYX | 03/22/2024 | 07/11/2024 |  | 06/10/2024 |  |  | 03/22/2024 | 06/10/2024 |  |  | 10/31/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024418651 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The complainant reported a patient (unknown race) in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support due to there being no exit strategy on extracorporeal membrane oxygenation. during support, the impella rp appeared to have fallen back into the right ventricle spontaneously when placing. the physician attempted to advance but was unsuccessful with torqueing and forward tension. pigtail appeared past the pulmonary valve, but the outlet appeared to be at the valve. the position of rp flex on x-ray showed probable placement of outlet at the pulmonary valve and pigtail remains in pulmonary artery (pa). color doppler shows forward flow into pa and decision to leave as is. care was withdrawn and the patient expired. there is no association between impella use and outcome. | [5 empty] | According to the clinical, during the first day of impella rp support the pump was inserted without difficulty. a couple hours later the pump appeared to have fallen back into the rv spontaneously. multiple attempts were made to reposition the pump without success. shortly, after the positioning issue was observed the patient also began to develop signs of hemolysis. due to the patient's declining condition the decision was made to withdraw the patient from care. the impella rp device was not received from the customer, and therefore an evaluation of the device was not possible. however, data logs were received for evaluation/analysis and confirmed suction alarms with no evident positioning issues. impella position was unknown per logs. product history review revealed the pump passed all post-sterile inspection checks. the most likely cause of the positioning issue was use related. the most likely cause of the hemolysis was improper positioning. correction was made to the following field(s)/section(s): f6: date removal. f8: date removal. g4: pma number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25517 | 1220648-2024-12243 | 19498183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498183&pc=OZD | 06/01/2024 | 07/18/2024 |  | 06/10/2024 |  |  | 06/01/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1382144 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/01/2024 |  |  |  |  | The user facility reported a 58-year-old male with pre-operation placement was implanted with an impella 5.5 for mechanical circulatory support. it was reported a purge disc not detected alarm occurred during support. despite troubleshooting, the issue remained. the medical staff then switched to a backup console. no patient harm was reported. | [5 empty] | The investigation for the console (aic) purge flag/disc issues has been completed. data logs were reviewed which confirmed that the purge disc couldn't be detected was confirmed. the device was returned for evaluation and the purge disc retainer was found to be broken. the root cause of the issue was a broken purge disc retainer. added- d9 device return date f6/f8 should have been left blank in the initial report. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the aic was not possible. upon investigation closure, a supplemental mdr will be filed.
25518 | 1220648-2024-12244 | 19498192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498192&pc=OZD | 07/07/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 03/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060577 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult or Delayed Positioning | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that implantation was very difficult due to the patient's anatomy. the patient was kept on support for a few days until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25519 | 1220648-2024-12247 | 19498210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498210&pc=OZD | 07/07/2021 | 05/13/2025 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 12/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022032153 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it is reported that arteriosclerosis obliterans was identified in both lower limbs making insertion of the peel away sheath difficult. the medical staff managed to insert the sheath utilizing a buddy guide wire and plain old balloon angioplasty, but the impella pump catheter became stuck at the aortic arch due to the severe tortuosity and calcification of the patient¿s vasculature. due to unsuccessful attempts at delivery, the medical staff made the decision to abort the process and remove the impella. the device was explanted; however, the procedural outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25520 | 1220648-2024-12248 | 19498213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498213&pc=OZD | 06/04/2024 | 10/25/2024 |  | 06/10/2024 |  |  | 06/04/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1394733 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the pump stopped, and that the impella aic (automated impella controller) shut off. the aic was removed from service. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) pump stop has been completed. data logs were reviewed for evidence of the reported complaint. the reported pump stop was confirmed. the reported unexpected shutdown was not confirmed. there were at least seven instances of an impella stopped event due to the epc detecting a high motor current (event set #13). real time logs confirmed that when the impella stopped events occurred, pump flow rate dropped to 0. the console was returned for evaluation and the reported complaint could not be reproduced in the lab during testing. the reported complaint was determined to have been potentially caused by intermittent malfunctions of the impellatronic pca, however this was not conclusively determined. added- d9 device return date corrections to the initial mfr report: d4-model number, f6/f8 should have been left blank.
25521 | 1220648-2024-12245 | 19498226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498226&pc=OZD | 05/01/2024 | 06/10/2024 |  | 06/10/2024 |  |  | 06/04/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1313442 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Open or Close | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/04/2024 |  |  |  |  | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp placement the automated impella controller (aic) purge door was unable to open. the aic was removed from service. there was no patient harm. the investigation results indicates that the severity of this complaint is s5. | [5 empty] | Data review: no data review was needed as this complaint was only made for an issue with the consoles inability to open the purge door. device analysis: console was tested after arriving in field service. after examination of the purge door assembly, it was evident that there was dextrose buildup on the purge door and around the purge door latch. this dextrose buildup is what caused the purge door to be stuck. before sending this console back to the customer, field service was instructed to clean the purge door and area around the purge door latch, perform steps 10-12 of the preventative maintenance (pm) procedure and perform a full functional check on the console and optical system. root cause: the root cause of this issue was dextrose buildup on the purge door and around the purge door latch. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
25522 | 1220648-2024-12246 | 19498230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498230&pc=OZD | 06/29/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 06/29/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078702 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was placed on venoarterial extracorporeal membrane oxygenation and transfused blood products due to continuous bleeding from the access site. the patient was transported to a different medical facility for additional treatment and escalation of care from the impella cp to the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25523 | 1220648-2024-12249 | 19498244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498244&pc=OZD | 10/09/2023 | 10/24/2025 |  | 06/10/2024 |  |  | 10/09/2023 | 06/10/2024 |  |  | 07/31/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024376238 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/09/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following implantation of impella cp, the automated impella controller (aic) did not display metrics. the white connector was reported to be correctly locked. the impella was explanted and a new one was implanted without any issues. there was no known harm or injury and no known patient use. | [5 empty] | The impella device was received from the customer on 25-oct-2023. impella cp was returned for evaluation. pump connected to console and no placement signal issue alarm was reproduced. visually inspected and no catheter kink or other abnormalities were seen. optical bench 5s parameters were verified to be within normal range. optical light continuity test was done and passed without any issues. no other abnormalities were found. data logs were not returned for review. as per clinical, impella cp implanted with no metric seen on console with impella sensor failure alarm and new impella cp was replaced with current cp and worked without any complications. the cause of the optical signal issue was not determined since data logs were not returned and due to inconclusive results based on the field investigation. | D6a: correct implant date. d6b "if explanted, give date" corrected. | G3 ¿date received by manufacturer¿ corrected.
25524 | 1220648-2024-12250 | 19498247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498247&pc=OZD | 03/22/2024 | 06/12/2024 |  | 06/10/2024 |  |  | 03/22/2024 | 06/10/2024 |  |  | 10/31/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420287 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. pump is running at p9 with 3.9 l/min, patient is stable underneath, signs of hemolysis continue, billirubin and liver values are rising, just checked the position, pump is supposedly in a good position, very difficult echo conditions, as the patient is extremely obese and very difficult to position. duplex mentioned again. catecholamines could already be reduced. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and positioning has been completed. the complaint device was not returned for investigation. there were no data logs to review. the cause of the hemolysis issue cannot be determined since insufficient clinical data provided and no data logs/product evidence to conclude. the cause of the positioning issue most likely was use related due to improper technique as the pump was pulled out by the patient. impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ added- b1 selected product problem d4-updated model number d6-implant and explant dates g1-corrected mfr site name and address h5 corrected to yes
25525 | 1220648-2024-12251 | 19498266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498266&pc=OZD | 07/07/2021 | 05/27/2025 |  | 06/10/2024 |  |  | 07/07/2021 | 06/10/2024 |  |  | 12/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022025042 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that bleeding was observed at the groin access sited due to the patient not being provided enough sedation. a c-clamp was placed to control the bleeding. additionally, suction alarms were received, and echocardiogram was utilized to confirm good pump position. it was also reported that the patient experienced a cardiac device infection. there is no indication of how hemostasis was achieved regarding infection. weaning was successful, the device was explanted, and the procedural outcome is the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25526 | 1220648-2024-12252 | 19498281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498281&pc=OZD | 05/31/2024 | 07/30/2024 |  | 06/10/2024 |  |  | 05/31/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451712 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Ischemia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 56-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that pulses in the patient¿s bilateral extremities were not present. additionally, there were ongoing suction alarms and the pump was not in the proper position. the medical staff decided to explant the pump and a thrombus was observed coming out of the access site. the patient survived the event. | [5 empty] | Zthe impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the limb ischemia, low pump flow, and positioning issues has been completed. product was returned in damaged condition (without pump plug). visual inspection found biomaterial inside the pump housing. no other issues were found on the rest of the pump. data logs were reviewed which showed 14.5 hours into support, low flow occurred that triggered suction alarms. this event remained to the rest of the case.the root cause of the low flow was biomaterial ingestion. the root cause of the limb ischemia could not be determined without additional clinical details. the root cause of the positioning issue could not be determine without additional details. f6/f8 should have been left blank in the initial report.
25527 | 1220648-2024-12253 | 19498284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498284&pc=OZD | 07/08/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/08/2021 | 06/10/2024 |  |  | 03/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1477975 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/08/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding, including a hematoma, which required blood transfusion. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25528 | 1220648-2024-12254 | 19498290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498290&pc=OZD | 03/22/2024 | 06/12/2024 |  | 06/10/2024 |  |  | 03/22/2024 | 06/10/2024 |  |  | 10/31/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024409466 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted escalation of support. as a result of the noted resistance, the decision was made to abort the implant. the patient remained supported with an already implanted impella cp pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since resistance was met due to patient anatomy. in accordance with updated procedures, sections d6 have been revised from the initial submission.
25529 | 1220648-2024-12255 | 19498310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498310&pc=OZD | 07/08/2021 | 09/23/2025 |  | 06/10/2024 |  |  | 07/08/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079792 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella pump pulled out of the aortic valve and was unable to be repositioned. the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25530 | 1220648-2024-12256 | 19498325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498325&pc=OZD | 07/11/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/11/2021 | 06/10/2024 |  |  | 03/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069067 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that no doppler pulse was found in the left foot. it is unknown if/how this issue was resolved, but the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25531 | 1220648-2024-12257 | 19498343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498343&pc=OZD | 07/15/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/15/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079014 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a perclose was used at the access site, but it continued to ooze, even with manual compression. a femstop was then used with light pressure and subsequently the device was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25532 | 1220648-2024-12258 | 19498345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498345&pc=OZD | 06/29/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 06/29/2021 | 06/10/2024 |  |  | 05/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 06/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump could not be placed through the prosthetic aortic valve. there was too much length to place the impella back through the new prosthetic valve. the decision was made to cut the cp catheter. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
25533 | 1220648-2024-12259 | 19498380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498380&pc=OZD | 01/31/2024 | 07/18/2024 |  | 06/10/2024 |  |  | 05/16/2024 | 06/10/2024 |  |  | 11/30/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured artery puncture with a "definite" relationship to the impella device used: ¿underwent impella-assisted rotational atherectomy/pci complicated by l femoral artery perforation. stent placed in left femoral artery in order to control the bleeding¿. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
25534 | 1220648-2024-12260 | 19498513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498513&pc=OZD | 06/30/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 06/30/2021 | 06/10/2024 |  |  | 05/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while inserting the impella, there was moderate difficulty getting wire/pigtail into descending aorta due to severe calcification at the aortic bifurcation. there were numerous attempts made to insert the impella, but all were unsuccessful crossing the aortic bifurcation. ultimately, the surgeon decided to abort impella placement and sent the patient to the cardiac care unit for monitoring. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
25535 | 1220648-2024-12261 | 19498533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498533&pc=OZD | 05/31/2024 | 07/23/2024 |  | 06/10/2024 |  |  | 05/31/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481924 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 67-year-old female with non-ischemic cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the implant was only an attempt because the pump was unable to advance past the patient¿s anastomosis. an impella cp was instead implanted successfully. no patient harm has been reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the 5.5 was unable to be advanced past the anastomosis as per the clinical description. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number
25536 | 1220648-2024-12262 | 19498543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498543&pc=OZD | 05/21/2024 | 08/01/2024 |  | 06/10/2024 |  |  | 05/21/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the automated impella controller (aic) indicated it had an unexpected shutdown. there was no patient harm. the aic was swapped out with another aic. the aic is being returned for investigation. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) shutdown or restart issue has been completed. upon reviewing the data logs, no evidence was found to confirm the reported failure mode. this console was returned for evaluation and upon running the test optical pump for more than 16 hours, it was unable to reproduce the reported failure mode - unexpected shutdown. there was no issue found with the console hardware. the device functioned/operated to specification. added- b5 mso review, d9 device return date f6/f8 should have been left blank in the initial report.
25537 | 1220648-2024-12263 | 19498603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498603&pc=OZD | 05/30/2024 | 07/24/2024 |  | 06/10/2024 |  |  | 05/31/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025470724 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The complainant reported a male patient of an unknown age in st (sinus tachycardia) elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient was lysed due to lost pulse at the limb. then, the patient developed a compartment in the right thing, which had to be surgically split. blood circulation in the leg was successfully restored. the pump was successfully explanted two days later. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 health effect - impact code 4324, component coded added-3038 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12263: d4 model number. f6/f8-should have been left blank . g1 manufacturing site name and address.
25538 | 1220648-2024-12264 | 19498618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498618&pc=OZD | 05/31/2024 | 07/15/2024 |  | 06/10/2024 |  |  | 05/31/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482947 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 54-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the left axillary placement was aborted due to the impella being unable to pass through the patient¿s vasculature. no patient harm has been reported. | [5 empty] | The root cause of the delivery issue was most likely patient condition since the pump wasn't able to cross secondary due to stenosis in left subclavian. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25539 | 1220648-2024-12265 | 19498639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498639&pc=OZD | 06/01/2024 | 07/23/2024 |  | 06/10/2024 |  |  | 06/01/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490843 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2024 |  |  |  |  | The complainant reported a 64-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the pump experienced continuous suction during support. medical staff added volume and repositioned the pump, but the issue persisted. medical staff then decided to replace the pump. no patient harm has been reported, and the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for low pump flow has been completed. compliant device not returned for investigation. data log reviewed: motor current spikes observed. many suction alarms occurred. drop in flow, at p1 and 2 flow became zero. no positioning issue observed. the root cause of the low pump flow most likely was biomaterial ingestion. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12265: f6/f8-should have been left blank. h.6 code 4445, 4627, 4641, and 4628 were reported incorrectly. new code was added to health effect - clinical code.
25540 | 1220648-2024-12266 | 19498772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498772&pc=OZD | 06/02/2024 | 07/22/2024 |  | 06/10/2024 |  |  | 06/02/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473906 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2024 |  |  |  |  | The complainant reported a 64-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported, despite all efforts from the medical staff, delivery and wire exchanges failed. medical staff aborted the 5.5 procedure and implanted an impella cp instead. | [5 empty] | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. an impella 5.5 was attempted to be delivered surgically through the right axillary artery to the patient. initially, a tunneling method was utilized, but the pump was unable to cross the anastomosis. once dissecting the tissue in between to create one large insertion pocket, the surgeon could pass the anastomosis but was then unable to make the turn from the axillary artery into the ascending aorta. data logs were not applicable for this case. the root cause of the delivery issues was most likely patient condition and more specifically, diseased/small vessels. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25541 | 1220648-2024-12267 | 19498891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498891&pc=OZD | 07/01/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/01/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081989 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were no pulses on the right leg. the femoral-femoral (fem-fem) bypass was clotted, so it was flushed, and the flow was restored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25542 | 1220648-2024-12268 | 19498918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19498918&pc=OZD | 05/30/2024 | 07/16/2024 |  | 06/10/2024 |  |  | 05/30/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465645 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous intervention was implanted with impella cp device for mechanical circulatory support and after the implant, low pump flows were noted. the flow decreased with flattened motor current. the pump was repositioned to no avail, and multiple attempts were made to ramp up and down on p-levels without success. consequently, the impella cp pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Data review: data logs confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. product was not returned for review. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review g1 reporting contact fax number missing.
25543 | 1220648-2024-12270 | 19499013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499013&pc=OZD | 07/20/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/20/2021 | 06/10/2024 |  |  | 01/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022041506 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2021 |  |  |  |  | The complainant reported a patient of unknown race and ethnicity in st elevation myocardial infarction and cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. a femoral angiogram revealed the patient had small vessel size and tortuous anatomy. the impella sheath was attempted to be inserted, but was unsuccessful. the impella cp insertion was aborted due to the patient's femoral anatomy. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records.
25544 | 1220648-2024-12269 | 19499015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499015&pc=OZD | 06/30/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 06/30/2021 | 06/10/2024 |  |  | 05/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1478929 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the urinary output was dark red in color and the patient was neurologically unresponsive. groin site bleeding was observed, and the access site was re-sutured and tightened with a perclose. the patient was placed on continuous renal replacement therapy after declining overnight. the patient¿s condition gradually improved | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25545 | 1220648-2024-12271 | 19499127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499127&pc=OZD | 06/02/2024 | 07/30/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471162 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock implanted with an impella 5.5 pump for mechanical circulatory support experienced drain site bleeding and a stroke. midway through time on support, heparin was paused due to bleeding at the pericardial drain site and was later resumed. additionally, during explant of the device (after approximately 30 days of support), the patient experienced a stroke. the status of the patient was noted to be stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the stroke and vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke and vascular damage could not be determined. added-h6 component code 925.
25546 | 1220648-2024-12272 | 19499139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499139&pc=OZD | 07/02/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/02/2021 | 06/10/2024 |  |  | 12/31/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022021299 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention (aortic valvuloplasty) was implanted with an impella cp device for mechanical circulatory support. after support, when attempting to close the access site, 4 perclose sutures failed. manta deployed to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction : section h1 : type of reportable event should have been malfunction but serious injury was selected in error in the initial report.
25547 | 1220648-2024-12273 | 19499278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499278&pc=OZD | 07/01/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/01/2021 | 06/10/2024 |  |  | 02/28/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump flows were low, and volume was required. the patient was transferred to intensive care where a gastrointestinal (gi) bleed was observed. the heparin was removed from the purge solution due to bleeding concerns. the patient also went into atrial fibrillation (afib) during the course of the night and volume was provided to the patient. heparin was restarted and the groin site was stated to be warm, dry and intact with some bruising. the patient was transferred to a different medical facility for further treatment as their condition improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25548 | 1220648-2024-12274 | 19499324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499324&pc=OZD | 07/16/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/16/2021 | 06/10/2024 |  |  | 06/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092297 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. while inserting the impella, the physician was unable to cross the transverse arch. additional efforts were made with different guidewires, but all attempts were unsuccessful, and the process was aborted. the patient outcome is unknown. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25549 | 1220648-2024-12275 | 19499386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499386&pc=OZD | 07/01/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/01/2021 | 06/10/2024 |  |  | 11/30/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021018542 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the flows were low, and it was unclear if the impella was properly positioned and patient required multiple shocks, so impella was removed, and veno-arterial extracorporeal membrane oxygenation was initiated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25550 | 1220648-2024-12276 | 19499451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499451&pc=OZD | 07/01/2021 | 03/20/2025 |  | 06/10/2024 |  |  | 07/01/2021 | 06/10/2024 |  |  | 03/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022062025 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the impella leg pulses were unobtainable. the patient is doing well from a cardiac standpoint. no additional information was provided regarding the reported issue. | Additional information noted the device was explanted to address the issue. the patient survived the procedure. |  |  |  |  | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. correction: section b5 : additional information was updated which was inadvertently left out in the initial report. section h6 : the health impact code was updated which was inadvertently left out in the initial report.
25551 | 1220648-2024-12277 | 19499530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499530&pc=OZD | 07/02/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/02/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069839 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hematoma and urine was pink. systemic heparin was stopped. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25552 | 1220648-2024-12278 | 19499680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499680&pc=OZD | 07/20/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/20/2021 | 06/10/2024 |  |  | 01/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022038638 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient¿s urine was stated to be tea colored with hemolyzed lab results. it was recommended to the medical team to confirm impella cp position with transesophageal echocardiography. the patient's urine was reported to remain amber colored throughout impella support until the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25553 | 1220648-2024-12279 | 19499688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499688&pc=OZD | 06/03/2024 | 07/29/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506828 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a 77-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the physician discovered the patient¿s aortic stenosis during the implant procedure and therefore could not cross the aortic valve. the procedure was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to cross av due to severe aortic stenosis present as per the clinical description. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25554 | 1220648-2024-12280 | 19499817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499817&pc=OZD | 06/03/2024 | 07/24/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock implanted with an impella cp device experienced an access site complication. there was right groin site bleeding, and a hematoma was present. to address, sutures were tightened, the site redressed, and two units of packed red blood cells were given for 3pt hemoglobin (hgb) drop. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since the hemostasis was achieved by tightening the sutures. h6 component code 4755 changed to 925.
25555 | 1220648-2024-12281 | 19499819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499819&pc=OZD | 05/30/2024 | 07/15/2024 |  | 06/10/2024 |  |  | 05/30/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1078703 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a battery communication failure (yellow alarm). the aic was replaced. no patient harm has been reported. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the data logs confirm the failure mode of controller error alarm. the aic was returned for evaluation. the controller error and loss of placement signal was due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by the embedded personal computer (epc). the aic software operated as designed. f6 and f8 should have been blank. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
25556 | 1220648-2024-12282 | 19499835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19499835&pc=OZD | 06/30/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/02/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022073698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/02/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had little to no pedal pules and a cold leg. the patient was monitored overnight and the plan of care is to be determined. no additional information provided. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25557 | 1220648-2024-12283 | 19500017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19500017&pc=OZD | 07/02/2021 | 02/21/2025 |  | 06/10/2024 |  |  | 07/05/2021 | 06/10/2024 |  |  | 02/28/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050841 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, heparin induced thrombocytopenia, and acute renal failure was noted. heparin was removed from the purge solution. sodium bicarbonate was added in the purge solution and helped very well. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombocytopenia and renal failure has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25558 | 1220648-2024-12284 | 19500048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19500048&pc=OZD | 06/03/2024 | 07/30/2024 |  | 06/10/2024 |  |  | 06/03/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462115 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during left anterior descending artery surgery. the impella cp was explanted after a day of support due to the visualization under echo of new left ventricle thrombus. patient survived event and is supported now with an intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for thrombus has been completed. no device was received. as the necessary information was not provided, the root cause of the thrombus was not determined. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12284 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12284 was submitted. h.6 code 4755 was entered incorrectly on the previously submitted manufacturer device report number 1220648-2024-12284. a new code was added.
25559 | 1220648-2024-12285 | 19500129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19500129&pc=OZD | 05/30/2024 | 08/01/2024 |  | 06/10/2024 |  |  | 05/30/2024 | 06/10/2024 |  |  |  |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025451664 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 35-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support and was scheduled for a heart transplant. the patient was complaining of pain and decreased sensations in lower extremities. the patient was taken for a computed tomography where they revealed some, ¿ischemic issue with patient¿s spinal cord.¿ the patient's heart transplant was cancelled. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added-b5 mso review added-d6b, f6/f8 should have been left blank in the initial report.
25560 | 1220648-2024-12286 | 19500977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19500977&pc=OZD | 02/16/2024 | 09/26/2025 |  | 06/10/2024 |  |  | 02/26/2024 | 06/10/2024 |  |  | 11/30/2025 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that the patient had low platelets. the patient was provided one unit of plasma for high international normalised ratio and low platelets. patient is stable post issue. | The complainant also reported a placement signal unreliable alarm. correct pump position was confirmed. care was eventually withdrawn, and the patient expired. |  |  |  |  | The investigation has been completed. according to the clinical information, it was reported that the patient had low platelets. the patient was provided one unit of plasma. the patient was then deemed unfit for surgery and an upgrade or switch. no swaps were made and the pump remained in the patient until they withdrew from care. no product was returned for a visual inspection. the cause of the issue could not be determined due to insufficient clinical details. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for a visual inspection. data log analysis confirmed normal 5s parameters and a negative downward trend in the placement signal days before the placement signal unreliable alarm. this placement signal trend is consistent with sensor drift. the root cause of the issue could not be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b2 updated to death. b5 updated to include placement signal issue description and death. h1 updated to death. h6 updated to include death and placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-12286. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email updated.
25561 | 1220648-2024-12287 | 19501319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19501319&pc=OZD | 07/14/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/16/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081055 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding, and heparin was discontinued to resolve the issue. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25562 | 1220648-2024-12288 | 19502908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19502908&pc=OZD | 07/21/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/21/2021 | 06/10/2024 |  |  | 03/31/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057163 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2021 |  |  |  |  | The complainant reported that a patient was brought to the operating room to have a their existing impella cp removed and replaced with a new impella cp as a bridge to transplant. after the second impella cp was placed, the patient experienced oozing and systemic anticoagulants were held until bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25563 | 1220648-2024-12289 | 19503256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19503256&pc=OZD | 07/21/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/21/2021 | 06/10/2024 |  |  | 11/30/2022 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022026794 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, grossly bloody urine was noted and the patient received blood products as a result of the hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25564 | 1220648-2024-12290 | 19503546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19503546&pc=OZD | 07/22/2021 | 06/10/2024 |  | 06/10/2024 |  |  | 07/22/2021 | 06/10/2024 |  |  | 02/28/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022053991 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, systemic heparin was withheld after a the patient's platelet count dropped and the physician was concerned the patient was experiencing heparin induced thrombocytopenia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25565 | 1220648-2024-12291 | 19503964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19503964&pc=OZD | 07/22/2021 | 03/21/2025 |  | 06/10/2024 |  |  | 07/22/2021 | 06/10/2024 |  |  | 04/30/2023 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support during a ventricular tachycardia ablation. it was reported that while the patient was on impella cp support, the patient experienced bleeding and protamine was administered to the patient to reverse the effects of heparin the the bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25566 | 1220648-2024-12292 | 19506917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19506917&pc=OZD | 04/01/2022 | 03/21/2025 |  | 06/10/2024 |  |  | 04/01/2022 | 06/10/2024 |  |  | 01/31/2024 |  | 06/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184910 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the insertion procedure, the impella device placement was halted because of issues related to the patient's anatomy. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25567 | 1220648-2024-12198 | 19495246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495246&pc=OZD | 03/15/2022 | 03/31/2025 |  | 06/09/2024 |  |  | 03/15/2022 | 06/09/2024 |  |  | 11/30/2022 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021020259 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2022 |  |  |  |  | The user facility reported a 70-year-old male on impella cp for mechanical circulatory support. it was reported there were complications during closure after the impella cp device was removed. during explant, the perclose failed, and bleeding occurred. additional attempts were made to close the access site, however the vessel appeared to be damaged. a balloon tamponade was performed through via alternate access, but no improvement was noted. an angiogram was performed which revealed a substantial leak post balloon inflation. the leak was covered by an atrium medical peripheral stent deployed in the right common femoral artery. the device was successfully explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25568 | 1220648-2024-12199 | 19495251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495251&pc=OZD | 03/15/2022 | 03/25/2025 |  | 06/09/2024 |  |  | 03/15/2022 | 06/09/2024 |  |  | 12/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175990 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2022 |  |  |  |  | The user facility reported a 59-year-old patient was implanted with an impella cp for mechanical circulatory support as a high-risk percutaneous coronary intervention. the decision was made to implant the impella cp prior to an urgent coronary artery bypass graft being placed. upon arrival at the catheterization laboratory, the patient¿s lab results were positive for hemolysis. the impella cp device position was observed to be too deep, and the device was pulled back. it was reported that as urinary output (uop) was marginal, the patient was started on dialysis and continuous renal replacement therapy (crrt). sepsis was reported as a complication. one (1) unit of packed red blood cells was transfused to the patient. the issue was managed with fluid resuscitation pressors. it was reported that the device was successfully explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: renal failure section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ infection/sepsis section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25569 | 1220648-2024-12200 | 19495259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495259&pc=OZD | 03/24/2022 | 04/03/2025 |  | 06/09/2024 |  |  | 03/24/2022 | 06/09/2024 |  |  | 10/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155411 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a 82 year old male patient was implanted with an impella cp for mechanical circulatory support. it was reported that there was thrombosis in the left anterior descending from prior intervention. dr. preformed thrombectomy of left anterior descending followed by balloon angioplasty of mid/distal left anterior descending. three additional stents were deployed to mid/distal left anterior descending to restore flow. patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25570 | 1220648-2024-12201 | 19495264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495264&pc=OZD | 03/25/2022 | 06/09/2024 |  | 06/09/2024 |  |  | 03/25/2022 | 06/09/2024 |  |  | 11/30/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162335 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2022 |  |  |  |  | The complainant reported that a 70 year old male patient was implanted with an impella cp for mechanical circulatory support. during support, a hematoma at the impella insertion site was noted. femstop was placed in addition to 1 unit of blood that was given. the impella was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25571 | 1220648-2024-12202 | 19495266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495266&pc=OZD | 04/01/2022 | 03/25/2025 |  | 06/09/2024 |  |  | 04/01/2022 | 06/09/2024 |  |  | 11/30/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169806 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the procedure was aborted due to mitral valve positioning with pump. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
25572 | 1220648-2024-12203 | 19495268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495268&pc=OZD | 04/01/2022 | 03/25/2025 |  | 06/09/2024 |  |  | 04/01/2022 | 06/09/2024 |  |  | 12/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181133 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. delivery issues were experienced when attempting to implant the impella pump. due to size of the left ventricle (lv) cavity, the pump continued to dive deep. the pump was slowly pulled back and secured into place. the patient was successfully weaned and the device was explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
25573 | 1220648-2024-12204 | 19495287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495287&pc=OZD | 04/01/2022 | 03/25/2025 |  | 06/09/2024 |  |  | 04/01/2022 | 06/09/2024 |  |  | 10/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159049 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. the impella was attempted to be inserted into the sheath, but unable to advance beyond the valve in the sheath. decision was made to place 14fx25cm impella sheath. guide-wire and catheter were advanced into the left ventricle. the impella was inserted but tortuosity prevented advancement beyond bifurcation. in addition, one unit of packed red blood cells were administered due to blood loss. decision was made to remove impella and place intra-aortic balloon pump (iabp). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (delivery issues) are as follows: ¿section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25574 | 1220648-2024-12205 | 19495435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495435&pc=OZD | 04/02/2022 | 03/25/2025 |  | 06/09/2024 |  |  | 04/02/2022 | 06/09/2024 |  |  | 12/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170714 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the patient had notable bleeding at axillary impella site. the patient was taken to the operating room for exploratory washout of axillary site. in addition, it was reported that the patient was at intermediate risk for heparin induced thrombocytopenia (hit) due to low platelet count. heparin was removed from the purge solution to resolve the reported issue. the impella was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (thrombocytopenia) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25575 | 1220648-2024-12206 | 19495443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495443&pc=OZD | 04/04/2022 | 03/25/2025 |  | 06/09/2024 |  |  | 04/04/2022 | 06/09/2024 |  |  | 01/31/2024 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185229 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was showing signs of hemolysis, pink-tinged urine and hemolyzed labs. the impella cp was explanted and the patient was escalated to an impella 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25576 | 1220648-2024-12207 | 19495453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495453&pc=OZD | 04/13/2022 | 03/24/2025 |  | 06/09/2024 |  |  | 04/13/2022 | 06/09/2024 |  |  | 12/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5. |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023171580 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a patient noted as high-risk surgery was implanted with an impella cp device for mechanical circulatory support. patient was escalated from impella cp to impella 5.5 pump. during support, it was noted that pump had high purge pressure alarm. sodium bicarbonate was running through purge system and there were no visible kinks. the pump was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25577 | 1220648-2024-12208 | 19495460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495460&pc=OZD | 05/12/2022 | 03/21/2025 |  | 06/09/2024 |  |  | 05/12/2022 | 06/09/2024 |  |  | 12/31/2022 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022023851 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during explant, it was reported that insertion/access site was surgically sutured. weaning was successful, the device was explanted. patient tolerated the procedure and condition was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25578 | 1220648-2024-12209 | 19495538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495538&pc=OZD | 08/02/2022 | 03/21/2025 |  | 06/09/2024 |  |  | 08/02/2022 | 06/09/2024 |  |  | 10/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143688 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a cerebrovascular event (hemorrhagic) that occurred during impella insertion, however it was not discovered until after the removal of the impella. as a result, anticoagulant therapy was discontinued. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The cause of the stroke/ cva was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25579 | 1220648-2024-12210 | 19495554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495554&pc=OZD | 03/28/2022 | 10/07/2025 |  | 06/09/2024 |  |  | 03/28/2022 | 06/09/2024 |  |  | 06/30/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103657 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported that a 56 year old male patient was implanted with an impella cp. while on support, the patient experienced some bleeding and hematoma at the insertion site but was manageable due to an unspecified intervention. no other information was provided, however the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d5 operator of device was inadvertently omitted on the initial manufacturer device report number. f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g2 report source; foreign was missing from manufacturer device report h6 component code revised from the initial submission in accordance with updated procedures.
25580 | 1220648-2024-12211 | 19495638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495638&pc=OZD | 03/28/2022 | 10/30/2025 |  | 06/09/2024 |  |  | 03/28/2022 | 06/09/2024 |  |  | 08/16/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119963 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Unknown |  |  | Asian | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported that the patient was implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock. while on support, the complainant reported ventricular fibrillation did not stop after seven times of direct current defibrillation. while compressing the chest, the impella slipped into the aorta, so it was removed. veno-arterial extracorporeal membrane oxygenation (va-ecmo) was planned for hemodynamic stabilization. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25581 | 1220648-2024-12212 | 19495727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495727&pc=OZD | 03/27/2022 | 06/09/2024 |  | 06/09/2024 |  |  | 03/28/2022 | 06/09/2024 |  |  | 11/30/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169165 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a 53-year-old male patient was implanted with an impella cp for mechanical circulatory support. while on support, an echocardiogram revealed that the placement was slightly deep, however, it was decided to maintain the impella device in the current position. vascular surgery later removed the device due to concerns for high bifurcation. in addition, there was bleeding at the access site, so systemic heparin was discontinued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25582 | 1220648-2024-12213 | 19495743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495743&pc=OZD | 03/29/2022 | 04/16/2025 |  | 06/09/2024 |  |  | 03/29/2022 | 06/09/2024 |  |  | 12/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180706 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2022 |  |  |  |  | The complainant reported a 78 year old male patient needing an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock. after several attempts and catheter and wire exchanges, the physician was unable to gain access to the left ventricle. due to the difficulty and patient anatomy, the surgeon decided to abort the impella insertion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25583 | 1220648-2024-12214 | 19495768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495768&pc=OZD | 03/30/2022 | 06/09/2024 |  | 06/09/2024 |  |  | 03/30/2022 | 06/09/2024 |  |  | 12/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023179561 |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/30/2022 |  |  |  |  | The complainant reported a 66 year old female patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced multiple episodes of bradycardic events and the physician believed it to be due to a blocked rhythm. epinephrine and dobutamine drips were started. the patient was escalated to a 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25584 | 1220648-2024-12215 | 19495777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495777&pc=OZD | 03/30/2022 | 03/21/2025 |  | 06/09/2024 |  |  | 03/30/2022 | 06/09/2024 |  |  | 10/31/2023 |  | 06/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152485 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/30/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced premature ventricular contractions (pvcs) and abnormal arrhythmias. the patient was successfully weaned and the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25585 | 1220648-2024-12187 | 19493132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19493132&pc=OZD | 09/22/2021 | 03/24/2025 |  | 06/08/2024 |  |  | 09/22/2021 | 06/08/2024 |  |  | 06/30/2023 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108061 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after peeling away the sheath during device placement, an angiogram was performed that revealed no flow distal to the impella repositioning sheath on the right leg. multiple attempts to place an external bypass were unsuccessful. the physician was aware of no doppler pulse of the popliteal artery, and the right leg was much cooler than the left leg. due to this, the impella was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25586 | 1220648-2024-12181 | 19493175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19493175&pc=OZD | 08/02/2021 | 06/07/2024 |  | 06/08/2024 |  |  | 08/02/2021 | 06/08/2024 |  |  | 04/30/2023 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072137 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced peripheral vascular damage that was resolved with balloon tamponade and stenting. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25587 | 1220648-2024-12195 | 19494101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19494101&pc=OZD | 03/16/2022 | 06/07/2024 |  | 06/08/2024 |  |  | 03/16/2022 | 06/08/2024 |  |  | 12/31/2023 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175411 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2022 |  |  |  |  | The user facility reported a patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was revealed that the right coronary artery (rca) was full of thrombus. pneumbra was wired to clear the rca but was unsuccessful as an impella low flow alarm was noted in the aorta. the decision was made to replace the original device with a new impella cp as clotting was still evident. as the clot began to resolve, ballooning was introduced, and flows were stated to have increased. it was reported that there was thrombus left within the patient despite efforts to eradicate. weaning was successful, the device was explanted, and the procedural outcome is the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
25588 | 1220648-2024-12189 | 19494235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19494235&pc=OZD | 03/25/2022 | 03/24/2025 |  | 06/08/2024 |  |  | 03/25/2022 | 06/08/2024 |  |  | 11/30/2023 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165814 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion, the impella got stuck and was difficult to insert, but insertion was completed. subsequently, cardiopulmonary withdrawal occurred. due to difficulty in passing the bioprosthetic valve, there was no aortic regurgitation, but there was concern that the valve leaflets were being pushed, and it looked like impella could be withdrawn, so the patient was subsequently transferred to the icu. the pump was weaned and removed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25589 | 1220648-2024-12194 | 19494473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19494473&pc=OZD | 03/27/2022 | 06/07/2024 |  | 06/08/2024 |  |  | 03/27/2022 | 06/08/2024 |  |  | 10/31/2023 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158391 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/27/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported to be some oozing around sheath. a mattress suture and fem stop was placed. additionally, the patient experienced new onset atrial fibrillation and levophed was restarted. the physician was able to reposition pigtail into free left ventricle space. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25590 | 1220648-2024-12191 | 19494497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19494497&pc=OZD | 03/26/2022 | 03/31/2025 |  | 06/08/2024 |  |  | 03/26/2022 | 06/08/2024 |  |  | 12/31/2023 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172211 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, patient had sepsis. the pump was successfully weaned and explanted and the patient was stable. no further information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the infection/sepsis was not determined since product was not returned and all the necessary information was not provided.
25591 | 1220648-2024-12196 | 19494938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19494938&pc=OZD | 05/09/2022 | 03/31/2025 |  | 06/08/2024 |  |  | 05/09/2022 | 06/08/2024 |  |  |  |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1486611 |  | 0048-0032 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that there was oozing at the right groin, pressure was held with clamp in place for hemostasis. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25592 | 1220648-2024-12197 | 19495066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19495066&pc=OZD | 12/21/2023 | 11/04/2025 |  | 06/08/2024 |  |  | 12/27/2023 | 06/08/2024 |  |  | 09/30/2025 |  | 06/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405266 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombocytopenia | [12 empty] | Other | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/27/2023 |  |  |  |  | The user facility reported a 72-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced heparin induced thrombocytopenia (hit) so heparin was stopped and bicarb was added to the purge. when the hit assay was negative, the patient was restarted on heparin purge. the patient is stable post issue. | Additional information received that the decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The complainant also reported that an implanted impella 5.5 pump triggered a placement signal unreliable alarm during patient support. the alarm was disabled, and the staff was advised on how to monitor the patient via other means. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B5 additional information included. investigation results are still pending. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12197. f6/f8-should have been left blank. g1 reporting contact fax number missing. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump and purge cassette were returned for investigation, however, the catheter side of the pump with the cannula was missing. therefore, the pump was unable to be run in the lab. data logs were returned, however, only for a portion of the case. the first hour after implant and last hour before explant were missing from the logs. based on data log analysis of the partial logs, it was seen that there were placement signal not reliable (alarms #130 and #1036) and placement signal low alarms. there were placement signal spikes throughout the case. no motor current spikes were seen, with values consistent with p-level changes. gain increased slightly while light decreased slightly. lamp and contrast were stable. snr was low and dropped close to zero towards the end of the case. clinical details state that the patient showed signs of heparin induced thrombocytopenia so the purge was changed to sodium bicarbonate. placement signal unreliable alarms were seen during the case. patient was on crrt support during the case as well. patient had multiple prbc transfusions during support and their platelet count continued to decrease. placement signal alarms were continuous and echo was done to confirm placement of the device. optical signal issue: the root cause of the optical signal issue was unable to be definitively determined as insufficient product was returned for analysis, data logs were inconclusive and limited clinical information was provided thrombocytopenia: the root cause of the thrombocytopenia was not determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 and b2 updated. b5 updated. h1 and h2 updated. h3 updated. h6 updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
25593 | 1220648-2024-12019 | 19481618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481618&pc=OZD | 05/27/2024 | 07/23/2024 |  | 06/07/2024 |  |  | 05/27/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing st-segment elevation myocardial infarction was implanted with an impella cp pump for mechanical circulatory support. two days on support, left ventricle thrombus was noted on echo. consequently, the impella cp pump was explanted. the patient was noted to be in stable condition. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.7 revised as information was inadvertently omitted from the previously submitted manufacturer device report 1220648-2024-12019. h.6 add code 925 since initial manufacturer device report 1220648-2024-12019 submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report 1220648-2024-12019 in accordance with updated procedures. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
25594 | 1220648-2024-12052 | 19481622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481622&pc=OZD | 05/30/2024 | 11/26/2024 |  | 06/07/2024 |  |  | 05/30/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132456 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The complainant reported a 68-year-old male patient with a high risk of percutaneous coronary intervention was being prepped for an impella cp implant for mechanical circulatory support. the complainant reported that the staff received an optical sensor error and switch to backup controller from their automated impella controller, which was then replaced. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) loss of opm data has been completed. the console log was reviewed and confirmed the reported failure mode given there was a controller error alarm (placement signal not reliable (opm invalid signal, optical pump connected)- alarm issued) (imcgui: event set: #1036) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench times of the subsequent suspension of optical bench algorithms, the optical parameters were significantly lower than normal values. the console was returned for evaluation. the optical system was tested and all optical parameters were confirmed to be within specification. the reported complaint was unable to be reproduced. no issues were found with the console. added- d9 device return date corrections to the initial mfr report: d4-udi number f6/f8 should have been left blank.
25595 | 1220648-2024-12043 | 19481659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481659&pc=OZD | 05/29/2024 | 06/19/2024 |  | 06/07/2024 |  |  | 05/29/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a battery failure during an in-service appointment. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) red alarm has been completed. the console was returned and the battery failure issue was able to be reproduced. the data logs were reviewed and the reported battery failure issue was confirmed. there were multiple instances of battery failure alarms (transport connection blown/missing) (event set #360), battery level low (50%) (event set #23), and an instance of battery failure alarm (low voltage) [v < 12v] (event set #350), found on the reported occurrence date. results of running batttest on telnet revealed that vcell1 in battery 2 had a voltage that was significantly less than the rest of the cells in the battery (see attached for batttest results). the reported battery failure issue was determined to be caused by internal battery cell imbalance (found in vcell1 of battery 2). d4-corrected model number
25596 | 1220648-2024-12041 | 19481784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481784&pc=OZD | 05/28/2024 | 07/15/2024 |  | 06/07/2024 |  |  | 05/28/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | Inadequate User Interface | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, blood was actively oozing from the red plug of the impella 5.5 pump. it was noted that the surgeon had punctured the catheter with a suture needle resulting in blood entering the catheter and traveling up to the red plug. due to the noted issue, the decision was made to remove the pump. there was no patient harm resulting from the failure. | [5 empty] | The root cause of the product damage (hole in catheter) issue was not determined since product was not returned for investigation. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25597 | 1220648-2024-12013 | 19481919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481919&pc=OZD | 05/08/2024 | 07/09/2024 |  | 06/07/2024 |  |  | 05/08/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492942 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that a patient on impella cp support exhibited oozing at the access site. manual pressure was placed, and the patient was transfused with two units of packed red blood cells. | The decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. updated b5 to reflect patient outcome and mso statement. f6/f8-should have been left blank in the initial report g1-corrected manufacturing site name and address,
25598 | 1220648-2024-12072 | 19482302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19482302&pc=OZD | 05/15/2024 | 07/11/2024 |  | 06/07/2024 |  |  | 05/15/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467794 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and while on support, there was an anti-contamination sleeve tear distal and proximal. the tear was covered in betadine and a tegaderm. there was no report of harm to the patient. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported product damage has been completed. product was returned for investigation. the sterile sleeve was seen torn from the distal end of the catheter towards the red handle. on the proximal end, the sterile sleeve was completely torn from the catheter anchor, with the anchor itself missing from the returned product. in conclusion, it was determined that the cause of the sterile sleeve damage was excessive force when handling the device. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: cautions. ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ added-h6 impact code 4641 d4-corrected model number.
25599 | 1220648-2024-12089 | 19483015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483015&pc=OZD | 05/31/2024 | 01/31/2025 |  | 06/07/2024 |  |  | 05/31/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1687316 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The us complainant had a console with a battery/power/rebooting issue at the medical center. the console was removed from service and the ifu (instructions for use) will be followed and a backup aic (automated impella controller) used for patient use. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) a/c power issues has been completed. logs from the console were no retrievable due to corruption of the compact flash cards. however, review of the constellation logs was not conclusive. the console was returned for evaluation. the failure was reproducible upon boot up in the lab. swapping the compact flash cards with a known good one fixed the issue. during field service repair, misaligned lcd pins were discovered. the root cause for the constant rebooting of the console was due to the corruption of the compact flash cards of the epc preventing a successful initial boot up of the console.
25600 | 1220648-2024-12107 | 19483086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483086&pc=OZD | 08/03/2021 | 03/25/2025 |  | 06/07/2024 |  |  | 08/03/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089175 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2021 |  |  |  |  | The complainant reported the patient was on impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the patient experienced ventricular tachycardia that was resolved with defibrillation and anti-arrhythmia medication. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25601 | 1220648-2024-12117 | 19483330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483330&pc=OZD | 05/10/2024 | 07/15/2024 |  | 06/07/2024 |  |  | 05/10/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458127 |  | 10000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Diminished Pulse Pressure | Thrombosis/Thrombus | [10 empty] | Disability | Required Intervention | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient suffered pulseless electrical activity twice. an echocardiogram was performed and a thrombus in the left ventricle was noted. a fem-stop was placed to address the access site bleeding due to the patient moving, and multiple units of packed red blood cells and fresh frozen plasma were transfused. the patient expired while on support. upon abiomed's medical safety officer review, there is not enough information to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the reported issues has been completed. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus and the ventricular fibrillation/ventricular tachycardia/atrial fibrillation/atrial tachycardia was not determined. the root cause of the access site bleeding issue was most likely patient movement. a femostop was placed to address access site bleeding due to the patient moving as per the clinical description. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12117 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12117 was submitted.
25602 | 1220648-2024-12119 | 19483430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483430&pc=OZD | 07/13/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085745 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was concerns for left ventricle clot with embolization. as a result, the impella was removed and replaced with intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | Thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
25603 | 1220648-2024-12118 | 19483440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483440&pc=OZD | 05/11/2024 | 06/20/2024 |  | 06/07/2024 |  |  | 05/11/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported one day following implant, no distal pulse was noted in the impella leg. an arterial ultrasound was requested, but no plan was initiated to address the condition. three days later, a computed tomography scan was performed, and it was noted that the patient had suffered a cerebral infarction. the patient's overall condition continued to decline, and the patient subsequently passed away. per abiomed's medical safety officer review, there is no association between impella use and the patient's outcome. the patient's peri-procedural condition was poor. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported limb ischemia and stroke has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and stroke was not determined.
25604 | 1220648-2024-12122 | 19483729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483729&pc=OZD | 05/08/2024 | 09/19/2024 |  | 06/07/2024 |  |  | 05/15/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691229 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | No patient involvement: the complainant reported that there was a broken part in the purge cassette area. there was no reported patient harm. | [5 empty] | The investigation for automated impella controller (aic) - purge disc/flag issues has been completed. the console was tested after arriving and the purge disc retainer was confirmed to be broken (broken blue disc retainer). the cause of the issue was a defective component. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12122: d4 model number. f6/f8-should have been left blank . g1 manufacturing site name and address. | The impella device was received from the customer on 04-jun-2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
25605 | 1220648-2024-12123 | 19483753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483753&pc=OZD | 05/16/2024 | 07/18/2024 |  | 06/07/2024 |  |  | 05/16/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449799 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the 71-year-old male patient experienced cold foot on impella sided leg. the patient's acts were between 160/180. the patient's condition was poor and the ejection fraction was at 10%. therapy was noted to be continued until relatives said goodbye, and care was withdrawn one day post implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- h6 component code 925 d4-corrected model number. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and number.
25606 | 1220648-2024-12124 | 19483776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483776&pc=OZD | 04/10/2024 | 06/11/2024 |  | 06/07/2024 |  |  | 05/20/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457192 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The user facility reported a 57-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that an impella cp was successfully implanted in a patient and developed a high purge pressure alarm. sodium bicarb was flowing through the cassette. the cassette was changed to resolve the alarm. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. result of investigation: product was not returned for investigation. data logs were returned and reviewed. there was rise in purge pressure for about 5 minutes before the alarm was seen. there was also a motor current spike corresponding to the rise in purge pressure. the pump was plowing at p2 level at that time the purge solution was of dextrose solution and bicarb. saturated flow was also noticed during the spike. per the clinical information provided and data log review, the root cause of high purge pressure issue was most likely biomaterial. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | D4-updated model number. d6- implant and explant date. g1-corrected mfr site name and address. g2-should not have selected foreign h6-type of investigation 10 is changed to 4114. the device was not returned.
25607 | 1220648-2024-12126 | 19483814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483814&pc=OZD | 05/30/2024 | 07/24/2024 |  | 06/07/2024 |  |  | 05/30/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1112622 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The complainant reported an (b)(6) male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported that the associated automated impella controller (aic) had a console failure. the aic was exchanged. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The console was received for investigation. console logs from the reported day of event are consistent with complaint. during testing of the console, and particularly of the optical bench, it¿s been observed that the output of the ccd (analog output of the bench) indicated possible defect of ccd. when monitored over several hours, the output of the ccd was observed to be spontaneously changing, without any interaction with the console. this phenomenon likely affected placement signal calculation, and the observed likely defect of ccd could have additionally affected optical bench internal algorithms. the cause of the automated impella controller (aic) issue was a defective optical bench. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12126: d3 manufacturer name, city and state missing. f6/f8-should have been left blank. g1 reporting contact fax number missing.
25608 | 1220648-2024-12127 | 19483828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483828&pc=OZD | 05/30/2024 | 07/17/2024 |  | 06/07/2024 |  |  | 05/30/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025471283 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that no flow was seen through the introducer during impella cp implant. both the pump and peel-away were left in the patient and a fem-fem external bypass was placed to address the flow issue. support with the implanted pump continued as planned following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4-corrected model number f6/f8-should have been left blank in the initial report. g4-added pma number.
25609 | 1220648-2024-12129 | 19483839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483839&pc=OZD | 03/24/2024 | 07/16/2024 |  | 06/07/2024 |  |  | 05/14/2024 | 06/07/2024 |  |  | 11/30/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435561 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support post coronary artery bypass grafting. abiomed received a notification via abiomed's long-term outcome and quality indication impella registry that the patient endured acute renal failure with a "possible" relationship to the impella cp device used: ¿began continuous renal replacement therapy then transitioned into hemodialysis with perma-cath placement." it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure could not be determined. f6/f8 should have been left blank in the initial report.
25610 | 1220648-2024-12134 | 19483938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483938&pc=OZD | 07/13/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the impella pump was being repositioned and came out of position. the decision was made to explant the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25611 | 1220648-2024-12135 | 19483945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483945&pc=OZD | 06/03/2024 | 07/30/2024 |  | 06/07/2024 |  |  | 06/03/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474207 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient experiencing acute myocardial infarction/cardiogenic shock implanted with an impella cp pump was escalated to an impella 5.5 pump. during impella 5.5 pump support the left ventricle wave forms were vastly negative and diminished in the bottom of the screen, which was noted to have immediately started upon the impella 5.5 pump start. proper position was verified in multiple angles on transesophageal echocardiogram. however, despite correct position, diminished wave forms persisted. the pump was explanted. there was no report of harm to the patient. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Impella 5.5 pump was returned. visually inspected and no physical damages, kinks or abnormalities found. optical bench 5s parameters were verified to be in normal range and light continuity test passed without any issues. no placement signal issue alarm reproduced. borescope view of cannula obtained, and biomaterial was observed inside the cannula. low flow reproduced with biomaterial during testing. no other issues were found. data logs were reviewed, and several suction alarms and low flow observed during most of the case. saturated flows were found at p-4 for a duration of 1 hour post 1.5 hours of pump support. remaining whole case, it was found to be low flows and suction alarms. no optical signal issues were found. the cause of the low pump flow issue was due to biomaterial per return product investigation and log review. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12135: d4 model number.
25612 | 1220648-2024-12136 | 19483954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483954&pc=OZD | 06/03/2024 | 08/11/2024 |  | 06/07/2024 |  |  | 06/03/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025480083 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. when placing the impella 5.5, the catheter kinked. a new impella 5.5 was inserted. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue was determined to be use issue because a new pump was placed later. a.1 and b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12136. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12136 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12136 was submitted.
25613 | 1220648-2024-12138 | 19483992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19483992&pc=OZD | 07/13/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 01/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022049255 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during impella explant, it was noted that the 4f sheath was left in place to be able to be pulled in the intensive care unit due to slight bleeding at the groin. the patient¿s condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25614 | 1220648-2024-12140 | 19484059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484059&pc=OZD | 09/25/2021 | 06/07/2024 |  | 06/07/2024 |  |  | 09/25/2021 | 06/07/2024 |  |  | 04/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072640 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/25/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was addressed with reduction of heparin dosage. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25615 | 1220648-2024-12141 | 19484104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484104&pc=OZD | 05/26/2024 | 08/01/2024 |  | 06/07/2024 |  |  | 05/26/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025471193 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | The user facility reported a 37-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that the sterile sleeve had a tear in the distal end. the tear was covered with betadine and a tegaderm. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage (sterile sleeve damage) has been completed. the root cause of the product damage (sterile sleeve) was not determined as no product was returned. b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12141. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12141 and should not have been. h.6 code 2199 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12141. a new code was added.
25616 | 1220648-2024-12144 | 19484163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484163&pc=OZD | 04/28/2024 | 07/09/2024 |  | 06/07/2024 |  |  | 05/10/2024 | 06/07/2024 |  |  | 11/30/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024433884 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported an 84-year-old female patient with pre-operation placement was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that the patient was on support for a cardiac rupture for 10 days before they expired. the patient also experienced a neurological cardiac infarction event. the time of the event is unknown, as medical staff started performing cardiopulmonary resuscitation. there is no association between the impella use and the patient¿s outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. f6/f8 should have been left blank in the initial report.
25617 | 1220648-2024-12145 | 19484208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484208&pc=OZD | 07/13/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 04/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070126 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was noted. no additional information related to the resolution of hemolysis was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25618 | 1220648-2024-12147 | 19484240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484240&pc=OZD | 05/30/2024 | 10/17/2025 |  | 06/07/2024 |  |  | 05/30/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1668177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) gave a message regarding optical sensor not working when it was connected to the impella cp pump. the staff tried to restart, disconnected/reconnected, the aic, and the same message displayed. also, an option to ¿optical bench service¿ appeared. consequently, the aic was replaced. there was no report of harm to the unknown age patient. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis imc logs confirmed that on the reported complaint date ((b)(6) 2024), during case start, there was an optical sensor system error that was triggered and displayed to the user. both preceding and following the reported complaint date, and across multiple boots of the console, there were multiple occurrences of suspensions of the optical bench algorithms. review of rt or lt logs did not reveal any information relevant to the reported complaint. device analysis console (b)(6) was sent back to field service for further investigation. the reported complaint could not be reproduced. the console¿s optical system was tested and confirmed to be working correctly. the reported complaint was determined to most likely have been caused by an intermittent malfunction of the optical bench. out of precaution, before returning the console to the customer, field service was instructed to replace the optical bench and the ribbon cable connecting the optical bench to the 4-in-1. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d1 common device name updated. d4 model number updated. g1 reporting contact email updated.
25619 | 1220648-2024-12149 | 19484429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484429&pc=OZD | 05/30/2024 | 07/18/2024 |  | 06/07/2024 |  |  | 05/30/2024 | 06/07/2024 |  |  | 11/30/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438752 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The complainant reported a 70-year-old male patient acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient experienced bleeding at the access site during the introduction of the repositioning sheath. medical staff held manual pressure and modified the purse string suture. the patient continued to have decreased platelets and hemoglobin despite having been administered four units of blood, overnight. medical staff withheld anticoagulants for several days, and, after the impella was removed, a large clot was noted in the inlet. the patient survived the event. | [5 empty] | The investigation for access site bleeding, hemolysis, and thrombus has been completed. data review: data logs showed pump ran without issue during support. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to patient condition (obesity and tortuous vessel). the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. h5 added yes h8 added initial use f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing h6 health effect - impact code omit 4624 from initial report | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25620 | 1220648-2024-12151 | 19484513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484513&pc=OZD | 07/14/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/14/2021 | 06/07/2024 |  |  | 04/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped which was resolved by replacing the catheter. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. section: troubleshooting the purge system¿ unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25621 | 1220648-2024-12155 | 19484538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484538&pc=OZD | 03/18/2022 | 03/21/2025 |  | 06/07/2024 |  |  | 03/18/2022 | 06/07/2024 |  |  | 07/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, an intravascular ultrasound was performed an identified a mid-left anterior descending (lad) artery dissection. an attempt was made to wean the impella, but the patient's anatomy could not tolerate it. the patient became hypotensive with minimal pulsatility. the decision was made to leave the impella in overnight. it was then found that the patient was bleeding from multiple access locations, though there was no indication as to where the bleed originated. the device was explanted, and the patient survived the procedures. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25622 | 1220648-2024-12158 | 19484606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484606&pc=OZD | 03/19/2022 | 03/21/2025 |  | 06/07/2024 |  |  | 03/19/2022 | 06/07/2024 |  |  | 12/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169892 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced distal limb ischemia. the device was explanted. it was reported that the patient was stable. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25623 | 1220648-2024-12156 | 19484607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484607&pc=OZD | 07/22/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/22/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089744 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding from multiple sites, no pulsatility, red urine, and purge pressure alarm was noted. systemic heparin was stopped, and 1l of volume given. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25624 | 1220648-2024-12160 | 19484727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484727&pc=OZD | 03/19/2022 | 03/21/2025 |  | 06/07/2024 |  |  | 03/19/2022 | 06/07/2024 |  |  | 08/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the site continued oozing with two perclose applications. an 8fr angioseal was placed with good results. additionally, it was reported that the patient also lost distal pulsatility, indicating limb ischemia; therefore, the impella was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25625 | 1220648-2024-12065 | 19484793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484793&pc=OZD | 05/24/2024 | 07/18/2024 |  | 06/07/2024 |  |  | 05/24/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485986 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while the patient was on support, they went into ventricular tachycardia/torsades and the patient was shocked. the impella was repositioned. additionally, the patient had a large hematoma at the access site. the team decided to replace the impella to the opposite femoral artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported ventricular fibrillation (vf) and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf and access site bleed could not be determined as the device was not returned nor sufficient clinical details. added- h6 component code 925 f6/f8 should have been left blank in the initial report.
25626 | 1220648-2024-12165 | 19484955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484955&pc=OZD | 03/19/2022 | 03/21/2025 |  | 06/07/2024 |  |  | 03/19/2022 | 06/07/2024 |  |  | 12/12/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced compartment syndrome of the leg, along with bleeding and continuous transfusions. heparin was removed from the purge cassette. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25627 | 1220648-2024-12170 | 19484996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19484996&pc=OZD | 03/09/2024 | 07/16/2024 |  | 06/07/2024 |  |  | 05/15/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447911 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ongoing hemolysis with a "possible" relationship to the impella device used: ¿3/9/24 hgb 9.4 this morning, down from 11.1 yesterday, was 14.9 on admission on 3/7. ldh elevated 2565, suspect pt likely having a component of hemolysis from impella. remains on impella with stable maps. concern for worsening anemia with elevated ldh of 2500- concern for ongoing hemolysis. 3/11/24 hgb 7.3 this morning. ldh elevated over 2500, haptoglobin <30, total bilirubin 1.2. concern for ongoing hemolysis from impella. 3/13/24 hgb 7.3 -> 7.1 3/12. ldh elevated over 2500, haptoglobin <30, total bilirubin 1.2. concern for ongoing hemolysis from impella. s/p 1u prbc on 3/11 up to hgb 7.3¿. the patient outcome is not known at this time. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. f6/f8 should have been left blank in the initial report. h5 changed to yes.
25628 | 1220648-2024-12171 | 19485033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485033&pc=OZD | 03/24/2022 | 03/21/2025 |  | 06/07/2024 |  |  | 03/24/2022 | 06/07/2024 |  |  | 12/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171259 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, oozing was noted at the access site and a cardiac device infection was observed. the pressure dressing was changed and angle matching techniques were utilized to resolve this issue. it was reported that a hematoma formed at the insertion site and the patient¿s blood pressure decreased. therefore, systematic heparin was halted, manual pressure was held for one hour and the patient was provided a fluid bolus. weaning was successful, the device was explanted, and the procedural outcome is the patient survived surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25629 | 1220648-2024-12173 | 19485074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485074&pc=OZD | 04/30/2024 | 06/21/2024 |  | 06/07/2024 |  |  | 05/15/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471189 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation of care from an impella cp device. abiomed complaints received a notification via abiomed's long-term outcome and quality indicator impella registry reporting the patient endured non-sustained ventricular tachycardia (recurring) with a "definite" relationship to the impella device used. ¿the patient developed recurring runs of non-sustained ventricular tachycardia and singular ventricular extra systolic beats. antiarrhythmic drugs started¿. the patient outcome is not known at this time. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (non-sustained ventricular tachycardia and singular ventricular extrasystole beats) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
25630 | 1220648-2024-12178 | 19485200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485200&pc=OZD | 05/10/2024 | 07/09/2024 |  | 06/07/2024 |  |  | 05/10/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025477599 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported a patient of unknown race, age and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had an ischemic right leg. the patient was placed on femoral-femoral bypass and the ischemia was resolved. the patient expired two days later. abiomed does not have enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report. g1-corrected mfr site address line 1 and remove mfr site address line 2
25631 | 1220648-2024-12179 | 19485260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485260&pc=OZD | 08/25/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 08/25/2021 | 06/07/2024 |  |  | 06/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was found to have cold feet and poor distal perfusion. the patient was later successfully weaned and the device explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25632 | 1220648-2024-12182 | 19485669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485669&pc=OZD | 08/26/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 08/26/2021 | 06/07/2024 |  |  | 06/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had cold extremities and narrow pulse pressure. it is unknown if/how this condition was resolved. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25633 | 1220648-2024-12184 | 19485910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485910&pc=OZD | 08/02/2021 | 12/05/2024 |  | 06/07/2024 |  |  | 08/02/2021 | 06/07/2024 |  |  | 01/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022048546 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Renal Impairment | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was suspected of having a middle cerebral artery ischemic stroke. an electroencephalogram was performed which identified moderate cerebral dysfunction with no epileptiform discharges detected. a repeat brain scan was performed that noted an embolic stroke. it is unclear how hemostasis was achieved. weaning was successful, the device was explanted. | It was additionally reported on that the patient had an acute kidney injury. it is unknown if the use of impella was related to the acute kidney injury. no information was provided if any treatment was performed for the acute kidney injury. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported stroke/cva and renal failure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the stroke/cva and renal failure was unable to be determined as insufficient clinical details were provided. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ additional information was provided in sections b5 and h6 for acute kidney injury. this report is being filed as part of a retrospective review of historical records.
25634 | 1220648-2024-12078 | 19485914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19485914&pc=OZD | 07/25/2021 | 03/21/2025 |  | 06/07/2024 |  |  | 07/25/2021 | 06/07/2024 |  |  | 01/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060041 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2021 |  |  |  |  | The complainant reported a patient in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the heart was ¿asleep¿ and the cath cannot be placed in a good position. p6 caused a lot of premature ventricular contractions, and the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25635 | 1220648-2024-12150 | 19486430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19486430&pc=PYX | 05/11/2024 | 09/19/2025 |  | 06/07/2024 |  |  | 05/11/2024 | 06/07/2024 |  |  | 01/31/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024322357 |  | 0046-0035 |  | No answer provided |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/11/2024 |  |  |  |  | The user facility reported a 77-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. it was reported the patient was on bipella support and the access site in the right groin was found to have saturated dressing. the dressing was taken down, a flow stasis device was applied to the mattress sutures, sutures were tightened down, and dressing was applied. the bleeding resolved. a placement signal alarm, preceded by suction, also sounded. flow calculations were disabled, and the motor current remained unchanged. the patient eventually expired. upon review by abiomed's medical safety officer, there is no association between impella use and the patient¿s outcome. | Us complainant reported the patient was on bipella support with an impella rp. while on rp support, the impella rp's access site in the right groin was noted to have saturated dressing. the dressing was taken down, a flow stasis device was applied to the mattress sutures, sutures were tightened down and dressing was applied. the bleeding resolved. additionally, there was a placement signal alarm noted. it was preceded with suction on rp. flow calculations were disabled and the motor current remained unchanged. patient expired while on support. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the access site bleed has been completed. no product was returned for evaluation. clinical details noted oozing from rp access site, dressings were removed, flow stasis device was used on mattress sutures and bleeding was resolved. the root cause of the access site bleeding - major cannot be determined. f6/f8 should have been left blank in the initial report. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. data logs were reviewed and placement signal not reliable (psnr) (dp signal out of range) alarm occurred multiple times towards the end of the case. cvp and dp sensor both went out of range multiple times throughout end of case with mean flow values spiking. motor current and optical signal remained unchanged throughout case. clinical details noted that patient had suction prior to psnr alarm occurring with intermittent placement signals then none being shown on the automated impella controller (aic). additionally, flow calculations were disabled but motor current remained unchanged. clinical details also noted oozing from rp access site, dressings were removed, flow stasis device was used on mattress sutures and bleeding was resolved. the root cause of the placement signal issue was most likely due to sensor drift. the root cause of the access site bleeding - major cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
25636 | 1220648-2024-12185 | 19487030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19487030&pc=OZD | 06/03/2024 | 08/12/2024 |  | 06/07/2024 |  |  | 06/03/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441237 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding. manual pressure and a femstop was placed. one unit of blood products was administered due to a drop in central venous pressure and pump suction events. the patient remained on support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. the data logs confirmed the failure mode and clinical narrative. logs showed after the pump started about 10 minutes. low flow occurred that triggered auto suction response and suction alarms. this event remained for about 24 hours and then was resolved. based on clinical description, the root cause of low flow was most likely due to patient condition as the central venous pressure was low during support. the root cause of the access site bleeding was not determined since neither the product nor sufficient clinical information were provided. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12185 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12185 was submitted. h.6 code 4641 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-12185.
25637 | 1220648-2024-12192 | 19487132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19487132&pc=OZD | 03/27/2022 | 06/07/2024 |  | 06/07/2024 |  |  | 03/27/2022 | 06/07/2024 |  |  | 04/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070261 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/27/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced persistent access site bleeding. systemic heparin was discontinued and blood products were administered. additionally, it was reported that the pump experienced persistent ¿impella in aorta¿ alarms. the impella was repositioned. alarms resolved, and waveforms improved with volume. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25638 | 1220648-2024-12188 | 19487241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19487241&pc=OZD | 01/23/2024 | 06/24/2024 |  | 06/07/2024 |  |  | 05/14/2024 | 06/07/2024 |  |  | 11/30/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424941 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | We teceived a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured bleeding (hematoma) with a "possible" relationship to the impella device used: ¿right groin hematoma starting after impella procedure on (b)(6). rfa site hematoma was treated with prn manual compression and transfusions ((b)(6)), progressively worsened getting bruised and firm. impella removed (b)(6) and pts hemoglobin persistently dropping while on heparin drip. pt underwent hematoma evacuation with r groin washout on (b)(6) and there was bleeding from both arteriotomies in the common femoral artery. these were reinforced with 5 0 prolene sutures. following this groin less firm and healing. hemoglobin trend: (b)(6) 14.3, (b)(6) 12.9, (b)(6) 10.0, (b)(6) 9.6, (b)(6) 10.4, (b)(6) 9.9, (b)(6) 9.5, (b)(6) 6.7, (b)(6) 6.3 > 10.5¿. the patient recovered with sequelae. we also received a notification via abiomed¿s long-term outcome and quality indicator impella registry that the patient that endured insertion site infection with a "possible" relationship to the impella device used: ¿went to operating theater on (b)(6) 2024 operative findings at the time of (b)(6) 2024 were significant for necrotic and infected soft tissue right groin. these specimens were sent for culture, which showed evidence of gram-negative rods. on inspection in the operating theater, the previous right femoral artery repair appeared intact. the necrotic skin and subcutaneous tissue was sharply debrided to healthy bleeding tissue. the wound measured 7 x 3 x 2 cm. patient was started on iv cefepime and vancomycin, and will complete a 4 week course at ltac. finished antibiotic course (b)(6)". the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the access site bleeding ¿ major and the infection/sepsis issues has been completed since the original report was submitted. the product was not returned for investigation. the root cause of access site bleeding is determined to use issue because the issue was resolved, and the site was healing after adding prolene sutures at the site. because this evidence has not been provided, the root cause of the infection cannot be determined.
25639 | 1220648-2024-12193 | 19487470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19487470&pc=OZD | 03/27/2022 | 03/21/2025 |  | 06/07/2024 |  |  | 03/27/2022 | 06/07/2024 |  |  | 01/31/2024 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181653 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient was brought to lab for positional issues with cp. echo was done and pigtail was noted to be under a pap muscle. the physician attempted to reposition under fluoroscopy and echo. the decision was made to remove impella and peel away sheath and replace with another impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. h1 type of reportable event: corrected from serious injury to product malfunction.
25640 | 1220648-2024-12152 | 19487801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19487801&pc=OZD | 03/16/2022 | 04/22/2025 |  | 06/07/2024 |  |  | 03/16/2022 | 06/07/2024 |  |  | 01/31/2024 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172837 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that bleeding was observed at the access site and at the extracorporeal membrane oxygenation cannulas. heparin was reduced in the purge. the patient continued on support until the patient was successfully weaned and bridged to transplant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25641 | 1220648-2024-12161 | 19487957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19487957&pc=OZD | 08/18/2021 | 03/20/2025 |  | 06/07/2024 |  |  | 08/18/2021 | 06/07/2024 |  |  | 06/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103684 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2021 |  |  |  |  | The user facility reported a 74-year-old female implanted with an impella cp for mechanical circulatory support due to acute systolic heart failure and severe mitral regurgitation. it was reported that the impella cp device unexpectedly alarmed motor current high, pump stop. the device was able to be started at p-2, but the motor current was very high and would stop after just a few moments. the patient was decompensating and the impella cp was withdrawn into the aorta and drips were titrated up to stabilize patient. the pump was then replaced due to patient's hemodynamic compromise. in addition, there was brisk bleeding noted around 11 french sheath and pressure was held. a new impella cp device was implanted, however, bleeding was still noted around the sheath. the patient was successfully weaned and the device explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ pump stop section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25642 | 1220648-2024-12131 | 19488002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19488002&pc=OZD | 07/13/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 02/28/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056188 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, left anterior descending artery dissection was observed, and the medical staff could not re-access. the patient was transported to a different medical facility for additional treatment. the patient was taken for a scan to evaluate neurological status. it was stated that the patient¿s neurological assessment was consistent with cerebral vascular accident. therefore, the impella pump was weaned and explanted the following day and escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related events are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the stroke/ cva was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25643 | 1220648-2024-12172 | 19488173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19488173&pc=OZD | 04/28/2024 | 07/17/2024 |  | 06/07/2024 |  |  | 05/15/2024 | 06/07/2024 |  |  | 09/30/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444562 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used: ¿hgb 9.4 on arrival to the hospital. 11 hours later, post-procedure, hgb dropped to 7.0 on the night of 12/8. anemia secondary to hemolysis from impella. patient transfused with 3 units prbc (packed red blood cells) and impella ultimately explanted. by 12/10 hemoglobin is improved and stable and continues to be monitored¿. the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
25644 | 1220648-2024-12125 | 19488796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19488796&pc=OZD | 05/29/2024 | 07/15/2024 |  | 06/07/2024 |  |  | 05/29/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1183259 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed a battery temperature high alarm when unplugged. the aic was immediately plugged back in to resolve the noted issue. there was no patient harm resulting from the failure. | [5 empty] | The investigation for the aic battery failure has been completed. data logs were reviewed and the reported battery temp high issue was confirmed. console logs from the complaint date revealed the ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. this battery temperature spike was recorded in the ltlog power data. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc. the console was returned. the console¿s batteries and power management system were tested, and no issues were observed. battery temp high alarm was due to a momentary communication glitch between the batteries and the pbm, where a singe sample from battery data (in this case value of battery temperature) was corrupt. added-d9 device return date d4-corrected model number. f6/f8 should have been left blank in the initial report. h6 component code changed to 765. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25645 | 1220648-2024-12128 | 19488819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19488819&pc=OZD | 05/27/2024 | 07/23/2024 |  | 06/07/2024 |  |  | 05/31/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478446 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the sterile sleeve was ripped from the hub during impella repositioning. the catheter was cleaned with chloraprep and tegaderms were used to reinforce. | [5 empty] | The investigation for sterile sleeve damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12128: f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25646 | 1220648-2024-12133 | 19488873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19488873&pc=OZD | 06/01/2024 | 08/12/2024 |  | 06/07/2024 |  |  | 06/03/2024 | 06/07/2024 |  |  | 10/31/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024431351 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/03/2024 |  |  |  |  | The user facility reported an 80-year-old female patient was implanted with an impella cp device for mechanical circulatory support. it was reported that on day 2 of support there was a purge leak from the sidearm. there was no lasting harm or injury and to date the pump has not yet been explanted or replaced. | [5 empty] | The investigation for the reported purge leak has been completed. pump was returned for evaluation. upon visual inspection, no abnormalities were noted. pump was purged for approximately 10 minutes and no leaks were observed. additionally, purge flow and pressure were within a normal range throughout purging. data logs were reviewed and no changes in purge pressure and purge flow were noted at all throughout support. clinical details noted that there was no increase in purge flow rate or decrease in purge pressure when leak was observed. no problem was found with the pump during the case. the pump functioned to specification in the field and in the lab. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ added- d9 device return date f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25647 | 1220648-2024-12186 | 19489044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489044&pc=OZD | 08/04/2021 | 05/14/2025 |  | 06/07/2024 |  |  | 08/04/2021 | 06/07/2024 |  |  | 07/31/2022 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1490580 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was a blood clot in the 16fr sheath and in the pigtail area. a new 16fr sheath was used. in addition, there were frequent suction alarms. the device was removed and replaced with a new device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25648 | 1220648-2024-12142 | 19489048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489048&pc=OZD | 03/16/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/16/2022 | 06/07/2024 |  |  | 11/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that while advancing the impella cp device over the 0.018 wire, the wire position in the left ventricle was lost. the pump was removed to rewire across the aortic valve. the impella cp device was inserted again successfully and started with flows in the expected range. it was also reported the patient experienced hemolysis. it was decided to remove the pump. the impella cp device was replaced with an intra-aortic balloon pump (iabp). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25649 | 1220648-2024-12162 | 19489096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489096&pc=OZD | 07/22/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/22/2021 | 06/07/2024 |  |  | 02/28/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056043 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. during implantation, the physician perforated the diagonal artery with the guide wire and the patient went into cardiac tamponade. flows decreased at the time. pericardiocentesis was performed and flows improved. the patient was stabilized and the impella cp device was explanted at the end of the case. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25650 | 1220648-2024-12177 | 19489267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489267&pc=OZD | 06/03/2024 | 08/12/2024 |  | 06/07/2024 |  |  | 06/03/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025471837 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella 5.5 device due to the inability to pass through the vascular system. the impella 5.5 pump met resistance just beyond the anastomosis during attempted delivery for patient support. upon attempts to advance it further, the shaft kinked. due to the noted damage, the impella 5.5 pump was removed and an impella cp device was subsequently implanted for patient support. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of small vessels of the patient. added- h6 med dev prob code 2889 d4-corrected model number. added- d6a/d6b.
25651 | 1220648-2024-12139 | 19489271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489271&pc=OZD | 05/16/2024 | 07/23/2024 |  | 06/07/2024 |  |  | 05/17/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488775 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when the impella cp wire was to be retracted after the pump was placed, it became stuck, and resistance occurred. after careful turning several times, the wire could be removed without any problems. so far, the pump has been running without any problems. the patient was stable. | [5 empty] | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use issue since the guidewire could not be removed with the proper technique. section: warnings and cautions: warnings. to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ d4 model number f6/f8-should have been left blank in the initial report. g1 manufacturing site name and address. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25652 | 1220648-2024-12148 | 19489273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489273&pc=OZD | 07/13/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083458 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had oozing from the access site and the patient was on integrlin drip. the dressing was re-secured. additionally, the patient¿s leg was warm to touch but pulses were lost since levophed was added for blood pressure. heparin was stopped systemically and in purge. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related events are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25653 | 1220648-2024-12130 | 19489464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489464&pc=OZD | 05/31/2024 | 08/08/2024 |  | 06/07/2024 |  |  | 06/01/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025477622 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2024 |  |  |  |  | The user facility reported a male patient of an unknown age was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed hemolysis during impella 5.5 support. it was noted that flow rates dropped from 4.0 lpm to 2.6 lpm following a scheduled thoracotomy and could not be increased. in response to the noted hemolysis, the support level was lowered, haptoglobin was administered, and the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12130: f6/f8-should have been left blank.
25654 | 1220648-2024-12132 | 19489589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489589&pc=OZD | 06/01/2024 | 07/23/2024 |  | 06/07/2024 |  |  | 06/01/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella cp pump for support to an unknown age patient experiencing acute myocardial infarction/cardiogenic shock. there was difficulty placing the impella cp pump across the valve and inability to maintain proper position. the decision was made to remove the impella cp pump and place an intra-aortic balloon pump. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance." | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the failure mode was determined to be patient condition as the clinical notes do not state any deviations from best practices. the clinical notes clearly state that the difficulty was with crossing the aortic valve (av) and maintaining proper position due to resistance from patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12132: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. d4: model number.
25655 | 1220648-2024-12159 | 19489840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489840&pc=OZD | 07/22/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/22/2021 | 06/07/2024 |  |  | 04/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079801 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with manual pressure and balloon occlusion. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25656 | 1220648-2024-12137 | 19489932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19489932&pc=OZD | 11/14/2023 | 06/07/2024 |  | 06/07/2024 |  |  | 05/15/2024 | 06/07/2024 |  |  | 07/31/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024377709 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a 83-year-old male noted as high-risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during impella 5.5 support, the patient pfhb increased up to 109.5 m/l on 17.11.2024 and the ldh (actate dehydrogenase) increased up to 3720 u/l on 17.11.202 and was diagnosed with hemolysis. the pump was ultimately explanted six days later, and the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25657 | 1220648-2024-12164 | 19490177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490177&pc=OZD | 07/25/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/25/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087401 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support an ischemic limb was noted. therefore, the device was explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25658 | 1220648-2024-12166 | 19490183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490183&pc=OZD | 05/30/2024 | 07/23/2024 |  | 06/07/2024 |  |  | 05/30/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482678 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon opening an introducer kit, the .035 stiff wire was found to be broken. the implant proceeded without use of the wire. there was no patient harm resulting from the broken stiff wire. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The root cause of the product damage (guidewire damage - peripheral vessel) issue was not determined since product was not returned for investigation. f6/f8-should have been left blank in the initial report. g1 manufacturing site name and address.
25659 | 1220648-2024-12121 | 19490264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490264&pc=OZD | 07/13/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/13/2021 | 06/07/2024 |  |  | 02/28/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055745 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2021 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing from all insertion sites therefore, red blood cells were given to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
25660 | 1220648-2024-12175 | 19490532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490532&pc=OZD | 07/29/2021 | 06/07/2024 |  | 06/07/2024 |  |  | 07/29/2021 | 06/07/2024 |  |  | 05/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had significant hemolysis in the intensive care unit. an echo was performed, finding the impella cp was 5 cm in ventricular. the patient was vomiting and attempting to sit up in bed. after repositioning the pump to 35 cm, the catheter was not in the mid ventricular space and was hugging the wall. after multiple attempts, the pump was repositioned under x-ray. another echo was completed after, confirming the pump was at 3.5 cm. later that evening as the hemolysis continued and the pump was still not in correct position. the physician chose to pull the impella back into the aorta, maintaining a p1 flow until stability can be confirmed or maintaining access for a potential placement of a new impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. cp with smartassist system instructions for use for the related event are as follows: section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
25661 | 1220648-2024-12146 | 19490691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490691&pc=OZD | 05/29/2024 | 08/01/2024 |  | 06/07/2024 |  |  | 05/29/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481923 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella 5.5 insertion the pocket was closed but with continual oozing, a decision was made to reopen the pocket. bioglue, gelfoam, and surgiflow were used to control bleeding at the anastomosis. the pocket was closed, and pressure dressing applied. it was further reported that despite correcting coagulation and replenishing with blood products, the patient continues to bleed from the right axillary pocket. the physician took the patient to the operation room (or) to explore and a washout. the patient is off anticoagulation due to bleeding concerns. the patient had improvement of axillary site bleeding. the physician took the patient again to or for hopefully the final washout and closure of the pocket. the patient is stable. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- b5 mso review, h6 component code 925, d4-corrected model number, f6/f8 should have been left blank in the initial report.
25662 | 1220648-2024-12183 | 19490747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490747&pc=OZD | 05/10/2024 | 06/19/2024 |  | 06/07/2024 |  |  | 05/11/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1234885 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The complainant reported a 63-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the automated impella controller (aic) switched off during the implant procedure. after medical staff restarted the aic, an unexpected console shutdown error message displayed. the patient eventually expired, but there is no association between the impella use and the patient¿s outcome. | [5 empty] | The investigation into the aic - shutdown or restart issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure mode was not reproduced during testing, but the audio sw process could be terminated causing the sw watchdog to trigger and resetting the console. the cause of the unexpected reboots was the audio server crashing triggering the sw watchdog to do partial and full resets of the console as intended. | The impella device was received, and an evaluation has begun. upon completion of the investigation, a supplemental mdr will be submitted.
25663 | 1220648-2024-12154 | 19490892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490892&pc=OZD | 03/18/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/18/2022 | 06/07/2024 |  |  | 11/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167887 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported a hematoma was observed at the access site and the patient did not have palpable or doppler pulses in the pedal, post tibial, or popliteal regions. the right femoral artery (rfa) was surgically exposed, and the impella cp device was turned off and explanted without issue. surgical wound closure with patch repair was performed on the rfa and a jackson-pratt drain was added. doppler was confirmed distal to patch prior to closure of the wound. hemostasis was achieved per surgical closure and the procedural outcome is the patient survived surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25664 | 1220648-2024-12190 | 19490978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19490978&pc=OZD | 09/24/2021 | 06/07/2024 |  | 06/07/2024 |  |  | 09/24/2021 | 06/07/2024 |  |  | 07/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105934 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with blood transfusion. additionally, the patient experienced sepsis from which the patient recovered with no specified intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
25665 | 1220648-2024-12143 | 19491448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19491448&pc=OZD | 03/08/2024 | 06/24/2024 |  | 06/07/2024 |  |  | 05/10/2024 | 06/07/2024 |  |  | 11/30/2025 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429406 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported a 55-year-old male patient with aortic valvuloplasty was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed a hematoma during support. medical staff applied manual pressure and administered multiple units of packed red blood cells. the patient recovered and survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the product was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
25666 | 1220648-2024-12169 | 19491914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19491914&pc=OZD | 07/28/2021 | 06/07/2024 |  | 06/07/2024 |  |  | 07/28/2021 | 06/07/2024 |  |  | 11/30/2022 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021009861 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/28/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with perclose and additional suturing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25667 | 1220648-2024-12153 | 19492069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492069&pc=OZD | 03/16/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/16/2022 | 06/07/2024 |  |  | 10/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced slight bleeding from the hemostatic valve. the patient was hemodynamically stable after a unit of blood. the physician weaned and explanted the impella cp device due to high risk of bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25668 | 1220648-2024-12168 | 19492072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492072&pc=OZD | 03/22/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/22/2022 | 06/07/2024 |  |  | 12/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the device had continuous suction alarms. it was set to power level p1 until the issue was resolved. it was then reported that the device may have been occlusive, as the patients right foot appeared mottled and the calf was very tight. the device was then explanted, and subsequently there was patent flow to the right lower extremity. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25669 | 1220648-2024-12120 | 19492095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492095&pc=OZD | 08/23/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 08/23/2021 | 06/07/2024 |  |  | 04/30/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070645 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2021 |  |  |  |  | The user facility reported a 67-year-old female needing impella cp for mechanical circulatory support. it was reported that a patient presented with non-q wave myocardial infarction- left main coronary artery, ostial left anterior descending artery, ostial circumflex artery, and mid section left anterior descending artery (mlad). it was reported that there was access site bleeding. the tamponade was crossed over with balloon and the patient was administered a unit of blood. the patient was sent to the operating room for a wire perforation that required a stitch. the impella cp device was explanted in procedural area and the patient is stable. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25670 | 1220648-2024-12174 | 19492424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492424&pc=OZD | 08/25/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 08/25/2021 | 06/07/2024 |  |  | 03/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022061918 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the physician suspected that the patient may be experiencing heparin-induced thrombocytopenia (hit). heparin was discontinued and replaced with another anticoagulant, and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The cause of the issue was not determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records.
25671 | 1220648-2024-12180 | 19492442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492442&pc=OZD | 05/11/2024 | 07/29/2024 |  | 06/07/2024 |  |  | 05/11/2024 | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025465779 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The user facility reported a patient of unknown race, age, and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the purge filter was leaking. during a follow up call, abiomed learned the pump stopped due to the leakage. the pump was explanted. the care was noted to have been withdrawn and the patient expired. per abiomed's medical safety officer review, there is not enough information to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. the pump was returned for investigation. data logs were not provided for this investigation. the pump passed electrical resistance testing. the pump was soaked in warm water and started for a test run. the pump failed to start several times, triggering a high motor current alarm. the purge was started for further investigation. a purge leak was noted from sidearm yellow luer crack running down the length of the luer. further investigation revealed blood in the purge line. the cause of the pump stop was damage to the sidearm yellow luer, which allowed a purge leak and blood to ingress into the motor housing due to the absence of purge. added- h6 component code 481 f6/f8 should have been left blank in the initial report.
25672 | 1220648-2024-12163 | 19492520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492520&pc=OZD | 07/22/2021 | 03/24/2025 |  | 06/07/2024 |  |  | 07/22/2021 | 06/07/2024 |  |  | 02/28/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022051646 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed bleeding at the groin. manual pressure was applied for three hours along with femstop. however, the bleeding did not stop. it was decided to explant the pump. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
25673 | 1220648-2024-12167 | 19492708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492708&pc=OZD | 03/22/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/22/2022 | 06/07/2024 |  |  | 07/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153093 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there were potential issues with blood flow to the lower right extremity. a fem-fem bypass was placed prophylactically for right lower extremity perfusion, resulting in positive pulses in the lower extremity. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
25674 | 1220648-2024-12157 | 19492807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492807&pc=OZD | 03/18/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/18/2022 | 06/07/2024 |  |  | 12/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178395 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | Material Twisted/Bent | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was found to have poor limb perfusion, and there was a kink in both the short and long sheath of the impella. it is unknown if/how these issues were resolved, but the patient continued on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25675 | 1220648-2024-12176 | 19492812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19492812&pc=OZD | 03/25/2022 | 03/24/2025 |  | 06/07/2024 |  |  | 03/25/2022 | 06/07/2024 |  |  | 12/31/2023 |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP- EU |  |  | ABIOMED, INC. |  |  | US |  | 2023163088 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/25/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after impella placement with moderate resistance, the graft and anastomosis were bleeding heavily. 6 units of packed red blood cells, 4 units of plasma, and 5 liters of fluid were administered to the patient. the pump and graft were surgically removed and the wound was closed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25676 | MW5156056 | 19503623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19503623&pc=OZD |  | 04/23/2024 |  | 06/07/2024 |  |  |  | 06/07/2024 |  |  |  |  | 06/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PERCUTANEOUS HEART PUMP |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [9 empty] | Voluntary report | Not Available | Other | [6 empty] | NA |  | Unknown |  |  |  | Injury | Yes | No | [7 empty] | Hemorrhage/Blood Loss/Bleeding | Hypovolemic Shock | [40 empty] | Initial submission | [9 empty] | It was further reported that peripheral bleeding from both the deep femoral artery and deep femoral vein occurred right after implantation of the percutaneous heart pump under cpr (cardiopulmonary resuscitation). the patient developed hemorrhagic shock from the bleeding two days later. the patient was transfused with a platelet concentrate and the bleeding sites were stitched up. this report reflects information received by fda in the form of a notification per 803.22 (b)(2).
25677 | 1220648-2024-12064 | 19475676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19475676&pc=OZD | 12/08/2023 | 10/27/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 09/30/2024 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023282032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 05/13/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used. hemoglobin (hgb) 9.4 on arrival to the hospital. 11 hours later, post-procedure, hgb dropped to 7.0. anemia secondary to hemolysis from impella. patient transfused with 3 units packed red blood cells and impella ultimately explanted. the patient recovered with sequelae. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding an impella cp that had a device deficiency. the "impella removed (b)(6) 2023 rota burr stuck occluding the lad." |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. the pump was not returned. data logs were reviewed finding many suction alarms occurred. no motor current spike observed. no positioning alarms occurred. no hemolysis related issue indicated. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. f6/f8 should have been left blank in the initial report. | Corrections to the initial mfr report: b1-selected product problem. b5-added additional failure mode. g1 mdr reporting contact email . added h6 component code 925, med dev prob code 1158. | The investigation hemolysis and positioning issue has been completed. the impella device was not returned for evaluation. hemolysis: no changed in the investigation results that was previously reported. positioning issue: the root cause of positioning issue most likely was patient condition related since rotablator burr stacked in the left anterior descending artery (lad) and impella removed.
25678 | 1220648-2024-12066 | 19475779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19475779&pc=OZD | 05/25/2024 | 08/01/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025474713 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The user facility reported a 69-year-old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported that the patient experienced oozing from their access site during impella 5.5 support. a washout of the site was completed to manage the condition and support with the implanted pump continued. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- b5 mso review, d6b, h6 component code 925 f6/f8 should have been left blank in the initial report.
25679 | 1220648-2024-12067 | 19475800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19475800&pc=OZD | 08/09/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/09/2021 |  |  |  |  | The user facility reported a 79-year-old patient implanted with an impella cp for mechanical circulatory support. the device appeared shallow due to the device being under the papillary structure. pump metrics were good. the patient was weaned and the device explanted. plans were made for coronary artery bypass grafting. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25680 | 1220648-2024-12068 | 19475896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19475896&pc=OZD | 05/27/2024 | 07/18/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The user facility reported a 39-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support did not have distal limb perfusion in either leg. bilateral foot wounds. impella cp was removed from the right femoral artery (rfa) and va ecmo (venoarterial extracorporeal membrane oxygenation) cannulated on the left with anterograde perfusion sheath. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
25681 | 1220648-2024-12069 | 19475941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19475941&pc=OZD | 05/08/2024 | 07/17/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468872 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Tear, Rip or Hole in Device Packaging | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during impella 5.5 support, a hole was found in the repositioning sleeve, approximately 25 centimeters away from the repo sheath. the hole reportedly did not affect the function of the pump, the patient expired on support. upon review by abiomed's medical safety officer, there is not enough information to associate the use of device with the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported product damage has been completed. no product was returned for investigation. the root cause of the sterile sleeve damage was not determined since the product was not returned and there was a lack of clinical details. e.1 facility name was revised as previously entered incorrectly on the initial manufacturer device report 1220648-2024-12069.
25682 | 1220648-2024-12070 | 19476101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476101&pc=OZD | 05/30/2024 | 08/08/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482932 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported delivery issues and vascular dissection during an attempt to implant an impella 5.5 pump in a patient for mechanical circulatory support. the impella 5.5 pump was being delivered to the patient being prepped for gastrointestinal surgery and was found to have left main coronary artery disease. the team placed the impella 5.5 pump but after the graft was placed a subclavian artery dissection was observed which was treated with two stents. the team then attempted to deliver the impella 5.5 pump, but native anatomy prevented placement. it was noted the entire pump was able to go through the graft and deliver except the motor. the decision was made to abort the use of the impella 5.5 pump and instead placed an impella cp pump without incident. the patient was noted in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | D.9 updated as the pump and purge cassette were returned for investigation. the investigation has been completed. the pump was returned for evaluation. upon visual inspection, a kink in the catheter just below the motor housing was present. additionally, scratches were present on the motor housing and outflow cage of the pump. clinical details noted that artery was 7.03 millimeters before insertion. additionally, the patient had a dissection of the axillary artery after the graft was attached and two stents were placed to resolve it. when inserting the pump, the physician could not pass the motor housing through the artery even with excessive force applied. the root cause of delivery issues were most likely due to the patient's condition as the stent placed likely narrowed the vessel diameter and became too small to accommodate the impella 5.5. the root cause of the vascular damage could not be determined as there was limited clinical information provided. h.6 code 4624 was reported incorrectly on the initial manufacturer device report. a new code was added to health effect - impact codes.
25683 | 1220648-2024-12071 | 19476223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476223&pc=OZD | 05/30/2024 | 07/22/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025489008 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 64-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that there was no doppler signal in the right leg. the decision was made to explant the impella due to ischemia and replace with an intra-aortic balloon pump. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised from the initial submission of manufacturer device report 1220648-2024-12071 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12071 and should not have been. h.6 added code 3038 since the initial manufacturer device report 1220648-2024-12071 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report 1220648-2024-12071 in accordance with updated procedures.
25684 | 1220648-2024-12074 | 19476344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476344&pc=OZD | 05/30/2024 | 07/15/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 03/31/2025 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024327149 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 68-year-old male patient of unknown ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after 3 days of support the pump was explanted due to concerns of hemolysis. the patient was reported stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for hemolysis has been completed. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. f6/f8-should have been left blank in the initial report.
25685 | 1220648-2024-12073 | 19476348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476348&pc=OZD | 05/23/2024 | 03/20/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025470701 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant or left ventricle assist device. the patient was suspected to have a perforated myocardium related to hospital-supplied 0.035" wire. this was noted following a cardiac surgical procedure. a sternotomy was performed to drain and patch the perforation. clinicians believe that the patient's myocardial infarction a few days prior, compromised the anatomy making the patient susceptible to tamponade. the patient is noted to be stable. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance." | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. based on the clinical detail an impella 5.5 was successfully delivered in the patient. during the procedure, the myocardium was perforated, and it was suspected to be caused by the hospital-supplied 0.035" guidewire. the patient's anatomy was also susceptible to tamponade due to a myocardial infarction (mi) a few days prior. logs were not applicable for this failure mode. the root cause of the ventricle perforation was most likely use related as a non-abiomed wire was the suspected cause of the perforation. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12073. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-12073. g1 reporting contact email revised on manufacturer device report 1220648-2024-12073 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-12073.
25686 | 1220648-2024-12077 | 19476356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476356&pc=OZD | 05/24/2024 | 07/11/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025505494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 72-year-old male patient of unknown race and ethnicity noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a slight dissection of the subclavian artery at the insertion point of the introducer. no intervention was performed on behalf of the dissection and support with the implanted impella cp pump was uninterrupted. the patient was reported stable. | [5 empty] | The investigation for the reported vascular damage has been completed. no product was returned for investigation. there was a slight dissection of the subclavian artery at insertion point of the 14f sheath which did not require intervention. the root cause of the vascular damage peripheral vessel issue was not determined since product was not returned and there is a lack of clinical details. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows section: warnings and cautions: to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25687 | 1220648-2024-12076 | 19476375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476375&pc=OZD | 08/19/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090926 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/19/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had runs of ventricular tachycardia, so the impella was repositioned. as the patient continued on support, it was reported that the patient went into atrial fibrillation/rapid ventricular response requiring cardioversion and also coded. the patient was evaluated and it was noted that the ejection fraction was decreased to fifteen percent and the impella appeared deep and was pulled back at bedside. the patient was escalated to impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25688 | 1220648-2024-12075 | 19476376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476376&pc=OZD | 05/24/2024 | 07/23/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482240 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during emergency coronary artery bypass grafting (cabg). the placement signal signifying pump was retracted back to the aorta along with impella red alarm with flows of 4.0 on auto mode. after transesophageal echocardiography was done, impella inlet was found in lvot. the impella was not in an optimal position. the surgeon proceeded to not advance the impella catheter with concerns of unsuccessful aorta cross clamping if catheter is across the aortic valve. impella was switched to surgical mode as the impella cp was not in the left ventricle (lv). at the time of the cross clamp, the physician retracted impella catheter to proximal aorta to cross clamp. remainder of cabg procedure ensued until the time to advance impella cp back into lv. the physician manually guided the catheter across the valve, proceeded to come off cardiopulmonary bypass support. however unable to achieve proper placement as catheter pointed to posterior lv with mitral valve leaflets occluding inlet. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for evaluation; however, data logs were made available. data log reviewed: positioning alarms in aorta occurred in the time of reported issue. suction alarms occurred. the cause of the positioning issue most likely was use related due to improper technique as the impella retracted back to aorta. f6 and f8 should be blank.
25689 | 1220648-2024-12079 | 19476492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476492&pc=OZD | 05/31/2024 | 08/11/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025481077 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that the sterile sleeve ripped during repositioning in intensive care unit (icu). it was further reported that the patient developed hematoma after starting heparin and was taken to the operation room for a left chest washout of the impella site. heparin was held. the patient remains on support and is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D.6b explant date was added. the investigation for the access site bleed and the product sterile sleeve damage has been completed. the impella device was not returned for evaluation. based on the clinical notes, all impella insertion and closure procedures followed best practices. the hematoma formed after beginning heparin infusions. this event was successfully managed and resolved by taking the patient off of heparin. the root cause was determined to be anticoagulation management. the clinical review confirmed that the sterile sleeve ripped while repositioning the patient in the intensive care unit due to excessive force. the root cause of the sterile sleeve damage was determined to be a use issue. b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-12079. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12079 and should not have been. h.6 codes 4114, 4624, and 4628 were reported incorrectly on the initial manufacturer device report number 1220648-2024-12079. new codes have been added. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2024-12079 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-12079 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-12079 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
25690 | 1220648-2024-12080 | 19476531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476531&pc=OZD | 05/29/2024 | 07/18/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490180 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that when inserting the impella cp the patient had a gap between the stents in the peripheral anatomy that kept impella from passing. attempt was made to dilate area but was unsuccessful. the impella cp case was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci . section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue is determined to be patient condition because of patient anatomy due to gap between the stents. instructions for use: section: warnings and cautions: to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ a5 ethnicity should have been blank in initial report as it is unknown. d4 model number, f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing, h.6 codes 2920 and 2338 were reported incorrectly on manufacturer device report 1220648-2024-12080. new code was added to medical device problem code.
25691 | 1220648-2024-12082 | 19476540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476540&pc=OZD | 05/30/2024 | 08/07/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453964 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Suction Problem | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a (b)(6) year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp had continued error of suction, unreliable placement signal, and glitching on the console. after the pci (percutaneous coronary intervention) procedure the impella was removed. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation placement signal issue has been completed. the imeplla device was received for investigation and no problem was found with the pump during the case as the device functioned to specification. d9 date device returned to manufacturer added b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12082. d.2 common device name was revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12082. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-12082 in accordance with updated procedures. d.9 revised as the pump was returned prior to submission of initial manufacturer device report number 1220648-2024-12082. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12082 and should not have been. g.1 facility name was inadvertently omitted from the initial manufacturer device report number 1220648-2024-12082. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12082 was submitted. g.1 manufacturing site fax number revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-12082. h.3 the device was returned at time of initial submission of the initial manufacturer device report number 1220648-2024-12082. if the device not evaluated, should of reflected that the device evaluation anticipated, but not yet begun. h.6 code 3009, 2170, and 403 were reported incorrectly on the initial manufacturer device report number 1220648-2024-12082. an ew code was added to medical device problem code.
25692 | 1220648-2024-12081 | 19476542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476542&pc=OZD | 05/31/2024 | 08/30/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171604 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The complainant reported that the impella aic had a controller error. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) yellow alarm has been completed. data logs from the console indicate that numerous kbd-communication loss alarms were issued by the console. the console was visually inspected and internal circuitry found to be of good integrity. the console was then setup with an oscilloscope monitoring the data communication between the epc and glass keyboard and rebooted 200 times without being able to replicate the failure mode. the root cause could not be determined due to issue not being reproduced. precautionary replacement of the subsystem components most likely to account for the failure mode was warranted.
25693 | 1220648-2024-12083 | 19476543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476543&pc=OZD | 07/26/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 01/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding, which required adding blood products, manual pressure, and deployment of closure device. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25694 | 1220648-2024-12084 | 19476551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476551&pc=OZD | 05/30/2024 | 07/17/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025490033 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the pulses in the right lower extremity were unable to be found via doppler. ultrasound was ordered. the patient continued on support despite noted issue and was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report.
25695 | 1220648-2024-12085 | 19476570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476570&pc=OZD | 05/30/2024 | 07/22/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 11/30/2025 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024429164 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 67-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient labs came back hemolyzed. a decision was made to explant the impella at bedside. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. f6/f8-should have been left blank in the initial report.
25696 | 1220648-2024-12086 | 19476584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476584&pc=OZD | 05/31/2024 | 07/12/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 07/31/2025 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443828 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 54-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported the patient experienced ventricular tachycardia (vt) storm. the impella was pulled back one centimeter, which resolved the issue. the patient remains on support and is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the patient's ventricular tachycardia (vt) has been completed. neither the impella products nor data logs were returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined. f6 and f8 should have been blank.
25697 | 1220648-2024-12088 | 19476665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476665&pc=OZD | 07/27/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced delivery issues that led to abandonment of the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25698 | 1220648-2024-12090 | 19476686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476686&pc=OZD | 05/26/2024 | 07/31/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025455721 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on impella cp support suffered a cerebral hemorrhage and expired. the physician stated that it is not believed that the cp was the direct cause of the hemorrhage given the patient¿s poor original diagnoses, but the direct cause cannot be determined. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12090: f6/f8-should have been left blank.
25699 | 1220648-2024-12091 | 19476701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476701&pc=OZD | 08/23/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104473 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experience bleeding from an unspecified source during support. it was reported that patient received three units of packed red blood cells. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25700 | 1220648-2024-12087 | 19476727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476727&pc=OZD | 08/22/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 07/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105330 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, distal perfusion of the left leg was observed. several days later, the patient was taken for removal of the impella, and it was reported that the patient's gastrointestinal bleeding showed improvement. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25701 | 1220648-2024-12092 | 19476787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476787&pc=OZD | 08/23/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101259 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the physician abandoned placing 2nd sheath, against field team recommendations, to perform femoral cutdown for explant. the physician was unable to accomplish cutdown, and a vascular surgeon was called in and successfully completed the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25702 | 1220648-2024-08142 | 19476837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476837&pc=OZD | 08/05/2022 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2024 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023218330 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient developed a hematoma at the cutdown sight. washout of the cutdown sight was performed and heparin was removed from purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25703 | 1220648-2024-12093 | 19476872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476872&pc=OZD | 08/23/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089768 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. on startup there was no flow. suspecting thrombus, it was immediately removed. they carefully rinsed and flushed inlet, outlet and cannula. there was a possible small thrombus flushed out. a transesophageal echocardiography was then performed and there was no apparent left ventricle thrombus. cp reinserted and started without issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
25704 | 1220648-2024-12094 | 19476981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19476981&pc=OZD | 08/25/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084729 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support, it was reported that the patient had gastrointestinal bleeding requiring discontinuation of heparin. the patient had an esophagogastroduodenoscopy which showed nasogastric tube trauma which was no longer bleeding. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25705 | 1220648-2024-12095 | 19477000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477000&pc=OZD | 08/27/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 04/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080392 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the physician contacted support for an ongoing suction alarm. the physician administered a fluid bolus and ordered an echocardiogram, but the outcome of these interventions is unknown. days later, the physician contacted impella support again to discuss the position of the device and reported that the placement signal and left ventricle signals were disassociated. it is unknown if/how these issues were resolved, but the patient remained on impella support for several more days. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25706 | 1220648-2024-12096 | 19477002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477002&pc=OZD | 07/27/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087299 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2021 |  |  |  |  | The complainant reported the patient was on impella cp support due to high-risk percutaneous coronary intervention. while on support, the patient experienced access site bleeding that was resolved with perclose, redressing, blood products, and withholding of heparin. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
25707 | 1220648-2024-12097 | 19477022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477022&pc=OZD | 09/22/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 07/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112421 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. after impella placement and support the patient become more stable. the decision was made to explant the device because the distal leg was not being perfused due to small/diseased artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25708 | 1220648-2024-12098 | 19477058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477058&pc=OZD | 08/02/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022086497 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported the patient was on impella cp support for high-risk percutaneous coronary intervention. while on support, the patient experienced access site bleeding that required pump removal and blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25709 | 1220648-2024-12099 | 19477072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477072&pc=OZD | 09/24/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090933 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing at the impella groin access site. the site was redressed by staff and oozing stopped. no heparin was administered. the patient was on effient increasing the risk of bleeding. additionally, the patient's platelets count was low and it was planned to transfuse platelets and remove the impella. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25710 | 1220648-2024-12100 | 19477172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477172&pc=OZD | 08/02/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 03/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported the patient was on impella cp support due to high-risk percutaneous coronary intervention. while on support, the pump stopped, which was resolved by replacing the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25711 | 1220648-2024-12101 | 19477230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477230&pc=OZD | 03/11/2022 | 04/16/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 11/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported it took about 30 minutes to gain access due to the vessel's depth. access was eventually achieved. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves warning: pre-support evaluation: axillary insertion of the impella 5.5 catheter. direct aortic insertion: use of impella with transcatheter aortic valves. caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance." section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25712 | 1220648-2024-12102 | 19477266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477266&pc=OZD | 03/11/2022 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 10/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158803 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/11/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was implanted and the patient's limb was instantly non pulsatile with support. dobutamine was increased with minimal effect on mean arterial pressure but increased left ventricle pressures. the pump was repositioned and the ischemia was resolved. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25713 | 1220648-2024-12103 | 19477308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477308&pc=OZD | 03/11/2022 | 05/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 10/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149249 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support it was noted that impella was too deep and retroperitoneal bleeding was suspected. pump repositioning resolved the issue. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25714 | 1220648-2024-12104 | 19477313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477313&pc=OZD | 08/02/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093193 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported the patient was on impella cp support for electrophysiology/ablation. while on support, the patient experienced access site bleeding that was resolved by giving protamine and volume, and removal of the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25715 | 1220648-2024-12105 | 19477384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477384&pc=OZD | 03/11/2022 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087425 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/11/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced heparin-induced thrombocytopenia. the purge solution was changed to d5w and argatroban was infused systemically. additionally, it was reported that there was high purge pressure. tissue plasminogen activator (tpa) protocol was initiated with slight initial improvement in pump flow and pressure. the plan of care changed to removal of the pump as tpa protocol did resolve the high purge pressure. the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25716 | 1220648-2024-12106 | 19477415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477415&pc=OZD | 08/02/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081992 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed bleeding in the groin and heparin had been reduced. afterwards, the groin was dry, and that systemic heparin had been significantly increased. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25717 | 1220648-2024-12108 | 19477849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477849&pc=OZD | 08/05/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088771 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/05/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced positioning issues. the pump was left in place, but eventually changed out due to being in the patient for five days. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25718 | 1220648-2024-12109 | 19477936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19477936&pc=OZD | 08/04/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that required blood products. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25719 | 1220648-2024-12110 | 19478058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19478058&pc=OZD | 08/11/2021 | 06/06/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097561 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support, and it was noted that the bioprosthetic valve made the impella implantation difficult. a jr4 plus a straight wire were used to recross with good placement | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25720 | 1220648-2024-12111 | 19478064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19478064&pc=OZD | 03/12/2022 | 05/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, continuous suction issues and deep pump migration were noted. attempts of repositioning and p level adjustments were initiated to resolve the issues. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25721 | 1220648-2024-12112 | 19478121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19478121&pc=OZD | 03/14/2022 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 10/23/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2022 |  |  |  |  | The user facility reported a patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, no distal pulses were found on the impella leg. the impella cp was removed secondary to peripheral vascular disease. it is unknown if the action resolved the ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25722 | 1220648-2024-12113 | 19478308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19478308&pc=OZD | 08/12/2021 | 06/06/2024 |  | 06/06/2024 |  |  | 09/22/2015 | 06/06/2024 |  |  | 05/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia. duplex studies of affected limb showed monophasic flow popliteal area and clinically, the affected foot was noticeably cooler to touch, however appeared capillary blush/return was intact. other extremity was only mildly better. the following day, the physician noted distal leg perfusion in question, possible impella removal if flows do not improve. the impella was explanted the next day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25723 | 1220648-2024-12114 | 19478639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19478639&pc=OZD | 08/12/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 03/31/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064972 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the patient's urine was pink in color, and labs test returned showing hemolysis. the site was redressed, and the patient's urine began to clear. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25724 | 1220648-2024-12115 | 19478882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19478882&pc=OZD | 06/02/2024 | 08/09/2024 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2025455698 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | No | Yes | Device Handling Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was placed but, patient had blood loss/bleeding observed a the sterile sleeve. the product malfunction prompted the team to give blood products. the pump was not replaced as the team observed the neuro status and possible escalation to the 5.5. the amount of blood products delivered is unknown. the patient remains on impella cp support on the date of this report. | Pump remained on for support for a total of 5 days when the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The pump was returned and a leak test was performed. the valve failed the leak test and upon inspection, a deformation around the outer perimeter of the o-ring valve was found. an abnormal groove along the side was also found which compromised the hemostatic seal by creating a leak path. the root cause of the access site bleeding - major was determined to be a damaged valve. b5 additional information added d6b added. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12115: f6/f8-should have been left blank h1 type of reportable event changed to serious injury h6 health effect - clinical code 1888 and health effect - impact code 4643 inadvertently not included.
25725 | 1220648-2024-12116 | 19479256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19479256&pc=OZD | 08/15/2021 | 03/21/2025 |  | 06/06/2024 |  |  |  | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092464 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, medical staff became unable to find a doppler pulse. the right lower extremity was later found to be cool, and still without a doppler pulse. the following day, the ischemia continued to worsen and show no doppler pulse, so the decision was made to explant the impella to resolve the issue. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25726 | 1220648-2024-12038 | 19481056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481056&pc=OZD | 08/17/2021 | 03/21/2025 |  | 06/06/2024 |  |  | 08/17/2021 | 06/06/2024 |  |  | 06/30/2023 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that while the doctor was preparing to complete the case the patient went into complete heart block. the impella cp flow was increased and the temperature pacer was placed in the right femoral vein. the impella cp device was removed, but the patient continued to decompensate so the device was reinserted, and support started was restarted. hemodynamics was obtained at the end of the case. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25727 | 1220648-2024-12045 | 19481181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481181&pc=OZD | 05/28/2024 | 07/18/2024 |  | 06/06/2024 |  |  | 05/28/2024 | 06/06/2024 |  |  |  |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494667 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Restricted Flow rate | Fluid/Blood Leak | Improper Flow or Infusion |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and after the implant the purge side arm was cracked and the staff replaced the impella. | [5 empty] | The data logs showed no issues with pump performance and there were a few suction alarms that were resolved. a leak to the purge system that may or may not have resulted in a low purge pressure and/or open purge alarms. this includes events where low purge pressure and/or open purge alarms trigger with no visible leak. the root cause of the purge leak - pump was determined to be due to a damaged sidearm as mentioned in the clinical details but the location of the damage is unknown. f6/f8-should have been left blank in the initial report. g1 manufacturing site name and address corrected. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
25728 | 1220648-2024-12040 | 19481465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19481465&pc=OZD | 05/23/2024 | 07/18/2024 |  | 06/06/2024 |  |  | 05/23/2024 | 06/06/2024 |  |  | 11/30/2025 |  | 06/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024436283 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 76-year-old male noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp pump had to be removed from the left ventricle as it was interfering with the planned percutaneous coronary intervention (pci) procedure. the pci was completed successfully without impella support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. product was not returned for investigation. data logs did not record any issues related to positioning. as per the clinical notes, the physician decided to remove the impella device and proceed with pci due to the pump obstructing the way while cannulating the left main coronary artery. the cause of the positioning issue was most likely use related. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address.
25729 | 1220648-2024-12012 | 19466779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19466779&pc=OZD | 05/08/2024 | 06/13/2024 |  | 06/05/2024 |  |  | 05/08/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492083 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | DISTAL REPERFUSION CATHETER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported an 86-year-old female of unknown ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was taken to the cardiac catheterization lab to place distal reperfusion catheter to improve blood flow to the right lower extremity. the patient expired while on support. per medical safety officer review, there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25730 | 1220648-2024-12014 | 19467089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467089&pc=OZD | 08/03/2021 | 03/24/2025 |  | 06/05/2024 |  |  | 08/03/2021 | 06/05/2024 |  |  | 12/31/2022 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022028595 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implantation, multiple attempts were made to advance the pump; however, the medical staff was unable to achieve position. a cook medical 14fr x 30cm introducer sheath was inserted and the impella was placed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
25731 | 1220648-2024-12015 | 19467093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467093&pc=OZD | 08/08/2021 | 03/24/2025 |  | 06/05/2024 |  |  | 08/08/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022077379 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s leg was cool and absent pulses were noted. the patient was successfully weaned off of support and the pump was removed. | [5 empty] | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: limb ischemia. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25732 | 1220648-2024-12016 | 19467095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467095&pc=OZD | 05/20/2024 | 07/16/2024 |  | 06/05/2024 |  |  | 05/20/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025475019 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient's access site had an ooze. there was a wound vac and two units of red blood cells infused back to the patient. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25733 | 1220648-2024-12017 | 19467176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467176&pc=OZD | 05/26/2024 | 07/18/2024 |  | 06/05/2024 |  |  | 05/26/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | The complainant reported a 59-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the day after the implant, the patient developed ischemia which prompted external bypass, fasciotomy, and planned amputation. the limb amputation had to wait for the liver team, and the family decided to withdraw patient care. the patient eventually expired. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- h6 component code f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25734 | 1220648-2024-12018 | 19467359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467359&pc=OZD | 05/25/2024 | 07/22/2024 |  | 06/05/2024 |  |  | 05/25/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469159 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The user facility reported a 58-year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had oozing from their access site following the implant. the medical staff repositioned the pump to resolve the oozing. leading up to the device explant, the patient began to complain about leg pain. the medical staff then noticed loss of pulses in the leg and took the patient to the cardiac catheterization lab for both the explant and thrombectomy, and both procedures were performed. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25735 | 1220648-2024-12020 | 19467490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467490&pc=OZD | 09/23/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 09/23/2021 | 06/05/2024 |  |  | 07/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115408 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2021 |  |  |  |  | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. the impella was inserted but was unable to pass through abdominal iliac anatomy. the impella was removed and high risk percutaneous coronary interventions was not performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25736 | 1220648-2024-12021 | 19467591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467591&pc=OZD | 08/11/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/11/2021 | 06/05/2024 |  |  | 11/30/2022 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021018907 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported the patient's partial thromboplastin time was greater than 200 and the patient experienced access site oozing. hemodynamics and hemostasis remained stable. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25737 | 1220648-2024-12022 | 19467666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467666&pc=OZD | 08/12/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/12/2021 | 06/05/2024 |  |  | 05/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088754 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was taken to cath lab for impella removal following development of hematoma which 1 unit of red blood cell was given. the decision was made to not do further intervention, dry close technique using 9mm balloon for tamponade, perclose x 2 deployed. angiogram revealed small amount of leaking right above bifurcation of superficial femoral artery and profunda, unable to use covered stent, manual pressure was held and femostop low pressure placed. the patient hemodynamics were stable on low dose levophed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25738 | 1220648-2024-12023 | 19467695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467695&pc=OZD | 08/10/2021 | 10/20/2025 |  | 06/05/2024 |  |  | 08/10/2021 | 06/05/2024 |  |  | 03/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022065695 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding was noted. therefore, impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25739 | 1220648-2024-12024 | 19467702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467702&pc=OZD | 08/12/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/12/2021 | 06/05/2024 |  |  | 06/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093183 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the access site. perclose sutures failed. manual pressure was initiated. the impella cp was explanted the same day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25740 | 1220648-2024-12026 | 19467755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467755&pc=OZD | 05/28/2024 | 07/16/2024 |  | 06/05/2024 |  |  | 05/28/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484214 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The complainant reported a 70-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient had access site bleeding. the patient was bleeding around the repositioning sheath, and the doctor removed the impella. the patient survived the event. | [5 empty] | The investigation into the access site bleeding has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. h6-added component code f6 and f8 should have been blank. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25741 | 1220648-2024-12025 | 19467756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467756&pc=OZD | 08/12/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/12/2021 | 06/05/2024 |  |  | 06/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101234 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding and hematoma. two units of blood were transfused to the patient to resolve this issue. the device was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25742 | 1220648-2024-12027 | 19467782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467782&pc=OZD | 08/12/2021 | 03/24/2025 |  | 06/05/2024 |  |  | 08/12/2021 | 06/05/2024 |  |  | 06/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101735 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced an aystolic episode and cardiac arrhythmia during insertion. no further information was provided. | [5 empty] | To make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25743 | 1220648-2024-12028 | 19467793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467793&pc=OZD | 05/29/2024 | 07/23/2024 |  | 06/05/2024 |  |  | 05/29/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025504342 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported a 61-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient¿s urine turned from pink to red. impella repositioned but the physician was unable to obtain inlet to mid ventricular space. labs sent came back hemolyzed. it was further reported that the patient developed bleeding at the impella access site. manual pressure was held but hemostasis was unable to be achieved. heparin purge changed to bicarbonate and femstop was applied but bleeding did not resolve so manual pressure continued. the patient was transferred to another facility. hemostasis was temporarily achieved but hematoma appeared again. the patient was escalated to 5.5 support. cp removed and vessel repaired surgically. the patient is stable. | [5 empty] | The investigation for access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added 925 f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
25744 | 1220648-2024-12029 | 19467831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467831&pc=OZD | 08/12/2021 | 03/24/2025 |  | 06/05/2024 |  |  | 08/12/2021 | 06/05/2024 |  |  | 05/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the complainant reported there was concern of the impella cp device being too deep in the left ventricle upon reviewing the echocardiogram. when attempting to withdraw the pump, the device was pulled into the aorta. after attempts to re-cross the valve, the physician made the decision to proceed towards removal of the impella cp device. the patient was stable throughout. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25745 | 1220648-2024-12030 | 19467852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467852&pc=OZD | 08/10/2021 | 03/24/2025 |  | 06/05/2024 |  |  | 08/10/2021 | 06/05/2024 |  |  | 06/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022099765 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced several episodes of ventricular tachycardia (vt) at the beginning of the case during rota ablation of the left main ventricle. it is unknown how hemostasis was achieved in regard to the vt. | [5 empty] | To make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. smart assist system with automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
25746 | 1220648-2024-12032 | 19467991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19467991&pc=OZD | 08/10/2021 | 03/24/2025 |  | 06/05/2024 |  |  | 08/10/2021 | 06/05/2024 |  |  | 01/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022042332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, oozing and bleeding was observed at the access site during bedside explant of the impella device, which was managed by tightening the preclose sutures as well as additional anchoring suture placed at the repo sheath. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25747 | 1220648-2024-12034 | 19468095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468095&pc=OZD | 08/04/2021 | 03/21/2025 |  | 06/05/2024 |  |  | 08/04/2021 | 06/05/2024 |  |  | 09/30/2021 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1445812 |  | 0048-0032 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was experiencing severe bleeding from multiple sites while receiving medical support. no further treatment was provided, and the patient was successfully weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25748 | 1220648-2024-12035 | 19468201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468201&pc=OZD | 08/10/2021 | 03/21/2025 |  | 06/05/2024 |  |  | 08/10/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078033 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, suction above p6 was noted. echocardiogram showed great impella positioning. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25749 | 1220648-2024-12036 | 19468344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468344&pc=OZD | 08/13/2021 | 03/21/2025 |  | 06/05/2024 |  |  | 08/13/2021 | 06/05/2024 |  |  | 05/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/13/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that the patient had a cold right leg. the impella cp device was exchanged for a perc ax cp in order to extubate and ambulate patient. the impella cp was successfully weaned and removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
25750 | 1220648-2024-12037 | 19468515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468515&pc=OZD | 08/11/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/11/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082064 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, oozing was noted at the groin access site. mattress sutures and a pressure dressing change were utilized, and hemostasis was achieved. suction alarms were observed the following day with low volume and pulsatility and the patient was transfused an unknow number of units of packed red blood cells and provided fluids. systematic heparin was held due to the reported bleed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25751 | 1220648-2024-12039 | 19468646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468646&pc=OZD | 08/11/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/11/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070601 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was noted that the patient's heart rate was erratic due to paroxysmal atrial fibrillation. the impella was weaned, and the patient remained stable, therefore, the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25752 | 1220648-2024-12042 | 19468713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468713&pc=OZD | 05/29/2024 | 07/29/2024 |  | 06/05/2024 |  |  | 05/29/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437478 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant, a left ventricle perforation occurred. the perforation was at the site of friable tissue/infarcted tissue. the perforation had surgical repair. the pump was explanted and patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12042: f6/f8-should have been left blank. g1 reporting contact fax number missing.
25753 | 1220648-2024-12044 | 19468772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468772&pc=OZD | 05/30/2024 | 09/23/2025 |  | 06/05/2024 |  |  | 05/30/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025449536 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella 5.5 smart assist and while on support, there was a pump stop. the patient was transferred to the ct (computed tomography) table and his body position was adjusted, the pump went from p-8 to p-0. attempts to restart the pump were made to no avail. the pump was removed. it is possible that the shaft or cord was pulled. there had been no increase in motor current, decrease in purge flow rate, or increase in purge pressure until just before the event, and anticoagulation had been within the appropriate range, so it is believed that the problem was caused by excessive force being applied to the pump. | The investigation uncovered optical/placement signal issues. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the pump stop has been completed. field data logs confirmed that alarm 115 triggered along with a motor current spike to 3000 ma on the day of the failure mode, indicating an electrical failure. additionally, at the same time stamp, all the optical 5s parameters went out of specification. the pump was returned for evaluation and upon visual inspection, all three motor wires had been broken along with the optical fiber inside of the red handle. 5s parameters were all out of specification when plugged into the aic in the lab. the root cause of the pump stop was determined to be electrical open due to excessive force. the root cause of the optical signal issue was a damaged fiber line due to excessive force. added- d9 device return date d4-corrected model number. f6/f8 should have been left blank in the initial report. | B.5. updated to include optical signal issue description. h.6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-12044. b.7. other relevant history, including preexisting medical conditions updated. d.10. concomitant medical products and therapy dates updated. f.6. and f.8. should have been left blank g.1. reporting contact email and reporting contact fax number updated.
25754 | 1220648-2024-12046 | 19468813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468813&pc=OZD | 05/30/2024 | 07/23/2024 |  | 06/05/2024 |  |  | 05/30/2024 | 06/05/2024 |  |  | 07/31/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442181 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The complainant reported an 89-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed uncontrolled hematoma. the physician then removed the impella, and the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding/hematoma has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12046: f6/f8-should have been left blank. g1 reporting contact fax number missing.
25755 | 1220648-2024-12047 | 19468817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468817&pc=OZD | 05/30/2024 | 07/23/2024 |  | 06/05/2024 |  |  | 05/30/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025493206 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 58-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after the implant low pump flow was noted. the impella powered on auto, ramped up, but never got above 1 lpm flow. left ventricle signal briefly appeared and immediately disappeared when suction alarmed on auto. attempted to take out of auto and only got suction to break at p1. attempted to reposition, but unable to break suction. the pump was successfully replaced with new impella cp. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | In reviewing the returned data logs from the case, there were no spikes in motor current that would explain why the pump was not able to produce more flow than it did around the time of the suction alarms. in investigating the returned product, there were no visual abnormalities such as impeller damage, kinks or biomaterial in the cannula or around the impeller. in the lab, the pump ran at p9 for several minutes achieving flow rates of 3.8 l/min. due to the inability to reproduce the failure and lack of further clinical detail, we were unable to determine a root cause for the low pump flow. d9 date device returned to manufacturer added he following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12047: f6/f8-should have been left blank.
25756 | 1220648-2024-12048 | 19468911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468911&pc=OZD | 05/29/2024 | 07/16/2024 |  | 06/05/2024 |  |  | 05/29/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467660 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella cp device for the support of a patient due to the patient's anatomy. the decision was made to have the impella cp pump advance through the axillary artery due to vessel size and tortuosity. the pump passed through the axillary artery; however, it was unable to advance through the subclavian even after multiple attempts. consequently, the case was aborted. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system: section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation into the delivery issues has been completed. abiomed products were not returned for evaluation and data logs were not received. the root cause of the delivery issue was most likely patient condition related, the impella was not able to delivered due to the patient's vessel size and tortuosity. d4 model number corrected.
25757 | 1220648-2024-12049 | 19468970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468970&pc=OZD | 05/08/2024 | 06/13/2024 |  | 06/05/2024 |  |  | 05/08/2024 | 06/05/2024 |  |  | 06/30/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427838 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced right foot ischemia since insertion of the impella cp pump and access site bleeding. upon arrival the patient became apneic and went into cardiac arrest. the patient responded to cardiopulmonary resuscitation. the right femoral artery was accessed, and the impella cp pump was inserted. percutaneous coronary intervention to the left main artery was completed. while on support in the unit, bleeding from the access site was noted. the patient was infused with multiple blood products. further information noted, there was concern of bleeding in the abdomen; bleeding from multiple other sources. per the physician the patient was in disseminated intravascular coagulation and not a candidate for escalation of care. the decision was made to withdraw care, and the patient subsequently expired. medical safety review outcome noted there is not enough information to associate the use of the device the patient's demise. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and access site bleed could not be determined. impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
25758 | 1220648-2024-12050 | 19468974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19468974&pc=OZD | 05/14/2024 | 07/18/2024 |  | 06/05/2024 |  |  | 05/14/2024 | 06/05/2024 |  |  | 09/30/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400269 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that 62-year-old male patient with history of cardiology and angiology came to hospital with allergic reaction to entresto and required resuscitation for ventricular fibrillation several times. the impella was implanted without issue with a short sheath. the initial act was 200, and patient received 15,000 iu of heparin. the patient's hemodynamics improved at low level. following check-in, the foot pulse on the impella side was not palpable. the patient condition at this time was moderate., but had increasing respiratory problems. patient condition deteriorated independent of ischemia, and patient expired. the impella ran over the entire period of time without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4-corrected model number f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address.
25759 | 1220648-2024-12051 | 19469027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19469027&pc=OZD | 05/30/2024 | 07/18/2024 |  | 06/05/2024 |  |  | 05/30/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478542 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.5 support had a stroke. this is a change from baseline an hour earlier. the patient was taken to stat ct (computed tomography) which showed ischemic stroke, and then to interventional radiology for mechanical thrombectomy of distal right m1. full flow reestablished but is still not following commands on left side. last echo was yesterday evening when device was pulled back slightly. it was further reported that impella was explanted the next day after the stroke, and no clots were observed on echo or device. the patient was able to move and lift left leg off the bed but cannot move left arm and has a left facial droop. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
25760 | 1220648-2024-11944 | 19469199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19469199&pc=OZD | 08/27/2021 | 03/21/2025 |  | 06/05/2024 |  |  | 08/27/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump did not start, positioning issues were noted and the patient¿s lab were hemolyzed. the outcome of the event was unknown. there was no additional information provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25761 | 1220648-2024-12053 | 19469325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19469325&pc=OZD | 05/21/2024 | 07/25/2024 |  | 06/05/2024 |  |  | 05/30/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025437612 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/30/2024 |  |  |  |  | The complainant reported a 70-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that the patient had coronary artery bypass surgery. heparin was administered post-surgery, and increased bleeding occurred. medical staff then withheld the heparin then slowly reintroduced it while monitoring it. it was believed the issue was related with heparin used in purging. (although it wasn't enough to be a problem.) the patient remains on impella support. | [5 empty] | The device was not returned for evaluation. upon completion of the investigation a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage - peripheral vessel was most likely use related due to improper anticoagulation. h6 component code added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-12053: d4 model number. e1 initial reporter name translation. f6/f8-should have been left blank. g1 reporting contact fax number missing. g2 report source: foreign missing. h.6 code 4441 was reported incorrectly.
25762 | 1220648-2024-12054 | 19469417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19469417&pc=OZD | 05/30/2024 | 07/22/2024 |  | 06/05/2024 |  |  | 05/30/2024 | 06/05/2024 |  |  | 11/30/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024429709 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 pump for mechanical circulatory support and the next morning, bloody stool was noted. heparin was withheld and the computed tomography scan showed no active bleed. however, endoscopy revealed bleed that was cauterized. blood volume was given. the patient remains stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. added d6b d4 model number.
25763 | 1220648-2024-12056 | 19469646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19469646&pc=OZD | 05/31/2024 | 07/30/2024 |  | 06/05/2024 |  |  | 05/31/2024 | 06/05/2024 |  |  | 10/31/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414777 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was a pump stop. while flowing on p9, a motor stoppage with alarm saying, "motor stop due to high motor current" occurred. the nurse restarted the impella on p5 with appropriate motor current, flow and waveforms. the pump was obtained to be sent back for investigation. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump stop and saturated flow issues has been completed. no product was returned for evaluation. data logs were reviewed and no pump stops were observed on the data logs. motor current was stable and consistent with p-level changes throughout case. higher than normal flows were observed during case. clinical details noted that the pump was able to be restarted at p-5 after pump stop and run for the remainder of the pci procedure. the root cause of the temporary pump stop cannot be determined as insufficient logs and pump were provided. the root cause of the saturated flow cannot be determined as insufficient logs and pump were provided. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address.
25764 | 1220648-2024-12055 | 19469666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19469666&pc=OZD | 05/31/2024 | 07/22/2024 |  | 06/05/2024 |  |  | 05/31/2024 | 06/05/2024 |  |  | 03/31/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324329 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 placement was aborted due to the impella being unable to pass through the anastomosis. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition due to stenosis which resisted the 5.5 to pass through. d4 model number corrected. f6/f8-should have been left blank in the initial report. h.6 codes 2920 and 2338 were reported incorrectly on manufacturer device report 1220648-2024-12055. new code was added to medical device problem code.
25765 | 1220648-2024-12057 | 19470231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19470231&pc=OZD | 05/31/2024 | 07/16/2024 |  | 06/05/2024 |  |  | 05/31/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467747 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2024 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient had a large hematoma in the right abdominal area and had bleeding from their impella cp access site following implant. as a result of the noted condition, the decision was made to remove the pump and place a new impella cp pump via the left femoral artery. a covered stent was placed in the right femoral artery and a balloon was inflated for dry closure of the site. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, the impella was in the aorta and reposition required. a covered stent was placed in the right femoral artery and a balloon was inflated for dry closure of the site. h6 added component code 925 f6/f8-should have been left blank in the initial report. g1 manufacturing site name and address. h6 corrected health effect - clinical code from 1886 to 1884.
25766 | 1220648-2024-11959 | 19470339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19470339&pc=OZD | 05/24/2024 | 08/01/2024 |  | 06/05/2024 |  |  | 05/24/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1631552 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 48-year-old male patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support. it was reported that on day 2 of support a the automated impella controller displayed a purge system blocked alarm. no kinks or blockages were observed. the impella at p9 showed concerns for pump saturation. purge cassette changed with no significant improvement to purge flows and purge pressure. the pump stopped with a successful restart, however shortly after the pump stopped again and failed to re-start. the pump was explanted, and the patient was reported hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation aic-pump stop has been completed. the impella device was received for investigation. data logs were reviewed, and visual inspection performed and there were no issues found. the device functioned/operated to specification. d9: date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11959: f6/f8-should have been left blank. h6: code 4629 was reported. new code was added to health effect - impact code.
25767 | 1220648-2024-12058 | 19471723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19471723&pc=OZD | 05/06/2024 | 08/02/2024 |  | 06/05/2024 |  |  | 05/06/2024 | 06/05/2024 |  |  | 10/31/2025 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024416928 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support for 57-year-old male patient, when changing from the peel away sheath to the repositioning sheath, the physician made an incision in the femoral artery. a groin bleed developed despite subsequent z-suture and cranial suture of the blue suture. the patient was noted to have expired during support due to global heart failure, which was not impella induced. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added-h6 component code 925 d4-updated model number. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address.
25768 | 1220648-2024-12059 | 19473105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473105&pc=OZD | 08/08/2021 | 03/21/2025 |  | 06/05/2024 |  |  | 08/08/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078911 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the original impella came out of the left ventricle and flipped on itself. the physician was able to remove the original pump and place a new pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: positioning impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25769 | 1220648-2024-12060 | 19473188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473188&pc=OZD | 08/09/2021 | 03/21/2025 |  | 06/05/2024 |  |  | 08/09/2021 | 06/05/2024 |  |  | 05/31/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089249 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were suction alarms if console was turned up to p-6. oozing was noted at the groin site and fem stop applied. there were attempts to reposition the previous night and they were unsuccessful. attending md did not want to re-position pump at that time. also, the patient became febrile. the patient was successfully weaned and the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25770 | 1220648-2024-12061 | 19473204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473204&pc=OZD | 08/09/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/09/2021 | 06/05/2024 |  |  | 04/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070412 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/09/2021 |  |  |  |  | The user facility reported a 79-year-old patient was admitted as a transfer from another hospital with a gastrointestinal (gi) bleed. the patient was placed on impella cp support due to acute respiratory arrest, followed by cardiac arrest. it was reported that while on support, gi bleed and pump stoppage was noted. heparin was put on hold due to the gi bleed. the patient was weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
25771 | 1220648-2024-12062 | 19473271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473271&pc=OZD | 08/09/2021 | 05/21/2025 |  | 06/05/2024 |  |  | 08/09/2021 | 06/05/2024 |  |  | 10/31/2022 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022035778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/09/2021 |  |  |  |  | The user facility reported a 68-year-old patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced a spasm in the vessel when the device was implanted. nicardipine was injected and the vessel relaxed and flow was restored. after the patient was stable, the impella was weaned from the patient and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25772 | 1220648-2024-12031 | 19473276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473276&pc=OZD | 05/26/2024 | 03/18/2025 |  | 06/05/2024 |  |  | 05/26/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1375165 |  | 0042-0000-AU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 58 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a placement signal failure occurred on the automated impella controller (aic) during patient support. consequently, due to the missing placement signal, the aic was replaced with a backup console. there was no report of patient harm. | [5 empty] | The automated impella controller was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the console logs from the reported day of event are consistent with complaint and show that during case start, a ¿placement signal not reliable¿ condition was detected, and case start wizard prompted operator to disconnect pump due to ¿optical sensor system error¿ during the device analysis the cause of the issue was not determined since issue could not be reproduced. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report (b)(4) in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report (b)(4) in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on manufacturer device report (b)(4). g2 report source; foreign was inadvertently omitted on manufacturer device report (b)(4). h6 component code was revised on manufacturer device report (b)(4) to reflect correct device.
25773 | 1220648-2024-12033 | 19473483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473483&pc=OZD | 05/29/2024 | 07/11/2024 |  | 06/05/2024 |  |  | 05/29/2024 | 06/05/2024 |  |  |  |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485148 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported an 81-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp was removed distal flow to the right lower extremity was absent via doppler. the physician was aware of the patient¿s peripheral artery disease, which was the reason for immediate removal post procedure. the patient was reported as stable, and flow returned several hours after explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report.
25774 | 1220648-2024-12063 | 19473894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19473894&pc=OZD | 08/09/2021 | 06/05/2024 |  | 06/05/2024 |  |  | 08/09/2021 | 06/05/2024 |  |  | 06/30/2023 |  | 06/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095368 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/09/2021 |  |  |  |  | The user facility reported a 63-year-old patient needing an impella cp for mechanical circulatory support. during implantation, the device was unable to be inserted as the patient¿s vessels were too small. the procedure was aborted. additionally, the left femoral artery was oozing slightly and femostop was applied. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ this report is being filed as part of a retrospective review of historical records.
25775 | 1220648-2024-11954 | 19458042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458042&pc=OZD | 05/14/2024 | 07/23/2024 |  | 06/04/2024 |  |  | 05/14/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492415 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during implant procedure for impella cp, a purge pressure high alarm occurred. there were no kinks or blockages noted. the purge cassette was exchanged but the issue did not resolve. therefore, the pump was also exchanged as the patient needed immediate support. there was no reported patient harm. | [5 empty] | The impella device was received from the customer on 31-may-2024 and therefore, an evaluation of the device is under way. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | Pump was returned for evaluation. upon visual inspection, biomaterial was present in the purge gap. high purge pressure reproduction testing was performed and high purge pressure did reproduce and no blood ingress was noted in purge line. pump catheter was cut right before motor housing and high purge pressure resolved. pump teardown was performed and no biomaterial was present inside the motor or on the bearing. biomaterial was found inside the purge gap. data logs were reviewed and soon after the pump was turned on, a motor current spike was noted and purge pressure began to rise over a period of 10 minutes before "high purge pressure" alarm was triggered. the root cause of the high purge pressure was due to biomaterial present in the purge gap. d4 model number. f6/f8-should have been left blank in the initial report. g1 manufacturing site name and address.
25776 | 1220648-2024-11955 | 19458160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458160&pc=OZD | 05/16/2024 | 08/10/2024 |  | 06/04/2024 |  |  | 05/16/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481534 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The complainant reported 54-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the physician reports an arterial bleeding next to the impella puncture site, so the staff weaned the impella and explanted in the operating room (or). the patient is hemodynamically stable with a good left ventricle (lv) function. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.5 unique identifier (udi) # revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11955. e.1 facility name revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11955. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11955 and should not have been. g.1 add information to reporting contact address line 1 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-11955 submitted. g.1 revised manufacturing site name and address section as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11955. h.6 added code 925 since initial submission of manufacturer device report number 1220648-2024-11955 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-11955 in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer .¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25777 | 1220648-2024-11956 | 19458181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458181&pc=OZD | 07/25/2021 | 07/09/2024 |  | 06/04/2024 |  |  | 05/17/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | On 2024-05-02 complaints team forwarded publication out of bonn regarding 60 impella patients, 40 of which were impella 5.0. the entitled: "role of acquired von willebrand syndrome in the development of bleeding complications in patients treated with impella rp devices" detailed complications as 9 major bleed, 5 ecmo placements, 0 access site bleed, and 10 vw syndrome. patients were those implanted from (b)(6) 2020. | [5 empty] | The investigation into the access site bleeding ¿ major issue been completed. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. b.7 revised as this information was entered incorrectly on manufacturer device report number 1220648-2024-11956. d.1 device name and d.2 common device name were revised as this information was previously entered incorrectly on manufacturer device report number 1220648-2024-11956. d.4 model number was revised as this information was previously entered incorrectly on manufacturer device report number 1220648-2024-11956. e.1 first name, last name and facility name were revised as previously entered incorrectly on manufacturer device report number 1220648-2024-11956. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11956 and should not have been. g.1 revised reporting contact fax number from manufacturer device report number 1220648-2024-11956 in accordance with updated procedures. g.3 date was entered incorrectly on the manufacturer device report number 1220648-2024-11956. the date should of been 02-may-2024. h.6 code 4114 was reported incorrectly on manufacturer device report number 1220648-2024-11956. a new code has been added to type of investigation codes h.10 additional manufacturer narrative instructions for use (ifus) were reported on manufacturer device report number 1220648-2024-11956 incorrectly. no ifus are needed to be reported in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿assess access site for bleeding and hematoma. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ oezkur, m., reda, s., rühl, h., theuerkauf, n., kreyer, s., duerr, g. d., charitos, e., silaschi, m., medina, m., zimmer, s., putensen, c., & treede, h. (2021). role of acquired von willebrand syndrome in the development of bleeding complications in patients treated with impella rp devices. scientific reports, 11(1). https://doi.org/10.1038/s41598-021-02833-8.
25778 | 1220648-2024-11957 | 19458261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458261&pc=OZD | 05/22/2024 | 07/18/2024 |  | 06/04/2024 |  |  | 05/22/2024 | 06/04/2024 |  |  | 03/31/2025 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024322197 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 pump implanted in the patient for mechanical circulatory support had to be replaced. upon plugging in the impella 5.5 pump, the console displayed a controller failure. the controller failure noted that the pump could be started but that the placement signal would not be available. the automated impella controller was swapped in an attempt to troubleshoot the failure, however, the issue persisted. the impella 5.5 pump was then replaced, and the failure resolved. there was no patient harm resulting from the failure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed. according to the clinical review, when the pump was plugged into the automated impella controller (aic) it read controller failure. the console was then switched and the same alarm was present. the next screen on the console read replace the catheter or start without placement monitoring. the decision was then made to replace the catheter and the new pump had no issues. product evaluation confirmed a fiber break on the patient cable side within the red handle. data log analysis confirmed no other abnormalities. the cause of the optical signal issue was a damaged fiber line due to an error in the manufacturing process. additional information: d9 added device return date d4 corrected model number.
25779 | 1220648-2024-11958 | 19458372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458372&pc=OZD | 09/25/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 09/25/2021 | 06/04/2024 |  |  | 06/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient lost pulses in the foot. the decision was made to wean and remove the device, and manual pressure was applied. the procedure was tolerated well and hemodynamics was obtained. | [5 empty] | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25780 | 1220648-2024-11960 | 19458533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458533&pc=OZD | 11/28/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 11/28/2021 | 06/04/2024 |  |  | 08/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125541 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed leg ischemia, and the physician chose to explant the impella. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25781 | 1220648-2024-11961 | 19458547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458547&pc=OZD | 05/25/2024 | 07/31/2024 |  | 06/04/2024 |  |  | 05/25/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479608 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient post in postcardiotomy cardiogenic shock (post open-heart surgery) was implanted with an impella 5.5 device for mechanical circulatory support and while on support, the patient had bleeding from the chest. many blood products were provided, anticoagulation was withheld, and the patient had a washout. of furthermore, the impella had high purge pressure. tissue plasminogen activator was added to the purge. although the purge pressure and flows normalized after some time, the team did not want to risk purge problems with the impella pump again in the late evening. therefore, the cardiology team decided to place an intra-aortic balloon pump and remove the impella 5.5 pump, which was completed without incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for high purge pressure has been completed. no product was returned for evaluation. the data logs were reviewed and purge pressure experienced a sudden increase likely due to biomaterial ingestion on day two of support. the increased purge pressure values were not high enough to trigger a "high purge pressure" alarm until day six of support. clinical details noted that the patient had just received a mitral valve replacement and tricuspid valve repair in open heart surgery and was receiving lots of blood products while on support. dextrose 5% in water with sodium bicarbonate was present in purge solution at time of ingestion. once high purge pressure alarms were triggered, tissue plasminogen activator (tpa) protocol was initiated and purge values were returned to a normal range. additionally, heparin was added to the purge after tpa was completed. the root cause of the high purge pressure was most likely due to biomaterial ingestion. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b.1 adverse event was selected on the initial manufacturer device report 1220648-2024-11961 and should not have been. b.5 adverse event of access site bleeding mentioned in narrative on initial manufacturer device report number 1220648-2024-11961 should be omitted due to event was unrelated to impella device. b.6/b.7 information in b.6 submitted on the initial manufacturer device report number 1220648-2024-11961 belongs in b.7. b.7 was updated on this report to reflect this information. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-11961 in accordance with updated procedures. h.1 type of reportable event was revised to reflect malfunction based off the information received that the adverse reported in the initial manufacturer device report number 1220648-2024-11961 was unrelated to the impella device. h.6 codes 1888, 4643 and 4641 were reported incorrectly on manufacturer device report number 1220648-2024-11961. new codes have been added to health effect - clinical code and health effect - impact code. h.10 instructions for use were updated based off the information received that the adverse reported in the initial manufacturer device report number 1220648-2024-11961 was unrelated to the impella device.
25782 | 1220648-2024-11962 | 19458616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458616&pc=OZD | 05/14/2024 | 08/05/2024 |  | 06/04/2024 |  |  | 05/14/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490242 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the patient by accident broke the purge side arm during movement of the other leg. the impella cp was exchanged to an impella 5.5 and the procedure was successfully completed. the patient was stable. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation has been completed. upon investigation of the product, the purge sidearm was confirmed to be broken in half. the root cause of the purge leak was damage to the joint reservoir to filter as stated in the clinical description to have been broken during patient movement. f.6 and f.8 dates were reported on manufacturer device report number and should not have been.
25783 | 1220648-2024-11963 | 19458672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458672&pc=OZD | 05/25/2024 | 07/17/2024 |  | 06/04/2024 |  |  | 05/25/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506470 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant in the hospital with ami/cgs (acute myocardial infarction/ cardiogenic shock) and stemi (st elevation myocardial infarction). the team placed a cp to support the patient and allow for the pci (percutaneous coronary intervention). the team observed cad (coronary artery disease) and mr (mitral regurgitation). the cp was placed and allowed for pci. after 3 days the pump came out of position and was removed. no harm or injury. | [5 empty] | The investigation info the positioning issues has been completed. the impella cp was pulled back into the aorta while attempting repositioning. the pump was replaced. data logs were not applicable for this investigation. no product was returned for this investigation. the cause of the positioning issue was use related. f6 and f8 should have been blank. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25784 | 1220648-2024-11964 | 19458678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458678&pc=OZD | 12/10/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 12/10/2021 | 06/04/2024 |  |  | 04/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was slight oozing from the insertion site. it was noted that the impella cp device was explanted the following day. it is unknown if the oozing influenced the decision to explant the impella cp. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25785 | 1220648-2024-11965 | 19458737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458737&pc=OZD | 08/23/2021 | 06/04/2024 |  | 06/04/2024 |  |  | 08/23/2021 | 06/04/2024 |  |  | 06/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient had access site bleeding while on impella cp support. hemostasis was achieved with surgical wound closure at explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25786 | 1220648-2024-11966 | 19458741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458741&pc=OZD | 05/13/2024 | 07/11/2024 |  | 06/04/2024 |  |  | 05/13/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025467376 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 implant procedure for 65-year-old male patient had to be aborted due to abnormal vessel anatomy. the impella could not pass through the vessel and was only inserted into the graft. the physician switched to lvad (the left ventricular assist device) and rvad (the right ventricular assist device) combination. the physician noted a spontaneous unseen dissection in the aorta, not related to impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation into the delivery issues has been completed. abiomed products were not returned for evaluation. per the clinical review: the root cause of the delivery issue was most likely patient condition related, the impella implant had to be aborted due to abnormal vessel anatomy. f6 and f8 should have been blank. d4 corrected model number. h6 health impact code added, omitted in the initial. g1 manufacturing site name and address corrected.
25787 | 1220648-2024-11967 | 19458750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458750&pc=OZD | 08/24/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 08/24/2021 | 06/04/2024 |  |  | 06/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095400 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2021 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia and ventricular fibrillation. chest compression was started and endotracheal intubation was performed, and the procedure continued. the patient remained on impella support. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25788 | 1220648-2024-11968 | 19458755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458755&pc=OZD | 05/16/2024 | 07/24/2024 |  | 06/04/2024 |  |  | 05/16/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502752 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported that during impella cp implantation, the pump could not be started in the left ventricle (lv) and the impella cp had to be removed again. the procedure was successfully completed with another impella cp. there was no reported patient harm. | [5 empty] | The pump and purge cassette were returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for pump did not start has been completed. no data logs were returned. the product was returned and upon inspection the full red lumen was found in the pump. the pump was also able to start and run in the lab after removing the red lumen. the root cause of the pump did not start issue was determined to be a use-related red lumen issue. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11968. f6/f8-should have been left blank.
25789 | 1220648-2024-11969 | 19458777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458777&pc=OZD | 05/16/2024 | 07/23/2024 |  | 06/04/2024 |  |  | 05/16/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488887 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The complainant reported an unknown race and gender patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. following successful implant procedure, the impella cp withdrew automatically after a few minutes. suction alarm was displayed and flow was reduced from 3.6 l/min to 0.8 l/h. x-ray was performed and impella was noted to be too far into the aorta without alarm. the impella was explanted, even though wire was noted to be difficult to remove during original implant. the same impella was implanted again after flushing, but faced the same issue. the device would remain in the correct position, before pulling itself back into the aorta. the pump was planted and a new one was placed to resolve the issue. there was no patient harm reported. there were no visible thrombus or visual damages noted on the explanted pump. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for positioning issue has been completed. pump was returned for analysis. visual inspection found no issues on the pump. data was not returned for review. based on clinical description and product analysis, the root cause of positioning issue was most likely use related the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11969.
25790 | 1220648-2024-11970 | 19458795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458795&pc=OZD | 05/21/2024 | 09/25/2025 |  | 06/04/2024 |  |  | 05/22/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025490118 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | ECMO | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/22/2024 |  |  |  |  | The user facility reported a 59-year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient experienced two episodes of ventricular tachycardia/ventricular fibrillation requiring multiple drugs and three shocks on external defibrillator. the impella was pulled back 1 centimeter due to the pump being against the posterior wall. the patient was reported as stable. | The complainant also reported that the patient was on impella cp support when the aortic sensor had an issue. there were no readings after insertion. the representative attempted to fix the issue, however, it was reading 20 points lower than the arterial line. the staff monitored and support continued. |  |  |  |  | The investigation for the ventricular tachycardia has been completed. according to clinical details, on the fifth day of impella cp support, the patient also suffered a large seizure. this was followed by two episodes of vtach/vfib that required multiple drugs and three shocks on the external defibrillator. the pump was then repositioned due to it being placed against the posterior wall. the next day the patient was escalated to an impella 5.5 device. the impella was returned for evaluation. product evaluation confirmed the presence of no abnormalities. data log evaluation confirmed all 5s parameters were within normal spec. as the necessary information was not provided, the root cause of the vt/vf was not determined. the instructions for use for the impella cp with smartassist system lists ventricular arrythmia has an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. according to the clinical, shortly after the impella cp device was placed no aorta (ao) reading was being recorded. the representative was called in and fixed the ao reading only for it to be 20 points lower than the radial arterial line. on the fifth day of impella cp support, the patient also suffered a large seizure. this was followed by two episodes of ventricular tachycardia/fibrillation that required multiple drugs and three shocks on the external defibrillator. the pump was then repositioned due to it being placed against the posterior wall. the next day the patient was escalated to an impella 5.5 device. product evaluation confirmed the presence of no abnormalities. data log evaluation confirmed all 5s parameters were within normal spec. however, the initial offset value was extremely high. upon further investigation it was discovered that there was an optical sensor offset error with this pump. both the pumps before and after did not have the offset error. the cause of the optical signal issue was an offset error. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. update b5 to include optical signal issue description. update h6 to include optical signal issue. b7 updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
25791 | 1220648-2024-11971 | 19458826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458826&pc=OZD | 08/25/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 08/25/2021 | 06/04/2024 |  |  | 04/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022073734 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced bleeding from the insertion site. the angle of the impella insertion site was adjusted and hemostasis was acheived. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25792 | 1220648-2024-11973 | 19458855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458855&pc=OZD | 05/23/2024 | 07/22/2024 |  | 06/04/2024 |  |  | 05/23/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025494897 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | DISTAL PERFUSION CATHETER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 67-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the while on support the patient lost distal flow to their lower right extremity during impella cp support. a fem-fem bypass was subsequently placed with a distal perfusion catheter to treat the condition. support with the implanted pump continued following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.7 revised from the initial submission of manufacturer device report 1220648-2024-11973 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11973 and should not have been.
25793 | 1220648-2024-11972 | 19458860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458860&pc=OZD | 05/28/2024 | 10/01/2025 |  | 06/04/2024 |  |  | 05/28/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458189 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the use of the device cannot conclusively be associated with the outcome. | The user facility reported the patient implanted with an impella 5.5. device for mechanical circulatory support encountered high purge pressure. purge composition was changed, and staff continued to monitor. there was no report of patient harm as a result of this event. | The complainant also reported that there were placement signal not reliable alarms. the staff made sure the motor current and flow were operating well. support continued. |  |  |  | The investigation for the high purge pressure has been completed. pump was not returned for investigation. data logs were reviewed and it was seen that there was a gradual buildup in the purge pressure. the purge flow tick seen in the data logs were similar to that of a biomaterial buildup. the root cause of the high purge pressure issue was most likely biomaterial. added- b5 mso review, d6b d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were reviewed, and placement signal was seen to be lost after hours of the pump start. the root cause of the optical signal failure was an air gap from between the automated impella controller and the patient cable plug or within the middle of the red handle ferrule. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 updated to death. b5 updated to include placement signal issue description. h1 updated to death. h6 updated to include death and placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11972. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
25794 | 1220648-2024-11974 | 19458867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458867&pc=OZD | 12/20/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 12/20/2021 | 06/04/2024 |  |  | 07/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115176 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during support, hematuria was observed in the patient and hemolysis was diagnosed through tinged urine in the foley catheter. a bedside echocardiogram was employed to assess the position of the impella and the patient remained on impella support. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25795 | 1220648-2024-11975 | 19458883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458883&pc=OZD | 05/25/2024 | 07/24/2024 |  | 06/04/2024 |  |  | 05/25/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474715 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support developed multiple complications and subsequently expired. oozing at the access site was noted and no pulses were present in the right lower extremity with mottling. the care team decided to allow the oozing while reinforcing the dressing. in regard to the ischemia, the condition had reached the point where amputation was necessary. the patient's family subsequently decided to withdraw care planned for the following day (approximately two days after implant). the patient's status was changed to do not resuscitate. the patient expired shortly thereafter. it is unknown if the conditions contributed to the patient's demise. | [5 empty] | The investigation for the limb ischemia has been completed. the impella has not been returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the limb ischemia has not been determined. added-h6 component code 925 d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25796 | 1220648-2024-11976 | 19458921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458921&pc=OZD | 05/25/2024 | 07/12/2024 |  | 06/04/2024 |  |  | 05/25/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456784 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support after cardiac surgery. after the impella cp procedure, a hematoma and active bleeding were noted upon arrival to intensive care unit (icu). ultrasound performed and no pseudoaneurysm noted. pressure held in icu but unable to obtain hemostasis with anticoagulation and anti-platelet on board. the last known activated clotting time was 391 in cath lab. the decision was made to take the patient back to the cath lab and remove the right sided impella cp and switch to the left side access site for impella cp. vascular surgery present and attempted two perclose but was unsuccessful. used dry closure and manual pressure with a pressure dressing applied. the patient is stable. | [5 empty] | The investigation into the access site bleeding has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25797 | 1220648-2024-11977 | 19458936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458936&pc=OZD | 05/28/2024 | 11/20/2024 |  | 06/04/2024 |  |  | 05/28/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1106952 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The complainant reported that an automated impella controller experienced an unspecified purge error. the console was taken out of service because of the noted issue. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) purge issues has been completed. data logs were reviewed which show that console was unable to proceed with the purge system priming of impella rp. purge priming is then reset and attempted multiple times with the same result. real time logs show that there is less than 21 flow ticks over more than 30 seconds leading to the issuance of the error message. the console was returned and was visually inspected with no observable anomaly of the internal circuitry. there was no observable purge issue upon the boot up of the console. a known good pump and purge line was run with console at p-9 level with no observable problem. the root cause of the error was no problem found since the console behavior was within specification.
25798 | 1220648-2024-11978 | 19458938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458938&pc=OZD | 05/06/2024 | 06/25/2024 |  | 06/04/2024 |  |  | 05/06/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485094 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The user facility reported a 54-year-old female of unknown race in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding and developed a hematoma, which led to the heparin to be discontinued. it was further reported that at the conclusion of the insertion of the impella part of the graft was left outside of the body and gauze and tegaderm was applied. the decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding major was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25799 | 1220648-2024-11979 | 19458944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458944&pc=OZD | 05/28/2024 | 07/16/2024 |  | 06/04/2024 |  |  | 05/28/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025449994 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The user facility reported a 39-year-old male patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 access was gained through the sewn-on graft. it was reported that the pump was accidentally removed from the left ventricle (lv). the pump was explanted and not replaced. the patient was reported as stable. | [5 empty] | The investigation into the positioning issue has been completed. the abiomed products were not returned for evaluation. no data logs were available. the cause of the positioning issue most likely use related due to improper technique as the pump accidentally pulled out from the left ventricle. f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25800 | 1220648-2024-11980 | 19458994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19458994&pc=OZD | 05/08/2024 | 06/26/2024 |  | 06/04/2024 |  |  | 05/08/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456602 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was urgently implanted with an impella cp device for mechanical circulatory support as a rescue mission. the patient went into ventricular arrhythmia immediately after groin puncture. the patient immediately began decompensating requiring cardiopulmonary resuscitation and defibrillation. a temporary pacer was placed with no improvement in vitals. prior to the implant, the patient's left leg already had questionable flow. immediately following implant via the right femoral artery, both of the patient's legs were noted to be cold with poor color. the vascular surgeon opposed the use of a reperfusion cannula and the decision was made to withdraw care the following day. there is not enough information to associate the use of device with patient's outcome. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 component code was entered in error on the previously submitted report. no component code was needed. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25801 | 1220648-2024-11981 | 19459079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459079&pc=OZD | 12/20/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 12/20/2021 | 06/04/2024 |  |  | 09/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144864 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was excessively heparinized (systemically) and was bleeding from the intrajugular swan. heparin was discontinued, and the patient remained on impella support. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25802 | 1220648-2024-11982 | 19459150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459150&pc=OZD | 05/29/2024 | 06/20/2024 |  | 06/04/2024 |  |  | 05/29/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025488223 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the implant there was a failure alarm and the automated impella controller (aic) and impella were successfully replaced. | [5 empty] | Please see the corrected data field below for the rationale pertaining to the closure of this case. upon review of this case, it was determined that this is a duplicate report of an existing investigation. this case will be closed as a duplicate. please see mfr 1220648-2024-12521 for any and all information pertaining to this event and/or the investigation associated with the involved impella device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25803 | 1220648-2024-11983 | 19459214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459214&pc=OZD | 05/26/2024 | 06/12/2024 |  | 06/04/2024 |  |  | 05/29/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1325613 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) failed to recognize an installed purge cassette during patient support. the purge cassette was replaced to troubleshoot the issue, however, the failure remained. the aic was swapped with no harm to the patient. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. added-h6 impact code 4629 d4-corrected model number.
25804 | 1220648-2024-11984 | 19459245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459245&pc=OZD | 05/05/2024 | 07/03/2024 |  | 06/04/2024 |  |  | 05/05/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454008 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device intraoperatively for mechanical circulatory support. while on support, approximately three days post implant, it was noted the patient developed a gastrointestinal bleed of an unknown origin. the patient deteriorated, the decision was made to withdraw care, and the patient subsequently expired. medical safety review outcome noted there is not enough information to exclude the patient's demise from use of device. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. d4: corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25805 | 1220648-2024-11985 | 19459276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459276&pc=OZD | 05/21/2024 | 07/08/2024 |  | 06/04/2024 |  |  | 05/21/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025475861 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient had a distal perfusion cannula in the impella leg to help with flows. the patient was reported stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
25806 | 1220648-2024-11986 | 19459290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459290&pc=OZD | 08/26/2021 | 06/04/2024 |  | 06/04/2024 |  |  | 08/26/2021 | 06/04/2024 |  |  | 02/28/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050683 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient lost pulses in the right foot at completion of intervention. angiogram revealed decreased flow. the right common iliac was ballooned and stented and pulses returned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25807 | 1220648-2024-11987 | 19459364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459364&pc=OZD | 12/21/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 12/21/2021 | 06/04/2024 |  |  | 07/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108816 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma and bleeding at the access site. the patient received a blood transfusion and remained on impella support. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25808 | 1220648-2024-11988 | 19459377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459377&pc=OZD | 05/29/2024 | 07/22/2024 |  | 06/04/2024 |  |  | 05/29/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1078703 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the implant there was a failure alarm and the automated impella controller (aic) and impella were successfully replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The console was received for investigation. unable to reproduce the reported failure mode. performed power cycling of the console ten times, and each time, the console could run the multiple test optical pumps without any issue or impella defective alarm. no hardware issue was found on the console. the root cause of the impella defective alarm was not determined due to the inability to reproduce. the communication issue was most likely due to improper i2c clock setup caused by the impellatronic crystal. d9 added f6/f8-should have been left blank in the initial report.
25809 | 1220648-2024-11989 | 19459387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459387&pc=OZD | 08/31/2021 | 06/04/2024 |  | 06/04/2024 |  |  | 08/31/2021 | 06/04/2024 |  |  | 04/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072715 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/31/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. low pump flow was reported. the pump was explanted and a new impella was inserted. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25810 | 1220648-2024-11990 | 19459432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459432&pc=OZD | 09/01/2021 | 06/04/2024 |  | 06/04/2024 |  |  | 09/01/2021 | 06/04/2024 |  |  | 05/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080927 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient's right leg was cold and had no pulses. the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25811 | 1220648-2024-11991 | 19459460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459460&pc=OZD | 05/23/2024 | 07/15/2024 |  | 06/04/2024 |  |  | 05/23/2024 | 06/04/2024 |  |  | 05/31/2025 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351889 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 52-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed hemolysis. the impella was a little deep, but the physician felt good with the placement. the impella depth and hemolysis were discussed with the heart failure team and repositioning was declined. the plan was to wean and explant the impella cp. the patient was reported as stable. | [5 empty] | According to clinical, the patient transferred from a hospital in cardiogenic shock. lactate was trending down from 2.3 to 1.4. cardiac index up to 2.3 from 1.1 the prior day. the patient was on p-level p-4 due to hemolysis. device measuring 4.02cm in the left ventricle. impella depth and hemolysis discussed with heart failure team. repositioning was declined at that time with plan to wean and see if patient remains stable, intra-aortic balloon pump/pressors if needed. escalation to impella 5.5 was discussed but on hold due to patient being fully loaded on brilinta and current patient stabilization. potential durable left ventricular assist device in the future is being considered. the impella cp device was not returned, and therefore an evaluation of the device was not possible. however, data logs were received for evaluation/analysis which showed no motor current spikes nor trends in the placement signal or positioning issues. the root cause of the hemolysis cannot be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25812 | 1220648-2024-11992 | 19459471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459471&pc=OZD | 09/03/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 09/03/2021 | 06/04/2024 |  |  | 03/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced oozing at the access site and a nosebleed was noted. the patient was reduced to half dose heparin in purge. one pack of platelets was also administered to the patient. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25813 | 1220648-2024-11993 | 19459500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459500&pc=OZD | 04/03/2022 | 03/21/2025 |  | 06/04/2024 |  |  | 04/03/2022 | 06/04/2024 |  |  | 01/31/2024 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186241 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient continued to bleed after a double perclose cinch. manual pressure was held for fifteen (15) minutes with successful hemostasis achieved. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25814 | 1220648-2024-11994 | 19459506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459506&pc=OZD | 12/29/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 12/29/2021 | 06/04/2024 |  |  | 07/31/2022 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1499664 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that pulses were not palpable. heparin was added to the purge solution. the impella was successfully weaned and the device explanted. patient's condition was stable post procedure. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25815 | 1220648-2024-11996 | 19459561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459561&pc=OZD | 09/11/2021 | 05/22/2025 |  | 06/04/2024 |  |  | 09/11/2021 | 06/04/2024 |  |  | 07/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107309 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support it was reported that the pump stopped; however, the impella was successfully weaned. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25816 | 1220648-2024-11995 | 19459573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459573&pc=OZD | 05/29/2024 | 07/23/2024 |  | 06/04/2024 |  |  | 05/29/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480400 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The impella device was explanted three days after the reported event. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. while on support, a fem-fem external bypass needed to supply blood flow to distal limb. the peel away sheath was removed at time of insertion. the patient is stable. the patient remains on impella cp support on the date of this report. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. a.5 ethnicity was selected incorrectly on initial manufacturer device report 1220648-2024-11995. ethnicity is unknown. b.5 the device was explanted at the time of the initial manufacturer device report 1220648-2024-11995 and was reported incorrectly on that report. b.7 added information as it was inadvertently omitted from initial manufacturer device report 1220648-2024-11995. d.6b revised from the initial submission of manufacturer device report 1220648-2024-11995 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report 1220648-2024-11995 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11995 was submitted. h.6 added code 3038 since initial submission of manufacturer device report 1220648-2024-11995 in accordance with updated procedures. h.6 codes 4114 and 4624 were reported incorrectly on the previously submitted manufacturer device report 1220648-2024-11995. new codes has been added to type of investigation codes and health effect - impact codes. h.10 additional manufacturer narrative on the previously submitted manufacturer device report 1220648-2024-11995 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿
25817 | 1220648-2024-11997 | 19459605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459605&pc=OZD | 09/15/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 09/15/2021 | 06/04/2024 |  |  | 06/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105992 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported the patient had an episode of ventricular tachycardia and was defibrillated. the patient was stabilized. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25818 | 1220648-2024-11998 | 19459609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459609&pc=OZD | 05/24/2024 | 07/25/2024 |  | 06/04/2024 |  |  | 05/24/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469076 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team had difficulty placing the impella cp due to kinking. the team replaced the impella cp with another with no issues. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues has been completed. the impella was returned for evaluation. the returned device was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related; the cannula was unable to pass through the introducer and unable to pass through the patient's vasculature. added-d9 device return date, h6 med dev prod code 2682 f6/f8 should have been left blank in the initial report.
25819 | 1220648-2024-11999 | 19459610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459610&pc=OZD | 03/04/2022 | 03/21/2025 |  | 06/04/2024 |  |  | 03/04/2022 | 06/04/2024 |  |  | 11/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166869 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced oozing from the bicuspid aortic valve (bav) contralateral groin site overnight. it was noted that heparin in purge was placed on hold. it was later noted that the bleeding resolved, no intervention mentioned further. additionally, the patient experienced limb ischemia, indicated by a ¿cool and cyanotic¿ right lower extremity, with no indication of intervention required to resolve. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25820 | 1220648-2024-12000 | 19459652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459652&pc=OZD | 09/16/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 09/16/2021 | 06/04/2024 |  |  | 05/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090959 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2021 |  |  |  |  | The user facility reported a patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient's urine was pink tinged. the p level was decreased and the position was confirmed under echocardiography. the device was weaned and explanted at beside. no further information was provided. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volume may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25821 | 1220648-2024-12001 | 19459660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459660&pc=OZD | 05/29/2024 | 07/31/2024 |  | 06/04/2024 |  |  | 05/29/2024 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. during the planned explant of the impella cp pump, a dissection was observed in the patient's vasculature. leading up to the discovery, an angiogram had revealed limited flow distal to the external iliac artery and distal pulses were no longer dopplerable. surgical intervention was required to provide adequate flow to the leg. | [5 empty] | The investigation for vascular damage - peripheral vessel issue has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage - peripheral vessel issue was not determined since the product was not returned and insufficient clinical information was provided. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-12001 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-12001 was submitted. h.6 code 925 was added in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-12001. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25822 | 1220648-2024-12002 | 19459685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459685&pc=OZD | 09/22/2021 | 06/04/2024 |  | 06/04/2024 |  |  | 09/22/2021 | 06/04/2024 |  |  | 05/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080838 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported the patient¿s legs were mottled and distal pulses were absent. nitropress was started. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25823 | 1220648-2024-12003 | 19459722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459722&pc=OZD | 04/04/2022 | 03/21/2025 |  | 06/04/2024 |  |  | 04/04/2022 | 06/04/2024 |  |  | 01/31/2024 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183814 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported that five percloses were deployed, where two finally caught. bleeding was still present after perclose singed down. an 8f angioseal was placed with manual pressure for twenty minutes. three percloses failed prior to two being fully utilized. the impella cp device was weaned and explanted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25824 | 1220648-2024-12004 | 19459755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459755&pc=OZD | 08/17/2021 | 03/21/2025 |  | 06/04/2024 |  |  | 08/17/2021 | 06/04/2024 |  |  | 05/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022086451 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, perclose x 3 deployed without success in achieving hemostasis. the decision was made to consult vascular for surgical closure due to angle of entry and low stick during intra-aortic balloon pump insertion. the impella was weaned successfully at the conclusion of the case. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25825 | 1220648-2024-12005 | 19459866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19459866&pc=OZD | 04/05/2022 | 03/24/2025 |  | 06/04/2024 |  |  | 04/05/2022 | 06/04/2024 |  |  | 01/31/2024 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABOIMED, INC. |  |  | US |  | 2023184897 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was bleeding at a rapid pace. heparin was discontinued and replaced with bicarbonate. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25826 | 1220648-2024-12006 | 19460092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19460092&pc=OZD | 04/10/2022 | 03/24/2025 |  | 06/04/2024 |  |  | 04/10/2022 | 06/04/2024 |  |  | 01/31/2024 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185866 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2022 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the patient¿s platelets and hemoglobin were low, and the patient was transfused one (1) unit of packed red blood cells. the patient¿s heparin-induced thrombocytopenia panel was returned with positive results, and hemolysis was also identified. it was also reported that heparin administration was removed from the purge system and converted to dextrose 5% in water. the patient remained on impella support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25827 | 1220648-2024-12007 | 19460324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19460324&pc=OZD | 09/30/2021 | 06/04/2024 |  | 06/04/2024 |  |  | 09/30/2021 | 06/04/2024 |  |  |  |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1466766 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/30/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. following support, the device was successfully weaned and explanted. the impella site was closed with two perclose and with a 6fr angio seal. some oozing remained. femstop and manual pressure were also applied and good hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25828 | 1220648-2024-12008 | 19460540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19460540&pc=OZD | 10/04/2021 | 03/24/2025 |  | 06/04/2024 |  |  | 10/04/2021 | 06/04/2024 |  |  | 05/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088854 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported that after initial stabilization and weaning of extracorporeal membrane oxygenation, the patient showed massive convulsions. cardiac computed tomography was performed and showed a large hemorrhage in the brain. therapy was frozen according to will of relatives. | [5 empty] | Explant date is unknown: to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
25829 | 1220648-2024-12009 | 19464025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19464025&pc=OZD | 08/17/2021 | 03/24/2025 |  | 06/04/2024 |  |  | 08/17/2021 | 06/04/2024 |  |  | 03/31/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a clot was detected. intervention was stopped and plan was made to rest heart on impella and aggrastat was used to dissolve clot. the patient was successfully weaned off of support. | [5 empty] | When making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: thrombus impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
25830 | 1220648-2024-12010 | 19465101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19465101&pc=OZD | 09/26/2022 | 04/07/2025 |  | 06/04/2024 |  |  | 09/26/2022 | 06/04/2024 |  |  | 04/30/2024 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226371 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/26/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced a stroke. support was withdrawn as the patient was not a candidate for advanced therapies due to the stroke. |  |  |  |  | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25831 | 1220648-2024-12011 | 19465203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19465203&pc=OZD | 08/08/2021 | 03/24/2025 |  | 06/04/2024 |  |  | 08/08/2021 | 06/04/2024 |  |  | 06/30/2023 |  | 06/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s myocardium was extremely irritable, so patient went into ventricular fibrillation (vf). a decision was made to remove the impella cp and replace intra-aortic balloon pump in order to keep the patient out of vf. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: arrhythmia impella cp impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
25832 | 1220648-2024-11918 | 19446372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446372&pc=OZD | 05/22/2024 | 07/10/2024 |  | 06/03/2024 |  |  | 05/22/2024 | 06/03/2024 |  |  | 09/30/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025477041 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Intraoperative Cardiac Injury | [11 empty] | Death | Required Intervention |  |  |  |  | INTRA-AORTIC BALLOON PUMP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery for ventricular septal defect was implanted with an impella 5.5 device for mechanical circulatory support during impella 5.5 placement, the physician noted on echo that guidewire crossed a ventricular septal defect into the right ventricle. the guidewire was repositioned into the left ventricle. the impella was inserted and ramped up. shortly after initiation, pulse pressure narrowed. physician noted effusion on echocardiogram. the physician suspected right ventricle perforation from the guidewire. the physician decided no intervention was needed. the patient is stable. the patient remains on impella 5.5 support on the date of this report. | It was further reported that the the care was withdrawn and the patient expired at explantation of the device. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough information to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported ventricle perforation has been completed. the product was not returned. based on the provided clinical details, the root cause of the ventricle perforation was most likely due to guidewire puncture due to improper technique. added-d6b, h6 component code 925 b2-selected death and added patient date of death b5-added mso review. f6/f8 should have been left blank in the initial report. h1-type of reportable event changed to death. added h6-impact code 1802.
25833 | 1220648-2024-11919 | 19446524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446524&pc=OZD | 05/23/2024 | 07/17/2024 |  | 06/03/2024 |  |  | 05/23/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480676 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The complainant reported a 47-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient did not have a history of heart disease but developed a groin bleed which caused a hematoma to develop. manual pressure was held, and the decision to explant the impella due to infection risk was made. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. f6/f8 should have been left blank in the initial report. g1 reporting contact fax number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25834 | 1220648-2024-11920 | 19446569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446569&pc=OZD | 05/24/2024 | 07/18/2024 |  | 06/03/2024 |  |  | 05/24/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494951 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 30-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and they had access site bleeding. two units of blood products were provided, and manual pressure was implemented. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25835 | 1220648-2024-11921 | 19446588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446588&pc=OZD | 05/21/2024 | 07/09/2024 |  | 06/03/2024 |  |  | 05/21/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505495 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 79-year-old female with a high risk of percutaneous coronary intervention shock was implanted with an impella cp for mechanical circulatory support. it was reported that the left axillary artery was dissected while implanting the cp via the left axillary/subclavian artery. the physician placed two covered stents to intervene. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage has been completed. the impella was returned for investigation. there were no physical abnormalities observed on the pump. the patient had diseased vessels which were found to be damaged during insertion. the insertion was smooth with no excessive force applied. per the clinical information provided, the root cause of the vascular damage - peripheral vessel issue was patient condition related to stenosis of the vessel near the insertion site. added- h6 impact code 4642 f6/f8 should have been left blank in the initial report 1220648-2024-11921-1.
25836 | 1220648-2024-11922 | 19446627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446627&pc=OZD | 05/26/2024 | 07/29/2024 |  | 06/03/2024 |  |  | 05/26/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458201 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump stopped unexpectedly during patient support. it was noted that the motor current on the pump spiked immediately prior to the stoppage. the pump was subsequently replaced as a result of the noted failure. there was no patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. pump was returned for evaluation. electrical testing was performed and all three motor cable lines were found to be measuring as open lines. the red handle was opened and all three motor cable lines were found to be broken. two of the motor cable lines were broken due to necking at the monkey fist and one motor cable line was completely snapped near the solder joint. additionally, the ground line was snapped completely with monkey fist fully separated from kink protector. data logs were reviewed and a motor current spike to 3000 with alarm #115 occurred followed by a sudden drop in motor current, motor speed, mean flow, and p-level to 0. clinical details did not note any patient movement at time of pump stop. the root cause of the pump stop is an electrical open with breaking of all motor cable lines due to excessive force. d4: corrected model number f6/f8 should have been left blank in the initial report. | The impella device was received from the customer on (b)(6)2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25837 | 1220648-2024-11923 | 19446692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446692&pc=OZD | 05/26/2024 | 08/01/2024 |  | 06/03/2024 |  |  | 05/26/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1692213 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Computer Software Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | The user facility reported a 52-year-old male patient of unknown ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) experienced a controller failure during patient support. the aic was requested for return for service evaluation. there was no patient harm. | [5 empty] | Data analysis: pump was run for 18 days before the user received an impella stopped alarm. the alarm never resolved for the next 7 hours until the pump was removed from the console. the lt logs confirm the motor current spiked before dropping to 0ma after the alarm triggered and motor speed and pump flow dropped to 0 as well. device analysis: the console was tested upon arrival, but the console performed as expected and had no problems operating at the p-level from the case (4) as well as 1 and 9. root cause: there was no issue found with the console that would have caused the pump stop. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11923. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25838 | 1220648-2024-11924 | 19446722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446722&pc=OZD | 05/21/2024 | 07/17/2024 |  | 06/03/2024 |  |  | 05/21/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. | [8 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a patient of unknown demographics was implanted with an impella 5.5 as a bridge to transplant. it was reported that the patient suffered an embolic event. the patient had a known history of embolic events prior to the recent occurrence. support with the pump was maintained despite the embolic event. the relationship between the pump and the noted condition remains unknown. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. f6/f8 should have been left blank in the initial report. | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
25839 | 1220648-2024-11925 | 19446747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19446747&pc=OZD | 08/17/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/17/2021 | 06/03/2024 |  |  | 06/30/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022096509 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the impella was prepped, however, due to the patient¿s anatomy it was not possible to place the sheath in the groin access site due to kinking. therefore, the procedure was aborted, and the intervention was performed without impella support. also, surgical wound closure was required to resolve site bleeding and achieve hemostasis. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25840 | 1220648-2024-11926 | 19447035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447035&pc=OZD | 08/26/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/26/2021 | 06/03/2024 |  |  | 06/30/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that impella was completely occlusive. the sheath was peeled away and it was still occlusive. the patient was placed on levo and the pump was removed. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25841 | 1220648-2024-11927 | 19447038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447038&pc=OZD | 08/17/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/17/2021 | 06/03/2024 |  |  | 03/31/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022065084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a large hematoma noted at the access site. manual compression applied and blood products were given. additionally, it was reported that the patient had post ventricular tachycardia in the cath lab after the impella cp was weaned and removed. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25842 | 1220648-2024-08541 | 19447069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447069&pc=OZD | 07/08/2022 | 03/21/2025 |  | 06/03/2024 |  |  | 07/08/2022 | 06/03/2024 |  |  | 08/31/2024 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189821 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a left ventricular thrombus. device was explanted to resolve the reported issue. patient's condition was stable post procedure. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (thrombus) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
25843 | 1220648-2024-11928 | 19447084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447084&pc=OZD | 08/27/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/27/2021 | 06/03/2024 |  |  | 05/31/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083040 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. three cardiac wounds caused tamponade and required repair. one cardiac wound was from perforation and two were from pericardiocentesis needle entering the right ventricle. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
25844 | 1220648-2024-11929 | 19447097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447097&pc=OZD | 05/22/2024 | 07/23/2024 |  | 06/03/2024 |  |  | 05/22/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2025479380 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient experiencing acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. twelve days after the impella cp pump was implanted, the pump stopped and was therefore explanted. there was no report of harm to the patient. | [5 empty] | The investigation into the pump stop has been completed. no product was returned for evaluation; however data logs were made available. the data logs confirm 'impella stopped - motor current high' pump stop. there were suction alarms, rising purge pressure, and a motor current spike before pump stop occurred. the root cause of the pump stop was most likely biomaterial ingestion, patient condition. d4 corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25845 | 1220648-2024-11384 | 19447120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447120&pc=OZD | 11/12/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 11/12/2021 | 06/03/2024 |  |  | 04/30/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071587 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient experienced loss of pulses in right arm. physician placed brachial sheath in the right arm and then did an external bypass from the sheath to the side part of the impella. pulses returned and flow restored. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (limb ischemia) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25846 | 1220648-2024-11930 | 19447183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447183&pc=OZD | 05/19/2024 | 07/29/2024 |  | 06/03/2024 |  |  | 05/18/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025459470 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Patient-Device Incompatibility | Device Damaged by Another Device | [5 empty] | Aortic Valve Insufficiency/ Regurgitation | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [5 empty] | MEDIKIT SHEATH | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information: after two days of support, the patient developed aortic regurgitation and the decision was made to explant the pump. extracorporeal membrane oxygenation was subsequently placed for ongoing support. | The user facility reported a patient with ischemic cardiomyopathy was implanted with an impella cp for mechanical circulatory support; however, delivery difficulty was encountered. upon insertion of the impella cp pump up to the left ventricle, it was noted the pump was in the wrong position and was pulled back into the aorta. the pump was torqued and re-advanced, but the shaft and wire became entangled, stuck making it impossible to proceed beyond the ascending aorta. the pump and wire were subsequently removed entirely, and a new wire was inserted for delivery. the introducer was then replaced with a 16fr medikit sheath, but the medikit sheath came off while positioning for the pump was being established. the delivery was eventually successful, and it was noted the impella cp pump worked without any issues. there was no report of harm to the patient. |  |  |  |  | Additional information has been received and following fields have been updated: b1. adverse event selected. b2. other serious (important medical events) selected. b5: additional information added. g3 date received by manufacturer updated. h1 type of reportable event. h6 health effect-clinical code. h6-health effect-impact code. h10 instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smart assist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for the heart valve damage and the delivery issues has been completed. the impella device was not returned for evaluation. data logs were not relevant to the failure mode. clinical details states that during insertion and positioning of impella in the lv, the 0.018" became stuck on the pump because the physician used torque to try and advance both back into the lv. the root cause of the delivery issue was use related because the physician attempted to advance using torque. additionally, trying to advance the pump when it is tangled with the guidewire is a use issue. because of the limited clinical notes and no product nor data logs were returned, the root cause of the heart damage cannot be determined. d4-corrected model number.
25847 | 1220648-2024-11931 | 19447205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447205&pc=OZD | 05/20/2024 | 07/12/2024 |  | 06/03/2024 |  |  | 05/20/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471120 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | It was further reported that the patient developed acute infection/abscess at r axillary surgical insertion site. elevated white blood cell (wbc) and exudates from site noted to develop over a 24 hour period. the patient was treated with a wound vacuum and eventually had an artery repair of the impella site. | The complainant reported a patient (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the device was placed, but the access site was saturated with blood. the dressing was adjusted and redressed and anticoagulation was removed. the patient remains on impella 5.5 support on the date of this report. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped to achieve hemostasis as per the clinical description. without additional clinical details, the root cause of the infection/sepsis cannot be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock state the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ added- b5 additional information, d6b, h6 impact codes 4642 and 4625 f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25848 | 1220648-2023-04670 | 19447206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447206&pc=PYX | 09/28/2023 | 09/09/2024 |  | 06/03/2024 |  |  | 09/28/2023 | 06/03/2024 |  |  | 08/31/2024 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289354 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 57-year-old male was implanted with an impella rp device for mechanical circulatory support. the complainant reported that the impella was on flows that were lower than recommended, and during this time the pump stopped. this was during a planned removal, so restarting the impella was not attempted. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. product evaluation confirmed no abnormalities with the exterior of the pump. however, once a teardown was performed retainer damage was found. this was confirmed via computed tomography (ct) scan imaging. data log analysis confirmed a high motor current pump stop occurred at the end of the case. the cause of the pump stop was bearing wear. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04670 d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h.6 code 4625 and 4627 were reported incorrectly. new code was added to health effect - impact code.
25849 | 1220648-2023-04666 | 19447207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447207&pc=OZD | 09/28/2023 | 06/05/2024 |  | 06/03/2024 |  |  | 09/28/2023 | 06/03/2024 |  |  | 05/31/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352069 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old female was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient¿s leg was ischemic. the impella was escalated from an impella cp to an impella 5.5. upon cutdown of the right femoral artery (rfa) and impella explanted, a dissection plane was found in the external illiac requiring stent. flow was restored down to the leg once impella was removed. | [5 empty] | The investigation into the limb ischemia ¿ reversible and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned for investigation. as per clinical findings, the root cause of the limb ischemia - reversible was most likely patient condition related. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
25850 | 1220648-2023-04669 | 19447208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447208&pc=OZD | 09/28/2023 | 10/01/2025 |  | 06/03/2024 |  |  | 09/28/2023 | 06/03/2024 |  |  | 01/31/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310018 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | Positioning Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Patient was transitioned to hospice care and the decision to withdraw care was made. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that the impella 5.5 was placed after left ventricle (lv) access was obtained using an angled glide wire and marked peripheral catheter. then in an attempt to optimize placement, placement was lost. access to lv was lost after peel away sheath was removed. after witless attempts using fluoroscopy and tee were unsuccessful and the doctor opted for sponge plug technique. this led to access using a 7fr sheath and heavy ties, then again using the same catheters and wires access to lv was obtained, and sponge plugs were attached to impella. impella was advanced, however wore access would be lost again. again, steps were completed with access to lv, sponge plug attached this time with success. unfortunately, blood products were needed post placement so that impella performance could be increased. impella was secured in place and patient was sent to icu. in addition, despite giving 10k units of heparin and 5k additional, a lv thrombus was noted. additionally, anticoagulants were halted due to access site bleeding. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04669 was provided in sections b2, b5, d6 h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding, thrombus, positioning issue, and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: the patient experienced oozing at the insertion site of the impella pump. the root cause of the access site bleeding was most likely use related due to repositioning to optimize placement. thrombus: the root cause of the thrombus was unable to be determined due to insufficient clinical details. positioning issue: the root cause of the positioning issue was most likely improper technique removing the sheath resulting in the pump to come out of the lv. low pump flow: the root cause for the low pump flow was most likely patient condition based on clinical details and data log analysis. g1 reporting contact email revised on manufacturer device report 1220648-2023-04669 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04669. h6 medical device problem code 1248 and 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04669. h6 medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2023-04669.
25851 | 1220648-2024-11932 | 19447212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447212&pc=OZD | 05/22/2024 | 07/17/2024 |  | 06/03/2024 |  |  | 05/22/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493200 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a 64-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the pump came out of the appropriate position in the left ventricle due to a tuohy valve failing to lock. the patient remained on impella support for 40 hours, after which the pump was explanted and not replaced. the patient remains on medical support. | [5 empty] | The investigation for the positioning issue has been completed. the pump was not returned for investigation. data logs were reviewed which show that the pump ran for 21 hours without any reported issues regarding positioning. according to the clinical report, at the start of the case, the pump was pulled back into the aorta after the patient moved while the touhy-borst was reported to be unsecured. the doctor was unsuccessful in tightening the touhy-borst multiple times. the cause of the positioning issue was most likely touhy-borst related, based on given clinical details. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25852 | 1220648-2024-11933 | 19447219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447219&pc=OZD | 08/27/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/27/2021 | 06/03/2024 |  |  | 06/30/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102695 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the staff was unable to find a pulse in the left leg and the pump was explanted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
25853 | 1220648-2024-11934 | 19447254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447254&pc=OZD | 08/18/2021 | 06/03/2024 |  | 06/03/2024 |  |  | 08/18/2021 | 06/03/2024 |  |  | 05/31/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080537 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, suction alarms and the impella was at p1 due to ventricle and aorta alarms were noted. the pump was repositioned. the decision was made to remove the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25854 | 1220648-2024-11935 | 19447255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447255&pc=OZD | 05/21/2024 | 07/18/2024 |  | 06/03/2024 |  |  | 05/21/2024 | 06/03/2024 |  |  | 09/30/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024415015 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 58-year-old female patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the patient experienced access site bleeding. as treatment a pocket revision was performed and a jackson pratt (jp) drain was placed, with an extra stitch added later. the patient remains on support and is reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. clinical details were sufficient to determine root cause. the root cause of the access site bleeding was most likely use related as the bleeding was resolved by revising the pocket. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25855 | 1220648-2024-11936 | 19447263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447263&pc=OZD | 05/28/2024 | 07/21/2024 |  | 06/03/2024 |  |  | 05/28/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1687316 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 05/28/2024 |  |  |  |  | The complainant reported the automated impella controller (aic) was unable to move through additional prompts to finish the purge cassette change. the purge cassette/disc was unrecognizable. the aic was replaced. no patient harm has been reported. | [5 empty] | A.1, a.4 and a.5 are unknown. the investigation into the aic - purge disc/flag issues has been completed. review of the logs confirmed the reported issue that a purge disc could not be detected. although it was reported to have occurred on ((b)(6) 2024), evidence in the imc logs suggest that it occurred on ((b)(6) 2024) and several pages of the data sheet contained multiple purge disk not detected alarms. the purge disc flag was returned and evaluated. it was found to be broken (missing at blue retainer). the root cause of the issue was a broken purge disc flag. a.3 revised as selection was inadvertently omitted from manufacturer device report 1220648-2024-11936. h.10 a.1, a.4 and a.5 are unknown statement was inadvertently omitted from manufacturer device report 1220648-2024-11936. d.2 common device name was revised as previously entered incorrectly on manufacturer device report 1220648-2024-11936. d.4 lot number was revised as previously entered incorrectly on manufacturer device report 1220648-2024-11936. d.9 revised as previously entered incorrectly on manufacturer device report 1220648-2024-11936. f.6 and f.8 dates were reported on manufacturer device report 1220648-2024-11936 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report 1220648-2024-11936 was submitted. g.2 revised selections as they were inadvertently omitted from manufacturer device report 1220648-2024-11936. h.3 if the device was not evaluated was entered incorrectly on manufacturer device report 1220648-2024-11936 and should of reflect device evaluation anticipated, but not yet begun. h.6 added code 4629 as it was inadvertently left off manufacturer device report 1220648-2024-11936. h.6 codes 4114 was reported incorrectly on manufacturer device report 1220648-2024-11936. the device was returned already. h.10 additional manufacturer narrative on manufacturer device report 1220648-2024-11936 stated that the device was not returned but the device was already returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H.3 device not returned to manufacturer for evaluation on original manufacturer device report 1220648-2024-11936 was selected incorrectly and should of reflected that the device was returned. h.6 codes 4121, 4112 and 4110 were inadvertently left off manufacturer device report 1220648-2024-11936 follow-up number 1 and were added to this report.
25856 | 1220648-2024-11937 | 19447272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447272&pc=OZD | 08/27/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/27/2021 | 06/03/2024 |  |  | 06/30/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108228 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into v-fib. the patient was shocked four times and was given cpr. additionally, the patient developed bleeding at the groin access site. pressure was applied, but bleeding continued. all best practices were attempted, but were not successful. the decision was made to wean patient off of support. weaning was successful and the procedural outcome was that the patient survived surgery. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ventricular fibrillation. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25857 | 1220648-2024-08537 | 19447286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447286&pc=OZD | 07/07/2022 | 03/10/2025 |  | 06/03/2024 |  |  | 07/07/2022 | 06/03/2024 |  |  | 03/31/2024 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023208096 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 56-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient did have to undergo cut-down to the lateral circumflex femoral artery (lcfa) post operation and a large hematoma was noted at the 5fr access site; which a cut-down was performed for hematoma evacuation and surgical closure of the artery. however, the patient received 4 units of packed red blood cells ( prbcs), due to the hemoglobin level dropped and the blood loss was noted in left groin upon drape removal and surgical cut-down was performed for repair. surgical cutdown was performed and the device was explanted without any issues and the artery was closed and repaired. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events : acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vascular damage - peripheral vessel issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
25858 | 1220648-2024-11938 | 19447288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447288&pc=OZD | 05/27/2024 | 07/18/2024 |  | 06/03/2024 |  |  | 05/28/2024 | 06/03/2024 |  |  | 06/30/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367294 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Fluid/Blood Leak | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The complainant reported a 15-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient had a history of arrythmia issues and ectopy, and the pump was repositioned due to the frequency of ectopy. purge pressure alarms occurred, and medical staff found that the yellow-to-yellow connection on the purge sidearm was wet. staff then replaced the purge cassette and tubing to resolve the issue. the patient remains on support. | [5 empty] | The investigation for vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. h.6 codes 1250 and 925 were reported incorrectly on manufacturer device report 1220648-2024-11938. new code was been added to medical device problem code. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
25859 | 1220648-2024-11939 | 19447322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447322&pc=OZD | 05/29/2024 | 07/11/2024 |  | 06/03/2024 |  |  | 05/29/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025504352 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The user facility reported a 70-year-old male patient of unknown race in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that pump stopped after 7 days of support and was explanted. the patient was not escalated as the patient had other non-cardiac issues. no patient harm was reported. | [5 empty] | The investigation for the pump stop has been completed. no pump was returned for investigation. data logs show a temporary pump stop with alarm 115 pump stop rpm and alarm 4 disconnected while running with mc and ps spikes. data logs do not show any other pump stops or pump stop alarms in the rest of the case including on the reported date. the root cause of the pump stop was not determined since product was not returned and data logs are inconclusive. f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: impella stopped: if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
25860 | 1220648-2024-11940 | 19447382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447382&pc=OZD | 05/29/2024 | 07/23/2024 |  | 06/03/2024 |  |  | 05/29/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474061 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. on day nine of the support, the team observed pus and redness at the impella 5.5 axillary site. the day prior, the jackson pratt drain and chest tubes had been removed. the team treated the patient with antibiotics. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The investigation for infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added a5 ethnicity. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smart assist system with automated impella controller & rp flex with.
25861 | 1220648-2024-11941 | 19447456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447456&pc=PYX | 05/27/2024 | 07/22/2024 |  | 06/03/2024 |  |  | 05/27/2024 | 06/03/2024 |  |  | 07/31/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024374138 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp device for mechanical circulatory support. the impella rp placement was aborted due to the impella rp being unable to pass through the introducer. the physician arranged venoarterial extracorporeal membrane oxygenation for the patient who was transported to another user facility. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely delivery issue since they were unable to advance rp into the sheath. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
25862 | 1220648-2024-11942 | 19447468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447468&pc=OZD | 08/19/2021 | 06/03/2024 |  | 06/03/2024 |  |  | 08/19/2021 | 06/03/2024 |  |  | 05/31/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093827 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/19/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was shocked after an episode of ventricular tachycardia and had suction alarms with decreased pressures. the pump speed level was set to p8, then the patient was taken for escalation of care to the impella 5.5. the impella cp was explanted, and the procedural outcome is the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records.
25863 | 1220648-2024-11943 | 19447469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447469&pc=OZD | 05/29/2024 | 07/25/2024 |  | 06/03/2024 |  |  | 05/29/2024 | 06/03/2024 |  |  | 09/30/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/29/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the team observed sheath damage/leak. the team replaced the sheath. the pump supported until percutaneous coronary intervention completed and pump explanted and groin closed. field representative reported that the hg drop prompted three units of blood to be given. | [5 empty] | The investigation for introducer damage ¿ mechanical and access site bleeding has been completed. upon visual inspection, no abnormalities were noted on the sheaths. pressure thresholds of both returned sheaths were measured against a control 14fr sheath from the lab. the control sheath was measured at 220.5 mmhg. the 14fr x 25cm sheath measured at 64.1mmhg and the 14fr x 13cm sheath measured at 156.1mmhg. 14fr introducers are tested to 180mmhg per dtm 0046-9910. clinical details noted that bleeding was occurring from the impella access site and was not able to be resolved with wedging towels under the introducer hub. additionally, the 14fr13cm and 25cm introducers were leak tested by physician and leaking was occurring from introducer hub when distal end of introducer was plugged, and introducer was flushed. the root cause of the access site bleeding was most likely due to a damaged valve on the introducer hub. the root cause of the introducer damage was most likely due to valve damage on the introducer hub. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11943: f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25864 | 1220648-2024-11894 | 19447476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447476&pc=OZD | 09/15/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 09/15/2021 | 06/03/2024 |  |  | 01/31/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022033535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had oozing around the repo sheath with the bleeding managed at bedside. the oozing stopped and active clotting time (act) was elevated. the patient was successfully weaned and the device explanted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25865 | 1220648-2024-11945 | 19447665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447665&pc=OZD | 05/27/2024 | 07/11/2024 |  | 06/03/2024 |  |  | 05/27/2024 | 06/03/2024 |  |  | 07/31/2025 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024388832 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The user facility reported a 63-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that patient had no distal flow noted on the right femoral artery. the device was explanted, and a clot was noted. the patient was escalated to an impella 5.5 due to limb ischemia. the patient was noted as stable. | [5 empty] | The investigation for the reported thrombus and limb ischemia has been completed. the device nor additional information was provided for evaluation. therefore, the root cause of the limb ischemia and thrombus could not be determined. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
25866 | 1220648-2024-11946 | 19447747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19447747&pc=OZD | 09/08/2022 | 03/21/2025 |  | 06/03/2024 |  |  | 09/08/2022 | 06/03/2024 |  |  | 07/31/2023 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121417 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that there was oozing at the access site and blood in the urine. heparin was withheld and echo was ordered to assess pump position. additionally, the patient experienced an episode of ventricular tachycardia. the patient was shocked once with resolution. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
25867 | 1220648-2024-11947 | 19448685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19448685&pc=OZD | 05/22/2024 | 07/30/2024 |  | 06/03/2024 |  |  | 05/22/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493208 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The user facility reported a 70-year-old female with a high risk of percutaneous coronary intervention and a need for a transcatheter aortic valve replacement (tavr) was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp failed delivery due to the 14 french sheath having been kinked. the sheath was removed and replaced, and the impella cp was successfully delivered. the impella cp was explanted three hours after the tavr procedure. no patient harm has been reported and the patient survived the event. | [5 empty] | D.9 updated as the introducer was returned for investigation. the investigation for the introducer damage. mechanical has been completed. impella pump was not returned. the 14 french introducer long sheath (25 centimeters) was returned for investigation. the introducer was visually inspected and sheath kink was observed on the sheath body. no other damages or abnormalities were found. the root cause of the introducer damage issue was an introducer sheath kink due to likely improper technique. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11947 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11947 was submitted. | The impella device was not received from the customer and thereforean evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25868 | 1220648-2024-08146 | 19448900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19448900&pc=OZD | 08/06/2022 | 03/21/2025 |  | 06/03/2024 |  |  | 08/06/2022 | 06/03/2024 |  |  | 02/29/2024 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196993 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. a balloon tamponade procedure was attempted but failed. a surgical cutdown was performed to resolve the issue. the patient's condition was stable post procedure. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25869 | 1220648-2024-11948 | 19448966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19448966&pc=OZD | 05/20/2024 | 07/30/2024 |  | 06/03/2024 |  |  | 05/23/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025459886 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was placed via the non-abiomed approved 16 french medikit sheath in the femoral for a patient admitted in with a needed high risk percutaneous cardiac insertion¿(hrpci). the pump was placed but the limb became ischemic and so the team placed an antegrade perfusion. the patient was additionally supported by extracorporeal membrane oxygenation (ecmo). | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4 corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25870 | 1220648-2024-11949 | 19449041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19449041&pc=OZD | 05/23/2024 | 07/24/2024 |  | 06/03/2024 |  |  | 05/23/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482243 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. an impella cp was being placed for support of a patient admitted in with cardiac comorbidities and a plan for a percutaneous coronary intervention. the pump was placed but the team found the sheath to be malfunctioning. the valve allowed for bleeding and so the valve/sheath was replaced with a new abiomed second peel away sheath. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the introducer damage has been completed. the introducer was not returned for investigation. the cause of the introducer damage issue was not determined since product was not returned. f6/f8 should have been left blank in the initial report.
25871 | 1220648-2024-11950 | 19449157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19449157&pc=OZD | 05/23/2024 | 07/10/2024 |  | 06/03/2024 |  |  | 05/23/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495285 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 45-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient¿s labs have been hemolyzing. after several attempts to reposition the pump, the decision was made to replace the impella cp with an intra-aortic balloon pump. the patient remained hemodynamically stable. | [5 empty] | The product was not returned for investigation. hemolysis was reported. data logs were provided finding no motor current spikes or major positioning-related issues were noted in the data logs. as per the clinical notes, the pump was explanted due to hemolysis and the pump was confirmed to be malpositioned at the time. the cause of the hemolysis was most likely improper positioning of the device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25872 | 1220648-2024-11951 | 19449314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19449314&pc=OZD | 08/18/2021 | 03/21/2025 |  | 06/03/2024 |  |  | 08/18/2021 | 06/03/2024 |  |  | 08/31/2022 |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1500242 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was an air in purge alarm that indicated that an embolism occurred. also, a small apical thrombus was reported that was noted in the left ventricle of the patient per an echocardiogram report. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
25873 | 1220648-2024-11952 | 19450038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19450038&pc=OZD | 05/23/2024 | 07/17/2024 |  | 06/03/2024 |  |  | 05/23/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. a patient with post-operative cardiogenic shock was in need of hemodynamic support. the placement of the impella 5.5 pump failed and so the team attempted to place an impella cp. the impella cp support also failed despite all attempts and measures such as wire exchange. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned and data logs were not available. per the clinical review, the root cause of delivery issue is determined to be patient condition causing the resistance for the pump to cross the anastomosis because of diseased vessels. f6 and f8 should have been blank. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
25874 | 1220648-2024-11953 | 19450588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19450588&pc=OZD | 05/24/2024 | 07/23/2024 |  | 06/03/2024 |  |  | 05/24/2024 | 06/03/2024 |  |  |  |  | 06/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464710 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. hemolysis occurred overnight with lactate dehydrogenase uptrend to 972 and plasma free hemoglobin 35.2 with tea-colored urine. the team had been aware of poor position pointing toward mitral valve but was hesitant to reposition since there were no issues. repeat transthoracic echocardiogram today showed impella at 5.3 centimeters and interacting with the mitral valve. the team was unable to reposition toward the septum but did pull back the device to try to minimize hemolysis. attempt was unsuccessful to reposition toward the septum. the patient is stable. | [5 empty] | The investigation into the patient's hemolysis has been completed. abiomed products and data logs were not returned for evaluation. based on the clinical review, the root cause of hemolysis was most likely due to pump in the improper position. f6 and f8 should be blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
25875 | 1220648-2024-11903 | 19441853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441853&pc=PYX | 12/01/2016 | 07/05/2024 |  | 06/01/2024 |  |  | 05/02/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Vascular Dissection | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A literature study entitled 'utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease" monitored 48 patients over two years who were implanted with a mechanical circulatory device. during the support period, an unspecified number of patients on impella rp support experienced conditions including limb ischemia, hematomas and axillary dissection. | [5 empty] | The investigation for the reported access site bleed, vascular damage, and limb ischemia has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the reported issues could not be determined. h6 type of investigation, corrected from 4115 to 4114. h6 health effect, corrected from 2199 to 4650. | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ impella rp smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ citation: kaki, a., alraies, m. c., blank, n., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c., elder, m., & schreiber, t. (2018). tct-757 axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. journal of the american college of cardiology, 72(13). https://doi.org/10.1016/j.jacc.2018.08.1983.
25876 | 1220648-2024-11904 | 19441999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441999&pc=OZD | 09/05/2018 | 07/11/2024 |  | 06/01/2024 |  |  | 05/02/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | UNKNOWN |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 02may2024, abiomed complaints team received a literature study entitled "'axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock" monitored 17 patients over 20 months who were implanted with a mechanical circulatory device. an unspecified number of those patients received an impella cp implant. during those 20 months, one patient expired from an intracranial hemorrhage and seven patients passed away after receiving renal replacement therapy for acute kidney injury. | [5 empty] | The investigation for the reported stroke and renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore the root cause of the stroke and renal failure could not be determined. b2-selected death d1-corrected brand name d4-corrected model number f6/f8-should have been left blank in the initial report g1 reporting contact fax number missing g4-added pma number h1-corrected type of reportable event to death h6-impact code changed to 1802 and 4641 | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ citation: kaki, a., blank, n., alraies, m. c., jani, a., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c. l., mohamad, t., elder, m., & schreiber, t. (2019). axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. the american journal of cardiology, 123(10), 1715¿1721. https://doi.org/10.1016/j.amjcard.2019.02.033.
25877 | 1220648-2024-11906 | 19442044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442044&pc=PYX | 05/02/2024 | 07/11/2024 |  | 06/01/2024 |  |  | 05/02/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Renal Failure | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A literature study entitled 'axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock" monitored 17 patients over 20 months who were implanted with a mechanical circulatory device. an unspecified number of those patients received an impella rp implant. during those 20 months, 1 patient expired from an intracranial hemorrhage and seven patients passed away after receiving renal replacement therapy for acute kidney injury. | [5 empty] | The investigation for the reported stroke and renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore the root cause of the stroke and renal failure could not be determined. the instructions for use for the impella rp smartassist system with automated impella controller in section: potential adverse events (united states) states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ added- h6 impact code 4641 g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ kaki, a., blank, n., alraies, m. c., jani, a., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c. l. f., mohamad, t., elder, m., & schreiber, t. (2019). axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. the american journal of cardiology, 123(10), 1715¿1721.https://doi.org/10.1016/j.amjcard.2019.02.03.
25878 | 1220648-2024-11905 | 19442071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442071&pc=OZD | 12/01/2016 | 07/05/2024 |  | 06/01/2024 |  |  | 05/02/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Stroke/CVA | Cardiogenic Shock | Multiple Organ Dysfunction Syndrome | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A literature study entitled 'utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease" monitored 48 patients over 2 years who were implanted with a mechanical circulatory device. during the support period an unspecified number of patients on impella 2.5 support expired from conditions including cardiogenic shock, stroke, and multiple organ failure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): acute renal dysfunction, acute renal dysfunction, aortic insufficiency, aortic valve injury, atrial fibrillation, bleeding, cardiogenic shock, cardiac tamponade, cardiopulmonary resuscitation, cerebral vascular accident/ stroke, death, device malfunction, failure to achieve angiographic success, hemolysis, hepatic failure, insertion site infection, limb ischemia, myocardial infarction, need for cardiac, thoracic or abdominal operation, perforation renal failure, repeat revascularization, respiratory dysfunction, sepsis, severe hypotension, thrombocytopenia thrombotic vascular (non-cns) complication, transient ischemic attack, vascular injury, ventricular arrhythmia, fibrillation or tachycardia. citation: alraies, m. c. (2019). utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease. 2019 scientific sessions. https://scai.confex.com/scai/2019/webprogram/paper3220.html. a2, a3, a5 unknown. b2: date of death unknown. d4: model number, catalog number, serial number, lot number, expiration date, and udi number are all unknown. a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | The investigation for the reported stroke, cardiogenic shock, multiorgan failure, and death has been completed. the impella device was not returned for evaluation. additionally, clinical details were insufficient to determine root cause. therefore, the root cause of the issues could not be determined. the instructions for use referenced in the initial report is for the impella 2.5 with smart assist system in section: potential adverse events (united states). d4-corrected model number. h6 clinical code changed from 1762 to 2262. f6 and f8 should have been blank. added attachment. inadvertently omitted from initial mdr.
25879 | 1220648-2024-11867 | 19442108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442108&pc=OZD | 07/26/2021 | 07/11/2024 |  | 06/01/2024 |  |  | 05/02/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | VA ECMO, VV ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On 2024-05-02 complaints forwarded publication entitled: "use of impella devices for acute cardiogenic shock in the perioperative period of cardiac surgery" which detailed 11 impella patients, 7 impella cp, 1 impella rp, 1 impella cp-impella rp bipella, and 2 impella 5.0 pumps. the publication noted "device related complications included varying degrees of hemolysis in 8 patients (73%) and device malfunction in 1 patient (9%). patient 1 (intraoperative) with the device malfunction had both an impella rp and impella cp support as well as va-ecmo"(venous-arterial extra-corporeal membrane oxygenation). patients underwent placement of impella devices preoperatively, intraoperatively or postoperatively with venous-arterial/venous-venous extra-corporeal membrane oxygenation (va/vv emco) utilization for some patients. | [5 empty] | The investigation for the reported hemolysis and the pump stop has been completed. the impella was not returned for evaluation. the root cause of the hemolysis cannot since the product was not returned and insufficient clinical details were provided. the root cause of the pump stop cannot since the product was not returned and insufficient clinical details were provided. d4-updated fields, expiration date unknown h10 corrected citation: sicouri, s., shah, v. n., buckley, m., imperato, n., mcgee, j., casanova, e., gnall, e., & plestis, k. a. (2023). use of impella devices for acute cardiogenic shock in the perioperative period of cardiac surgery. brazilian journal of cardiovascular surgery, 38(1). https://doi.org/10.21470/1678-9741-2021-0398. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. note(s): device-specific information other than brand name was not provided and therefore sections d4/h4 are respectively indicated as unknown. unspecified malfunction is captured to section h6 medical device problem code as pump stop. publication reference: serge sicouri1, md, et al. (2022). use of impella devices for acute cardiogenic shock in the perioperative period of cardiac surgery. braz j cardiovasc surg 2022-ahead of print: 1-8.
25880 | 1220648-2024-11907 | 19442678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442678&pc=OZD | 09/17/2021 | 03/21/2025 |  | 06/01/2024 |  |  | 09/17/2021 | 06/01/2024 |  |  | 06/30/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2021 |  |  |  |  | The user facility reported an 85-year-old female on impella cp for mechanical circulatory support. due to access complications, the pump was explanted in the procedural area. it was reported that the patient experienced loss of pulse in foot but had known severe peripheral vascular disease (pvd). | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25881 | 1220648-2024-11910 | 19442719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442719&pc=OZD | 09/22/2021 | 05/31/2024 |  | 06/01/2024 |  |  | 09/22/2021 | 06/01/2024 |  |  | 06/30/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100986 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2021 |  |  |  |  | The complainant reported the patient was implanted with an impella cp for mechanical circulatory support. while on support, ventricular tachycardia was noted. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25882 | 1220648-2024-11909 | 19442750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442750&pc=OZD | 09/18/2021 | 05/13/2025 |  | 06/01/2024 |  |  | 09/18/2021 | 06/01/2024 |  |  | 07/31/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022109405 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2021 |  |  |  |  | The complainant reports 86 year old male patient with critical coronary artery disease (cad) was presented with acute decompensated heart failure (adhf) which required intubation at the medical facility. an attempt was made to place the impella in the right femoral artery (rfa), but the medical professional had difficulty accessing the site. the left femoral artery (lfa) was prepped, and the device was implanted. it was reported that the left groin had a cardiac device infection (cdi) and that the patient¿s platelet counts continued trending downward. a heparin, induced antibody test, was performed and the decision was made to change the purge solution to sodium bicarbonate (bicarb) twenty-five (25) in one (1) liter of dextrose with five percent water (d5w). the patient subsequently transfused one (1) unit of packed red blood cells (prbc¿s). weaning was successful and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25883 | 1220648-2024-11908 | 19442936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442936&pc=OZD | 09/18/2021 | 03/21/2025 |  | 06/01/2024 |  |  | 09/18/2021 | 06/01/2024 |  |  | 07/31/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108725 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2021 |  |  |  |  | The user facility reported a 71-year-old female implanted impella cp for mechanical circulatory support. it was reported the patient experienced bleeding from insertion to removal of the impella cp device. the patient required blood transfusions. cardiac function was restored. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25884 | 1220648-2024-11912 | 19442947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442947&pc=OZD | 10/22/2021 | 05/31/2024 |  | 06/01/2024 |  |  | 10/22/2021 | 06/01/2024 |  |  | 07/31/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121574 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/22/2021 |  |  |  |  | The complainant reported that a 65 year old female patient was implanted with an impella cp for mechanical circulatory support. while on support, bleeding was noted at the impella access site. the oozing was resolved with re-suturing and higher elevation. patient to receive left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25885 | 1220648-2024-11913 | 19442991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19442991&pc=OZD | 11/11/2021 | 03/21/2025 |  | 06/01/2024 |  |  | 11/11/2021 | 06/01/2024 |  |  | 03/31/2022 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1480854 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. impella device was implanted without any issues. it was reported that the device was successfully weaned and explanted. post explant, the groin access site continued to ooze. an angiogram identified a dissection of the right common femoral artery (rcfa) from the perclose. a percutaneous transluminal coronary angioplasty (ptca) was performed on the rcfa, and manual pressure was held to resolve the reported issue. patient tolerated the procedure and was in stable condition. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (vascular dissection) are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
25886 | 1220648-2024-11911 | 19443042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19443042&pc=OZD | 09/24/2021 | 05/14/2025 |  | 06/01/2024 |  |  | 09/24/2021 | 06/01/2024 |  |  | 06/30/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103261 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2021 |  |  |  |  | The complainant reported a 73 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support, the impella cp moved out of position and the device was unable to be repositioned. the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction : section b1 and h1 were inadvertently reported incorrectly in the initial report which has been updated to correct information in this report.
25887 | 1220648-2024-11914 | 19443133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19443133&pc=OZD | 12/20/2021 | 06/01/2024 |  | 06/01/2024 |  |  | 12/20/2021 | 06/01/2024 |  |  | 07/31/2023 |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116654 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a 75 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support the patient had pink urine and there was bleeding at the impella insertion site. heparin was cut in half. the impella device was explanted and an intra-aortic balloon pump (iabp) placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: iimpella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25888 | 1220648-2024-11915 | 19443449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19443449&pc=OZD | 05/21/2024 | 09/30/2025 |  | 06/01/2024 |  |  | 05/21/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442123 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 51-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were concerns of hemolysis and multiple attempts were made to reposition the pump to remedy the situation. the decision was made to exchange the impella 5.5 after multiple failed re-positioning attempts. no patient harm was reported. | The complainant also reported that there was a placement signal not reliable. alarm occurred overnight. |  |  |  |  | The investigation for the positioning issue has been completed. the pump was not returned for investigation. the total case runtime was 140 hours. data logs show the placement signal was not reliable for 1.5 hours, with a reported downward trend in snr before reaching zero. the same trend was noted in contrast with minimum of 7. there were no changes noted in lamp, light, and gain values. placement signal values returned to normal after 1.5 hours. the placement signal was flat at the end of the case with reported positioning alarms. according to the clinical report, the placement signal reliability alarm was seen overnight. additionally, the doctor had difficulty repositioning the pump. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provide. f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was not returned for investigation. data logs show the placement signal was not reliable for 1.5 hours, with a reported downward trend in snr before reaching zero. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11915. b5 updated to include optical signal issue description. g1 reporting contact email updated.
25889 | 1220648-2024-11916 | 19443451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19443451&pc=OZD | 05/14/2024 | 08/02/2024 |  | 06/01/2024 |  |  | 05/14/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442072 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the automated impella controller (aic) displayed a suction alarm. the flow displayed was 0.41 liters. the decision was made to explant the pump. the patient was noted to have been in the end of organ failure prior to the incident. | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. motor current and placement signal curves flattened during support and a suction alarm occurred. the flow dropped to 0.4l, then the physician decided to explant the pump after verifying the position through echo. the pump was returned and evaluated. there were no cannula kinks, impeller damage, or foreign material found. a small amount of potential biomaterial was found on impeller. the low pump flow was not reproduced. no data logs were returned for this investigation. the root cause was not determined during product evaluation and no data logs were returned. the root cause was therefore not determined as the low pump flow could not be reproduced and insufficient clinical data was provided.
25890 | 1220648-2024-11917 | 19443458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19443458&pc=OZD | 05/14/2024 | 11/14/2024 |  | 06/01/2024 |  |  | 05/14/2024 | 06/01/2024 |  |  |  |  | 06/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, the automated impella controller (aic) displayed motor current and placement signal curves being flat, without alarm. a suction alarm occurred at the same time as the others. the pump's position was verified to be correct, via echocardiogram. there were no position alarms. the physician withdrew the pump to provoke position alarm but there was no alarm except for the suction alarm previously seen. the pump was ultimately explanted. the patient was noted to have expired, but the patient was at the end of organ failure prior to the incident. | [5 empty] | The investigation for the console (aic) buzzer/alarm malfunction has been completed. only the data logs were returned for evaluation. analysis of the data revealed that pump 467229 is consistent with complaint description ¿motor current and placement signal curves become flat without alarm¿. the pump associated with this failure was reported under mfr report #1220648-2024-11916. however, according to clinical details, there was no issue found with the console during the case. the device functioned/operated to specification. | D.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25891 | 1220648-2024-11842 | 19433321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433321&pc=OZD | 04/22/2024 | 06/24/2024 |  | 05/31/2024 |  |  | 05/01/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2025452390 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock undergoing percutaneous coronary intervention underwent placement of the impella cp device for mechanical circulatory support and thrombus was noted. during percutaneous coronary intervention, the patient suffered cardiopulmonary arrest and was given percutaneous cardiopulmonary support while being resuscitated. the impella cp was then inserted through the femoral artery. when an antegrade sheath was inserted to prevent lower limb ischemia, there was no retrograde blood, and it was confirmed that there was a thrombus in the superficial femoral artery. therefore, the impella cp pump was removed. it is unclear whether the blood clot existed before the impella insertion or whether it formed during the impella cp pump insertion. the patient was able to maintain blood flow in the lower extremities through collateral delivery from the deep femoral artery, and the percutaneous coronary intervention was completed with only percutaneous cardiopulmonary support only without any special treatment. the patient was transported to the unit, noted to be, in stable and "nimble" condition. it was noted subsequently that care was withdrawn, and the patient expired. medical safety review outcome noted there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to poduring repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the product was not returned for investigation. thrombus was also noted. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the thrombus was not determined.
25892 | 1220648-2024-11843 | 19433678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433678&pc=OZD | 09/30/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/30/2021 | 05/31/2024 |  |  | 07/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022120203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient's heart rate began dropping and the patient became hypotensive. the physician believed that the impella was causing mitral regurgitation (mr), so the impella was explanted prior to stent placement. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25893 | 1220648-2024-11844 | 19433683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433683&pc=OZD | 09/30/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/30/2021 | 05/31/2024 |  |  | 07/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2021 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, a large hematoma was found at the access site, and atrial fibrillation was noted. manual pressure was applied, followed by a pressure dressing. heparin was removed from the purge cassette due to the bleeding. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25894 | 1220648-2024-11845 | 19433704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433704&pc=OZD | 09/18/2021 | 10/20/2025 |  | 05/31/2024 |  |  | 09/18/2021 | 05/31/2024 |  |  | 01/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022048776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2021 |  |  |  |  | The user facility reported a 62-year-old male on impella cp support due to elevated troponin and elevated lactate. while on support it was noted the pump was in the chordae/mitral valve. the physician repositioned the device. there was successful removal of the impella cp device and hemostasis was achieved. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿
25895 | 1220648-2024-11846 | 19433739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433739&pc=OZD | 09/20/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/20/2021 | 05/31/2024 |  |  | 07/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114084 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Ventricular Fibrillation | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2021 |  |  |  |  | The complainant reported the 66-year-old male patient was implanted with an impella cp for hemodynamic support. during support the patient experienced ventricular fibrillation and sepsis. no further information was provided. the patient continued on support until the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
25896 | 1220648-2024-11847 | 19433783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433783&pc=OZD | 09/30/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/30/2021 | 05/31/2024 |  |  | 02/28/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022052411 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved by starting bicarbonates, removing anticoagulants, and adding fluid volume. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25897 | 1220648-2024-11849 | 19433915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433915&pc=OZD | 05/01/2024 | 09/10/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complainant reported the automated impella controller (aic) had battery charging issues. the aic was exchanged as a result. there was no reported patient harm. the patient and device information is not available at the time. | [5 empty] | Upon review, it was determined that the complaint reported had been created in error. there was no event nor allegation against the impella device or any of its component. | The aic was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the physician, patient and device information is not available at this time, additional report will be filed when this information is available.
25898 | 1220648-2024-11850 | 19433933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19433933&pc=OZD | 05/10/2024 | 03/14/2025 |  | 05/31/2024 |  |  | 05/13/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1433589 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The facility reported a 83-year-old female of unknown race and ethnicity noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the automated impella controller (aic) display froze during primary percutaneous coronary intervention (ppci). therefore, the aic was exchange and the procedure was successfully completed. there was no reported patient harm. | [5 empty] | The aic was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) ¿ display issue has been completed. there was no issue found during the case. based on clinical details, device analysis, and data logs, the device functioned/operated to specification. d2 common device name revised on manufacturer device report 1220648-2024-11850 in accordance with updated procedures. f6/f8-should have been left blank on manufacturer device report 1220648-2024-11850 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-11850 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on manufacturer device report 1220648-2024-11850. g1 manufacturing site name and address was inadvertently reported incorrectly on the manufacturer device report 1220648-2024-11850.
25899 | 1220648-2024-11851 | 19434093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434093&pc=OZD | 05/11/2024 | 07/29/2024 |  | 05/31/2024 |  |  | 05/13/2024 | 05/31/2024 |  |  | 02/28/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317497 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The complainant reported that impella cp was implanted on 55-year-old male patient with acute myocardial infarction complicated by cardiogenic shock. after successful implant, the patient experienced arterial bleeding around repositioning sheath in icu. the patient had received heparin during implantation and iv thrombolysis some hours before. the bleeding was unable to be controlled, leading to pump explant and blood cells infusion was needed. the pump will not be returned as it was discarded. patient is stable post explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 component code 925 f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25900 | 1220648-2024-11852 | 19434145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434145&pc=OZD | 05/22/2024 | 07/30/2024 |  | 05/31/2024 |  |  | 05/22/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465320 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was being prepped when the automated impella controller (aic) would only show motor current waveform and impella flow box. the placement signal and purge system boxes were blank. the impella was advanced and turned on briefly with noted good flows, but it was quickly removed to assess any issues for patient safety. the pump was tested in a bowl by the doctor with normal function. the aic and impella cp were replaced and procedure was successful. there was no patient harm related to this event. | [5 empty] | The investigation for the optical signal issue has been completed. no product was returned. data logs were reviewed and the optical 5s parameters signal-to-noise ratio, contrast was decreasing while the gain and lamp increased. this failure indicated the air gap characteristic which occurred immediately after plug connect. this issue resolved after the pump white plug was disconnected and reinserted. the cause of the optical signal failure was most likely an air gap from between the automated impella controller and the patient cable plug. b.7 revised as information was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-11852. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11852 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11852 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25901 | 1220648-2024-11853 | 19434238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434238&pc=OZD | 05/22/2024 | 10/30/2025 |  | 05/31/2024 |  |  | 05/22/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251887 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Component or Accessory Incompatibility | Computer Operating System Problem | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The user facility reported a 71-year-old white undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella cp was being prepped when the automated impella controller (aic) would only show motor current waveform and impella flow box. the placement signal and purge system boxes were blank. there was the yellow optical sensor error message at start of case, but the staff noticed that the white cable was not pushed in all the way, once it was disconnected and re-inserted, the error message went away, which led the team to believe issue was resolved. the impella was advanced and turned on briefly with noted good flows, but it was quickly removed to assess any issues for patient safety. the pump was tested in a bowl by the doctor with normal function. the aic was replaced. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that during case start, a ¿placement signal not reliable¿ condition was detected, and case start wizard prompted operator to disconnect pump due to ¿optical sensor system error¿ (rtlog data confirms very low snr and invalid placement signal). following case start wizard prompts, the pump was disconnected and reconnected within 1 second. this time eeprom hadn't been read, however g0 calibration succeeded with a new valid g0 value. after this moment, although optical snr was high, and raw optical signal was valid, the placement signal on console screen (and calculated placement signal in lt and rt logs) had persistent value of 0. console ic6527 was tested in abiomed singapore field service office, but the issue was unable to be reproduced. the console operated within specification during testing. investigation into a possible aic sw defect (gid-dft-3734730) concluded that the software operated as per design and correctly responded to all triggering criteria of pump disconnection and connection. there was no issue with data coming from optical bench. the console functioned/operated to specification regarding receiving opm data. there was no issue found with console cabling during the case. the device functioned/operated to specification. the cause of the optical signal issue was not determined since issue could not be reproduced a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9: updated. h3: updated. h6: updated. b7 other relevant history, including preexisting medical conditions updated. d2: common device name updated. d4: model number updated. d10: concomitant medical products and therapy dates updated. g1: reporting contact email updated. | The aic was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25902 | 1220648-2024-11854 | 19434274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434274&pc=OZD | 05/24/2024 | 07/15/2024 |  | 05/31/2024 |  |  | 05/24/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025506777 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 68-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that there were low pressures throughout the percutaneous coronary intervention (pci) and noted an access site vessel damage/perforation. the team explanted and placed a 14 french gore dyseal and stents to close and achieve hemostasis. blood products were given, and cardiopulmonary resuscitation (cpr) initiated along with the use of a lucas chest compression device. the patient expired despite efforts. | [5 empty] | The investigation into the vascular damage-peripheral vessel has been completed. a second failure mode of low flow has been added to the h6 coding. the data logs revealed low pump flows upon activation which increased gradually until p-level was changed. the logs also show suction alarms. the root cause of the low pump flow was most likely patient condition related. the root cause of the vascular damage - peripheral vessel was not determined due to lack of sufficient clinical information. b7 additional information added. h6 medical device problem code added. h6 component code added f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿.
25903 | 1220648-2024-11855 | 19434386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434386&pc=OZD | 10/07/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 10/07/2021 | 05/31/2024 |  |  | 03/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced ectopy and echocardiography showed the impella cp device at 5.7 centimeters (cm). the nurse practitioner pulled the device back 2 cm. it was also reported that impella was pushed in during surgery because of the patients very short aorta. it was advanced and repositioned. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25904 | 1220648-2024-11856 | 19434409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434409&pc=OZD | 08/26/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 08/26/2021 | 05/31/2024 |  |  | 06/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097566 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was oozing at the groin site. mattress sutures were placed and the bleeding was controlled. there were also signs of hemolysis in the urine. an echocardiogram was done with good positioning in the left ventricle. urine was concentrated, dark. there were plans to remove the impella. it is unknown if the hemolysis was removed or if the removal of the impella was influenced by the hemolysis. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25905 | 1220648-2024-11857 | 19434434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434434&pc=OZD | 10/09/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 10/09/2021 | 05/31/2024 |  |  | 03/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022062986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was noted that an arterial ultrasound was done since pulses were lost in the right leg. plan devised to decrease sedation as tolerated and gradually wean. the impella cp device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25906 | 1220648-2024-11858 | 19434530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434530&pc=OZD | 09/02/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/02/2021 | 05/31/2024 |  |  | 06/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107068 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2021 |  |  |  |  | The user facility reported a patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support clots to left anterior descending, golf ball sized thrombus in left ventricle, and hematuria were noted; therefore, the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smartassist catheter insertion: section: axillary insertion of the impella cp with smartassist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25907 | 1220648-2024-11859 | 19434551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434551&pc=OZD | 08/30/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 08/30/2021 | 05/31/2024 |  |  | 05/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080553 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient suffered with intermittent episodes of ventricular tachycardia. however, the medical staff decided to continue impella support. it was stated that the patient¿s urine was pink/red in color and that the abiomed representative advised the medical staff to be aware of active signs and symptoms of hemolysis. epinephrine was discontinued as the patient¿s condition was reported as stable. urinary output was clearing, and the medical staff was comfortable managing the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
25908 | 1220648-2024-11860 | 19434585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434585&pc=OZD | 09/03/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/03/2021 | 05/31/2024 |  |  | 03/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059317 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had an episode atrial fibrillation. the patient was placed on levophed support. additionally, there was a clot that was pulled out of endotracheal tube (et) tube. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25909 | 1220648-2024-11861 | 19434691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434691&pc=OZD | 04/24/2024 | 07/17/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449540 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Death | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient implanted with an impella 5.5 device preoperatively for mechanical circulatory support. the patient experienced renal failure and stroke during support and expired. upon removal of the impella 5.5 pump, a clear white blood clot was found inside the cannula. a definitive cause for the cerebral infarction and its relationship to the impella 5.5 pump could not be ascertained. it was later noted the patient subsequently passed as a result of the cerebral infarction and ongoing poor health. it was unknown if the date of the cerebral infarction and date of death was the same. a relationship between the impella 5.5 pump and patient's outcome was undetermined. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. d4-corrected model number.
25910 | 1220648-2024-11862 | 19434765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434765&pc=OZD | 05/01/2024 | 06/20/2024 |  | 05/31/2024 |  |  | 05/01/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464464 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. the staff redressed the site and then there was hematoma noted. the patient had reportedly, "incalculable hemoglobin" stats and the staff performed a massive blood transfusion. the doctor stated there was bleeding from an unknown source, however, it is unknown if the access site bleeding helped alter the patient's hemoglobin numbers leading to the transfusion. additionally, the patient had limb ischemia from the repo sheath being occlusive. ultrasound showed no arterial flow. it was recommended to remove the impella however, the patient was not stable enough to remove the pump. the repositioning sheath was repositioned and drips were decreased to improve flow. furthermore, the patient had two events of ventricular tachycardia. the pump's position was checked on echocardiogram and no effusion was noted. the relationship between the patient's ventricular tachycardia and the impella is unknown. additionally, the patient developed hemolysis while on support. the intervention for the hemolysis is unknown. the patient's family requested do not resuscitate status and the patient expired after impella support withdrawn. per abiomed medical safety officer: there is not enough information to associate use of device with outcome. | [5 empty] | The investigation for the access site bleed, hemolysis, limb ischemia, and ventricular tachycardia has been completed. no product was returned for a visual inspection. data log analysis confirmed the presence of suction alarms, motor current spikes, and trending placement signal. the most likely cause of the major access site bleeding was patient condition related. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. the root cause of the limb ischemia and the ventricular tachycardia could not be determined without additional clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
25911 | 1220648-2024-11863 | 19434805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434805&pc=OZD | 05/24/2024 | 07/17/2024 |  | 05/31/2024 |  |  | 05/24/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139059 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported their automated impella controller (aic) had a battery failure red alarm upon powering on the console. per the instructions for use, the complainant can use a backup console. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) battery failure alarm has been completed. data logs were reviewed which showed immediately after bootup, console displayed ¿battery failure (transport connection blown/missing) ¿ alarm,¿. then, the console shut down improperly after the user's shutdown request. after the shutdown sequence was initiated, logs were ended in a second, which was an abnormal symptom. the console was booted up six more times, and there was no ¿battery failure¿ alarm. the console was returned for evaluation. functional testing was unable to reproduce the reported failure mode. with the ac power, aic charged the battery to 100% and the batttest looks good. there was no issue found with the console hardware related to the reported failure mode. the reported failure mode was reproduced by turning on the aic with the battery switch toggled to off state. the root cause for the battery failure alarm was most likely a use issue (aic might be booted up with the battery switch disengaged). added- d9 device return date f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
25912 | 1220648-2024-11864 | 19434814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19434814&pc=OZD | 09/03/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/03/2021 | 05/31/2024 |  |  | 02/28/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043853 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site oozing, loss of pulsatility, and suspected gastrointestinal bleeding or ischemic bowel. a swan and a temp pacer were inserted. it was also noted that the patient's intravenous lines were compromised. after adjustments were made, then patient stabilized. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25913 | 1220648-2024-11868 | 19435121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435121&pc=OZD | 05/03/2024 | 06/28/2024 |  | 05/31/2024 |  |  | 05/03/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474126 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.5 support returned to the operation room for a washout after having a 5.5 placed via hemi sternotomy. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.4 revised model number from the initial submission in accordance with updated procedures. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
25914 | 1220648-2024-11869 | 19435135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435135&pc=OZD | 05/16/2024 | 07/17/2024 |  | 05/31/2024 |  |  | 05/16/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025446623 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Positioning Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 48-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a bedside echocardiogram was performed and was determined the pump placement was too deep. shortly after, the impella position went into the ventricle and the patient started to decompensate quickly. the cardiothoracic surgeon repositioned impella at bedside. the impella slowly drifted back to a shallow position where the inlet cage was at the level of the aortic valve, and the patient was not being properly supported. for the next 30 minutes the patient began to decompensate, and nine epinephrine bloused were administered. due to the patient having no corneal reflex, not responding to vasoactive drugs, and overall poor prognosis the patient was made do not resuscitate (dnr). the decision was made to withdraw care and the patient expired. | [5 empty] | The investigation for the positioning issue has been completed. clinical details state that everything was fixed in place but the pump still migrated to be shallow. this root cause of the positioning issue cannot be determined without the complaint product and it seems they followed best practices. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25915 | 1220648-2024-11870 | 19435149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435149&pc=OZD | 05/23/2024 | 07/22/2024 |  | 05/31/2024 |  |  | 05/23/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481881 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the insertion of the impella 5.5 was aborted as the impella was unable to pass through the patient's vasculature. the physician noted peripheral vascular disease was present. the team was planning to replace the device with an impella cp. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be advanced due to vasculature. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
25916 | 1220648-2024-11871 | 19435152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435152&pc=OZD | 05/25/2024 | 07/29/2024 |  | 05/31/2024 |  |  | 05/25/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507131 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a 76-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient experienced access site bleeding. medical staff surgically intervened by performing a cut down and administered two units of blood products to the patient. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the impella device was returned for evaluation and upon evaluation, there were no issues found with the device. additionally, clinical information was insufficient to determine the root cause. therefore, the root cause of the access site bleed could not be determined. added- f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address.
25917 | 1220648-2024-11872 | 19435158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435158&pc=OZD | 05/27/2024 | 07/22/2024 |  | 05/31/2024 |  |  | 05/27/2024 | 05/31/2024 |  |  | 08/31/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394943 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. when the pump was advanced, it was too deep in the ventricle. multiple attempts were made to pull the pump back. the impella continued to have suction and it looked like it was bending in the ventricle. the impella was pulled back into the aorta and replaced. there was no patient harm related to this event. | [5 empty] | Data review: data logs confirmed pump was in ventricle (about 1.5 hours into the case) that triggered pump position in ventricle and suction alarms. pump then was stopped manually and disconnected from the console. logs also showed low flow occurred during positioning issue with suction alarms at the end. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. the cause of the low pump flow issue was most likely due to improper positioning per clinical review and log analysis. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25918 | 1220648-2024-11873 | 19435170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435170&pc=PYX | 05/04/2024 | 07/05/2024 |  | 05/31/2024 |  |  | 05/04/2024 | 05/31/2024 |  |  | 08/31/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395776 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Death | Life Threatening | Required Intervention |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella rp support developed hemolysis. it was further reported that the impella rp had decreased flows since removing original extra-corporeal membrane oxygenation right ventricular assist device. throughout the night flows decreased while motor current increased. the team did not want to replace the pump. attempted to try tissue plasminogen activator protocol with no success. eventually the pump flow completely stopped and was removed. the patient's outcome was noted as critical. the team used high vasopressin/inotropic support to support the right heart while the family made a decision on plan of care. care was withdrawn and the patient subsequently expired. additional information provided indicated there was uncertainty if a clot was seen when the pump was eventually removed. there were no kinks seen. when the original right ventricular assist device (rvad) was removed, the impella rp motor current went up, and the physician did not want the impella rp flows decreased during removal of rvad. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hemolysis and low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started and ran for about 31 hours, p level was increased from p-5 to p-9 but flow did not reach to expected level. pump continued running on p-9 with low flow for about 20 hours, then the flow dropped to zero with multiple motor current (mc) spikes that triggered auto suction response. this event remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. the root cause of hemolysis was most likely biomaterial ingestion. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g1 reporting contact fax number missing in accordance with updated procedures, sections h10 have been revised from the initial submission.
25919 | 1220648-2024-11874 | 19435179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435179&pc=OZD | 05/28/2024 | 07/11/2024 |  | 05/31/2024 |  |  | 05/28/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ischemia Stroke | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/28/2024 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and while on support the patient suffered an ischemic cva (cerebral vascular accident) of the thalamus with associated edema. support continued. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. d4-corrected model number f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address h6 clinical code changed to 4418. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25920 | 1220648-2024-11875 | 19435197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435197&pc=OZD | 03/21/2024 | 07/23/2024 |  | 05/31/2024 |  |  | 05/28/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442127 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/28/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us based spouse of a patient supported by the direct placed 5.5 pump filed a medwatch regarding thrombosis in the internal jugular and cephalic veins. the pump support had been a successful one per abiomed field team for the patient for a total of 22 days til the patient was transplanted. the patient survived the support and final outcome was not known. no information regarding when the thrombosis occurred or how it as treated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | Attachments: medwatch report was inadvertently not attached to the initial manufacturer device report 1220648-2024-11875. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ b.2 selection was revised since manufacturer device report 1220648-2024-11875 based off of new health effect - impact code added. d.4 revised model number from the initial submission of manufacturer device report 1220648-2024-11875 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11956 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11875 was submitted. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report 1220648-2024-11875. new code has been added to type of investigation codes. h.6 added code 925 in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11875 was submitted. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report 1220648-2024-11875 in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
25921 | 1220648-2024-11876 | 19435213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435213&pc=OZD | 09/03/2023 | 06/05/2024 |  | 05/31/2024 |  |  | 05/03/2024 | 05/31/2024 |  |  | 06/30/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356955 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient admitted with st-segment elevation myocardial infarction was implanted with an impella 5.5 pump for mechanical circulatory support. a small purge leak was noticed at yellow luer. it was noted, the purge cassette was changed multiple times without an issue. there was no change in purge flow, purge pressure and guidance given was not to change out cassette. there was no report of harm to the patient. | [5 empty] | D4-updated model number. h2-changed device evaluation from no to yes. | Data download from the automated impella controller was expected but not returned for review. however, the impella 5.5 pump was returned, and visual inspection found no issues on the pump. the yellow luer was intact without damage, and there was no leak was found from yellow luer under bench test. the root cause of the purge sidearm leak was most likely use related as the failure mode could not be reproduced. product history was reviewed and revealed the pump passed all post sterile inspection checks. instructions for use for the related event are as follows: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
25922 | 1220648-2024-11877 | 19435290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435290&pc=OZD | 05/01/2024 | 06/20/2024 |  | 05/31/2024 |  |  | 05/01/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485981 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella cp support. the patient's labs continued to hemolyze, but repositioning was declined by the physician. discussions were made on potential escalation to impella 5.5 support; however, the cardiovascular surgeon was hesitant due to the patient's instability. ultimately, the decision was made to explant the pump and the patient was placed on hospice care. patient passed. | [5 empty] | The investigation for the reported hemolysis has been completed. the product was not returned. data logs were returned and reviewed finding no abnormalities were noticed. the root cause of the hemolysis was not determined since product was not returned and insufficient clinical information provided. the instructions for use referenced in the initial report is for the impella cp with smartassist system. d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25923 | 1220648-2024-11879 | 19435309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435309&pc=OZD | 09/01/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/01/2021 | 05/31/2024 |  |  | 04/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, after the peel away sheath was removed and repo sheath was advanced substantial bleeding was noted. the decision was made by the physician to remove the impella while in the lab. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25924 | 1220648-2024-11880 | 19435367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435367&pc=OZD | 09/05/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/05/2021 | 05/31/2024 |  |  | 04/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022073835 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for percutaneous coronary intervention (pci). the patient decompensated immediately after atherectomy and experienced thrombosis and gastrointestinal (gi) bleed. as a result, systemic heparin was withheld. additionally, the physician attempted perclose suture, which was unsuccessful, therefore, manual pressure was held. the impella was successfully weaned and explanted. the patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25925 | 1220648-2024-11881 | 19435431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435431&pc=OZD | 09/16/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/16/2021 | 05/31/2024 |  |  | 04/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079230 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/16/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that there was a soft hematoma from suspected venous ooze at right groin. there were no reports of medical intervention being required to resolve the hematoma. the patient was successfully weaned and the device was explanted due to discovery of left ventricle thrombus during echo. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
25926 | 1220648-2024-11882 | 19435437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435437&pc=OZD | 09/01/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/01/2021 | 05/31/2024 |  |  | 05/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081543 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2021 |  |  |  |  | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support, but implantation was aborted related to the patient¿s anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ this report is being filed as part of a retrospective review of historical records.
25927 | 1220648-2024-11883 | 19435559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435559&pc=OZD | 09/17/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/17/2021 | 05/31/2024 |  |  | 05/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089927 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/17/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, staff were unable to get pulses in the patients lower left extremity (non-impella leg). warm blankets were applied. the patient improved hemodynamically and continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. this report is being filed as part of a retrospective review of historical records.
25928 | 1220648-2024-11884 | 19435583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435583&pc=OZD | 09/02/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/02/2021 | 05/31/2024 |  |  | 03/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063588 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pigtail caught in mitral apparatus therefore, the impella was repositioned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
25929 | 1220648-2024-11885 | 19435751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435751&pc=OZD | 09/02/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/02/2021 | 05/31/2024 |  |  | 05/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085364 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed oozing at the impella cp access site. there was oozing from venous and arterial sites; therefore, heparin was started in purge at 25 iu/ml. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25930 | 1220648-2024-11886 | 19435861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435861&pc=OZD | 09/17/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 09/17/2021 | 05/31/2024 |  |  | 03/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/17/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced some oozing the access site. manual pressure held fifteen minutes with resolution of oozing and hemostasis. it was also decided to explant the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25931 | 1220648-2024-11887 | 19435942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435942&pc=OZD | 03/04/2024 | 01/24/2025 |  | 05/31/2024 |  |  | 03/04/2024 | 05/31/2024 |  |  | 11/30/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431079 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced heparin-induced thrombocytopenia which was resolved with plasmapheresis. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. f6/f8-should have been left blank. on manufacturer device report 1220648-2024-11887. g1 reporting contact email revised on manufacturer device report 1220648-2024-11887 in accordance with updated procedures. h6 component code revised on manufacturer device report 1220648-2024-11887 in accordance with updated procedures.
25932 | 1220648-2024-11888 | 19436008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436008&pc=OZD | 05/03/2024 | 06/13/2024 |  | 05/31/2024 |  |  | 05/03/2024 | 05/31/2024 |  |  | 08/31/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385070 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 escalation from impella cp and experienced access site bleeding. during impella 5.5 pump insertion, an artery was lacerated during the axillary cutdown. this led to some bleeding issues that had to be resolved prior to continuing the cutdown to the axillary artery. after the graft was sewed onto the right axillary artery, there was leaking noted around the 23 french introducer. there was some significant blood loss before the graft clamps were reseated over the introducer. however, the bleeding was resolved. the patient received three units of packed red blood cell and one unit of platelets. the patient was converted from venous-venous (vv)/venous-arterial (va) extra-corporeal membrane oxygenation (ecmo) to vv ecmo and on full impella 5.5 pump support when returning to the intensive care unit. following it was reported that overnight the patient's kidney function deteriorated, and the decision was made to withdraw care. per medical safety office review, there is not enough information to associated the use of the device with the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, the graft clamps were reseated over the introducer and the bleeding was resolved as per the clinical description. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states) states: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)".
25933 | 1220648-2024-11889 | 19436095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436095&pc=OZD | 09/05/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/05/2021 | 05/31/2024 |  |  | 07/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105282 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was discovered that the touhy borst lock was not locked and the blue-butterfly was not secured resulting in oozing at the insertion site. impella was successfully weaned and explanted at bedside with manual pressure and femostop to provide hemostasis. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25934 | 1220648-2024-08186 | 19436137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436137&pc=OZD | 08/15/2022 | 03/24/2025 |  | 05/31/2024 |  |  | 08/15/2022 | 05/31/2024 |  |  | 02/29/2024 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia. an external femoral bypass from a rewire port was placed to resolve the issue. patient tolerated the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (limb ischemia) are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25935 | 1220648-2024-11890 | 19436225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436225&pc=OZD | 09/05/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/05/2021 | 05/31/2024 |  |  | 06/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022099842 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the initial attempt to place the device was unsuccessful as the impella was unable to pass through the iliac artery. a longer sheath was utilized via the right femoral artery (rfa) and the sheath was implanted without issue. the patient was successfully weaned and the device explanted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: section: potential adverse events (united states) ¿section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
25936 | 1220648-2024-11866 | 19436284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436284&pc=OZD | 08/31/2021 | 05/31/2024 |  | 05/31/2024 |  |  | 08/31/2021 | 05/31/2024 |  |  | 06/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092449 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/31/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding from the endotracheal tube. therefore, heparin was withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25937 | 1220648-2024-11865 | 19436290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436290&pc=OZD | 05/01/2024 | 07/02/2024 |  | 05/31/2024 |  |  | 05/01/2024 | 05/31/2024 |  |  | 11/30/2025 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454021 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention (pci). the patient had an anomalous right coronary artery which precluded the placement of interventional products for the right coronary artery surgery. the team explanted the impella cp. the patient still required support for the pci and the team placed an intra-aortic balloon pump. once the pci was performed, the patient was implanted with a new impella cp. the patient was then cannulated for venoarterial extracorporeal membrane oxygenation. ultimately, the patient's family and team decided to withdraw care and the patient expired. per abiomed's medical safety officer, there was no patient injury, only a product malfunction. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. g1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11865 was submitted. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25938 | 1220648-2024-11891 | 19436619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436619&pc=OZD | 05/24/2024 | 07/16/2024 |  | 05/31/2024 |  |  | 05/24/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484103 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, suction alarms occurred at p9 however, within an hour , suction alarms were present from p3 to p9. the team implanted an impella rp and suction alarms were still present at p2 and p3. the staff repositioned the device to no avail. blood volume was given. additionally, the patient's lab were continually hemolyzed. the physician explanted the pump and replaced with a new impella cp. the patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the low pump flow and patient's hemolysis has been completed. the impella pump was returned and investigated. one of the impeller blade was found to be totally broken. clinical information mentions optimal flows were achieved at the start of the case. the patient then received cardiopulmonary resuscitation (cpr) since the patient arrested and after that there were constant suction alarms observed even at lower p-levels. therefore, per the clinical information provided and product investigation, the root cause of the low pump flow issue was fractured impeller blade due to cpr given to the patient. per clinical information, the labs were hemolyzed continuously. product investigation revealed a broken impeller blade. data logs were not available for review. per the clinical information, product evaluation and data log review, the cause of the hemolysis issue was fractured impeller blade due to cpr given to the patient.
25939 | 1220648-2024-11892 | 19436827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19436827&pc=OZD | 05/24/2024 | 07/11/2024 |  | 05/31/2024 |  |  | 05/25/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025467819 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The user facility reported a 67-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that low pulses were found in the patient¿s foot. the physician decided to remove the impella. the patient status is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- h6 component code 925 f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25940 | 1220648-2024-11893 | 19437143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19437143&pc=OZD | 09/15/2021 | 03/24/2025 |  | 05/31/2024 |  |  | 09/15/2021 | 05/31/2024 |  |  | 06/30/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2021 |  |  |  |  | The complainant reported a patient implanted with an impella cp for mechanical circulatory support. while on support, the patient was given blood products due to bleeding at groin site. the patient¿s urine was pink due to b-12, so heparin was reduced. the patient was cannulated for extracorporeal membrane oxygenation (ecmo) with placement of impella cp in left femoral artery (lfa). the impella was explanted and removed as a bridge to therapy. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25941 | 1220648-2024-07815 | 19437600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19437600&pc=OZD | 08/10/2022 | 03/24/2025 |  | 05/31/2024 |  |  | 08/10/2022 | 05/31/2024 |  |  | 10/31/2023 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022156134 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 70-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient was placed on impella cp support after the patient started to require increased pressor support overnight. the patient developed a small hematoma at the impella cp groin insertion site. angle matching dressing was in place. the plan was to add heparin inside the purge at a later time. the patient was later escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25942 | 1220648-2024-07981 | 19437939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19437939&pc=OZD | 10/21/2022 | 03/24/2025 |  | 05/31/2024 |  |  | 10/21/2022 | 05/31/2024 |  |  | 07/31/2024 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had oozing from the groin area. forward-facing sutures were placed to stop the oozing. the patient was successfully weaned from the impella and survived. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
25943 | 1220648-2024-11895 | 19438777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19438777&pc=OZD | 12/01/2016 | 07/03/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A literature study entitled 'utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease" monitored 48 patients over two years who were implanted with a mechanical circulatory device. during the support period, an unspecified number of patients on impella cp support experienced conditions including limb ischemia, hematomas and axillary dissection. | [5 empty] | The investigation for the reported access site bleed, vascular damage, and limb ischemia has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the reported issues could not be determined. f6 and f8 should have been left blank h6 type of investigation corrected from 4115 to 4114. | D.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ kaki, a., alraies, m. c., blank, n., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c., elder, m., & schreiber, t. (2018). tct-757 axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. journal of the american college of cardiology, 72(13). https://doi.org/10.1016/j.jacc.2018.08.1983
25944 | 1220648-2024-07926 | 19439701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19439701&pc=OZD | 08/29/2022 | 05/31/2024 |  | 05/31/2024 |  |  | 08/29/2022 | 05/31/2024 |  |  | 05/31/2024 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed a hematoma at the impella access site. the medical team decided to explant the impella device due to the hematoma. | [5 empty] | The investigation into the access site bleeding issue has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25945 | 1220648-2024-11896 | 19439783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19439783&pc=OZD | 09/05/2018 | 06/26/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A literature study entitled 'axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock" monitored 17 patients over 20 months who were implanted with a mechanical circulatory device. during the support period, patients on impella cp support experienced conditions including limb ischemia (1), bleeding (1), and upper extremity deep vein thrombosis (1). two units of blood were given to treat the bleed. | [5 empty] | D.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." kaki, a., alraies, m. c., blank, n., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c., elder, m., & schreiber, t. (2018). tct-757 axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. journal of the american college of cardiology,72(13).https://doi.org/10.1016/j.jacc.2018.08.1983 | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. as all the necessary information was not provided, the root cause of the limb ischemia and the thrombus was not determined.
25946 | 1220648-2024-11897 | 19439987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19439987&pc=PYX | 09/05/2018 | 07/09/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A literature study entitled 'axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock" monitored 17 patients over 20 months who were implanted with a mechanical circulatory device. during the support period, and patients on impella rp support experienced conditions including limb ischemia (1), bleeding (1), and upper extremity deep vein thrombosis (1). two units of blood were given to treat the bleed. | [5 empty] | The investigation into the reported issues has been completed. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. as all the necessary information was not provided, the root cause of the limb ischemia and thrombus was not determined. e.1 name prefix/title, first name and last name were revised since manufacturer device report number 1220648-2024-11897 was submitted. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11897 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report number 1220648-2024-11897 was submitted. | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ impella rp smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ citation: kaki, a., blank, n., alraies, m. c., jani, a., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c. l., mohamad, t., elder, m., & schreiber, t. (2019). axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. the american journal of cardiology, 123(10), 1715¿1721. https://doi.org/10.1016/j.amjcard.2019.02.033.
25947 | 1220648-2024-11898 | 19440308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19440308&pc=OZD | 05/26/2024 | 07/29/2024 |  | 05/31/2024 |  |  | 05/26/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469467 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2024 |  |  |  |  | The complainant reported a patient with plans to undergo cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the pump unexpectedly stopped functioning. multiple attempts to restart the pump were unsuccessful and the decision was made to explant the device. there was no patient harm resulting from the pump stop. the physician opted to diurese the patient for a few days prior to percutaneous coronary intervention. | [5 empty] | The investigation for the pump stop has been completed. the impella cp pump was not returned. data logs were reviewed, and motor current spike with saturated flows observed at the end. no pump stop alarm was found with likely missing logs for last 1.5 hour of pump support. the cause of the pump stop issue was most likely biomaterial per clinical and log review. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25948 | 1220648-2024-11899 | 19440653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19440653&pc=OZD | 05/27/2024 | 06/20/2025 |  | 05/31/2024 |  |  | 05/27/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466311 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | White | Death | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post percutaneous intervention. the patient developed bleeding from their groin site following impella cp implant. angle matching was applied to the site and a mattress suture was placed to achieved hemostasis. once in the intensive care unit, bleeding at the site returned, and two units of blood were given. a new hemostatic suture was placed to achieve hemostasis. the patient remains on impella cp support on the date of this report. | Once in the intensive care unit, bleeding at the site returned, and two units of blood were given. a new hemostatic suture was placed to achieve hemostasis. support with the implanted pump continued following the intervention with bleeding and organ failure as well as an observation made of a cva. additionally, there was air in purge system/purge system open issue, which was followed by pump stop, which pump stop was replaced. care was withdrawn and patient expired. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, pump stop, and purge leak ¿ pump has been completed. the investigation of stroke and vascular damage - peripheral vessel have not been completed, should any new information become available a supplemental report will be submitted. d9 device returned to manufacturer date added. access site bleeding - major: on the day of implant, pt lifted/bent his leg and bleeding occurred. angle matching done and mattress suture placed. hemostasis achieved. the root cause of access site bleeding is determined to be patient movement because the bleeding started upon moving patients leg and hemostasis was achieved by applying sutures. pump stop: after 3 days on support, purge cassette changed that began purge fluid leak at yellow luer. within 2 hours, pump stop occurred. data log review showed 3 attempted purge bag changes within 1 hour, at which time, purge pressure went to 0mmhg and purge flows went from 10ml/min to 30ml/min. 2 hours after purge issues began, mc trends up and saturated flow occurs followed by pump stop. the pump was returned and inspected. inspection revealed biomaterial wrapped around the impellor blades that was unable to be removed despite soaking in heparin and manual removal efforts. the pump was unable to be run as the pump stop reproduced. the cause of the pump stop was biomaterial due to the biomaterial found wrapped around the impellor blades that reproduced the pump stop and data log review showing lack of purge flow leading to mc and purge flows trending up before pump stop occurred. purge leak - pump: after 3 days on support, purge cassette changed that began purge fluid leak at yellow luer. ecmo rn reported that she believed the nurse accidentally cracked the luer on the impella side arm when changing the purge cassette. rep was able to confirm crack in luer. data log review showed 3 attempted purge bag changes within 1 hour, at which time, purge pressure went to 0mmhg and purge flows went from 10ml/min to 30ml/min. the pump was returned and inspected. inspection revealed purge leak at the yellow luer connection. when the investigator attempted to disconnect the luer, the purge sidearm was further damaged. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc. instructions for use: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ pump stop: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ purge leak-pump: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-11899. b2 death added. date of death is unknown. b5 additional information and patient outcome were inadvertently omitted on the initial manufacturer device report number 1220648-2024-11899. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-11899 g1 reporting contact email revised on manufacturer device report 1220648-2024-11899 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-11899. h1 revised based on patient outcome death. h6 health effect - clinical code 1770 and 3160, health effect - impact code 1802, 4643, and 4629 and medical device problem code 1503 and 1250 were inadvertently omitted on the initial manufacturer device report number 1220648-2024-11899. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2024-11899.
25949 | 1220648-2024-11878 | 19441036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441036&pc=PYX | 12/17/2020 | 06/25/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  | 08/31/2022 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1498341 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 complaints forwarded publication entitled: "percutaneous repositioning of impella rp: the snare¿manoeuvre¿prolapse technique¿a case report" in which a bipella patient (5.0 sf case (b)(4)) and rp placed to support a 49yo male. the rp was seen to be out of position "fluoroscopy demonstrated that the impella rp catheter had ¿fallen back¿ into the right ventricle. multiple attempts at manual manipulation were unsuccessful, and we elected to use snare-assistance to maneuver the pump." this method was effective and rp and 5.0 remained on until the patient expired when the family withdrew care. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25950 | 1220648-2024-07380 | 19441123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441123&pc=OZD | 10/13/2022 | 05/31/2024 |  | 05/31/2024 |  |  | 10/13/2022 | 05/31/2024 |  |  | 08/31/2024 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250813 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported a 78-year-old female requiring an impella cp device for mechanical circulatory support. it was reported the patient experienced bleeding due to the closure system, proglide. the access site was able to be successfully closed with a cover stent. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely due to use issue. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
25951 | 1220648-2024-07899 | 19441169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441169&pc=OZD | 08/26/2022 | 03/21/2025 |  | 05/31/2024 |  |  | 08/26/2022 | 05/31/2024 |  |  | 05/31/2024 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023230886 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 58-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella cp device was inserted, the position was found to be suboptimal. the impella cp device was removed and reinserted. while on impella cp support, the patient experienced a clot that was found and removed during mediastinal exploration surgery in the operating room. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
25952 | 1220648-2024-08466 | 19441259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441259&pc=OZD | 08/27/2022 | 03/12/2025 |  | 05/31/2024 |  |  | 08/27/2022 | 05/31/2024 |  |  | 04/30/2024 |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221127 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. access site bleeding was observed. perclose and manual pressure was applied together with angio seal. hemostasis was achieved and the issue resolved. the device was explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
25953 | 1220648-2024-11901 | 19441654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441654&pc=OZD | 05/02/2024 | 07/03/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ventricular Fibrillation | Cardiogenic Shock | Multiple Organ Dysfunction Syndrome | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A literature study entitled 'utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease" monitored 48 patients over 2 years who were implanted with a mechanical circulatory device. during the support period an unspecified number of patients on impella cp support expired from conditions including cardiogenic shock, stroke, and multiple organ failure. | [5 empty] | The investigation for the reported stroke, cardiogenic shock, multiorgan failure, and death has been completed. the impella device was not returned for evaluation. additionally, clinical details were insufficient to determine root cause. therefore, the root cause of the issues could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in section: potential adverse events (united states). d1-corrected brand name d2-correct common device name. h6 clinical codes corrected to 2130 to 2262 and added 3261. h6 corrected component code. h6 type of investigation-corrected from 4115 to 4114. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ kaki, a., alraies, m. c., blank, n., shemesh, a., kajy, m., laktineh, a., hasan, r., gade, c., elder, m., & schreiber, t. (2018). tct-757 axillary artery access for mechanical circulatory support devices in patients with prohibitive peripheral arterial disease presenting with cardiogenic shock. journal of the american college of cardiology, 72(13). https://doi.org/10.1016/j.jacc.2018.08.1983.
25954 | 1220648-2024-11902 | 19441671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19441671&pc=PYX | 09/05/2018 | 07/03/2024 |  | 05/31/2024 |  |  | 05/02/2024 | 05/31/2024 |  |  |  |  | 05/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Cardiogenic Shock | Multiple Organ Dysfunction Syndrome | Insufficient Information | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 04may2024, the abiomed complaints team received a literature study entitled 'utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease" monitored 48 patients over two years who were implanted with a mechanical circulatory device. during the support period an unspecified number of patients on impella rp support expired from conditions including cardiogenic shock, stroke, and multiple organ failure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ citation: alraies, m. c. (2019). utilization of axillary artery as access for mechanical circulatory support in cardiogenic shock patients with severe occlusive peripheral arterial disease. 2019 scientific sessions. https://scai.confex.com/scai/2019/webprogram/paper3220.html a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | The investigation for the reported stroke, cardiogenic shock, multiorgan failure, and death has been completed. the impella device was not returned for evaluation. additionally, clinical details were insufficient to determine root cause. therefore, the root cause of the issues could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in section: potential adverse events (united states). b2-selected death h6 clinical code corrected from 4580 to 3261. f6 and f8 should have been blank.
25955 | 1220648-2024-11795 | 19423247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19423247&pc=OZD | 05/17/2024 | 07/29/2024 |  | 05/30/2024 |  |  | 05/17/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1469546 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported in preparation for an impella 5.5 pump implant, the automated impella controller (aic) was powered on for support and abruptly shut down. the unexpected aic failure prompted the team to replace the console. there was no report or indication of patient involvement. | [5 empty] | The investigation for the console (aic) red alarm has been completed. console logs from the reported day of event were reviewed and are consistent with complaint and show the console was properly shut down, but the there were failed boot-ups, from which aic recovered by self-rebooting. this was followed by ¿unexpected shutdown¿ advisory alarm (#603). the issue did not reproduce upon the next power-cycle. the console was returned for evaluation. during functional testing, power-cycling 120 times was performed, and a single failed boot-up was observed ((b)(4) rate), from which the console successfully recovered but no alarm was presented during that time. failed boot-ups are rare, and if detected, the aic is designed to power-cycle and restore full operation. depending on small variations in timing (whether failure is detected before or after aic sw fully loaded) the alarm #603 may or may not trigger. therefore, there was no problem detected. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
25956 | 1220648-2024-11796 | 19423653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19423653&pc=OZD | 12/20/2021 | 03/21/2025 |  | 05/30/2024 |  |  | 12/20/2021 | 05/30/2024 |  |  | 09/30/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140094 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding complications. this was managed by a gauze bundle with loose pressure. the patient continued to have a slow, steady oozing at the access site, but the physician was not concerned. the patient was supported with nitroglycerin, esmolol, and aggrastat, and eventually a stitch was used to close the access site while the impella was still implanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
25957 | 1220648-2024-11797 | 19423656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19423656&pc=OZD | 05/17/2024 | 07/29/2024 |  | 05/30/2024 |  |  | 05/17/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489984 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted with st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed a hematoma at the access site. it was noted that bleeding was occurring around the repositioning sheath, and the patient's blood pressure began to drop. two units of blood were given to counteract the blood loss and the decision was made to replace the remove / reinsert the pump while utilizing an introducer in place of the repositioning sheath to stop the bleeding. support with the implanted pump continued following the intervention, and patient was transferred to the unit, noted to be, in stable condition. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical information was provided. h.6 code 4641 was reported incorrectly on initial manufacture device report number 1220648-2024-11797. code 925 was added since the initial submission of manufacture device report number 1220648-2024-11797 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25958 | 1220648-2024-11798 | 19423694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19423694&pc=OZD | 05/21/2024 | 07/17/2024 |  | 05/30/2024 |  |  | 05/21/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025460790 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 79-year-old female patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was enrolled in protect iv study and developed hemolysis while on support. the product insert stated the relationship to be definite to the use of impella. the impella was removed post percutaneous coronary intervention and the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were evaluated and there was no motor current spike or positioning issues observed in the data logs. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information was provided. f6/f8 should have been left blank in the initial report.
25959 | 1220648-2024-11799 | 19423735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19423735&pc=OZD | 12/20/2021 | 03/24/2025 |  | 05/30/2024 |  |  | 12/20/2021 | 05/30/2024 |  |  | 09/30/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145620 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, bleeding and hematoma were found at the access site. manual pressure was applied, the device was explanted, and a balloon tamponade was applied to the right common iliac artery. the vascular surgeon then repaired the tissue and closed the femoral artery after removal. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
25960 | 1220648-2024-11800 | 19423738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19423738&pc=OZD | 05/22/2024 | 06/28/2024 |  | 05/30/2024 |  |  | 05/22/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504299 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support as bailout during percutaneous coronary intervention. upon repositioning, the impella cp and pigtail were pulled across the aortic valve. the patient was taken to the catheterization lab for repositioning again. the patient is stable. the patient remains on impella cp support on the date of this report. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on the provided clinical details, the root cause of the positioning issue was most likely due to use as it was accidentally pulled back into the aorta during repositioning. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25961 | 1220648-2024-11603 | 19424027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19424027&pc=PYX | 05/02/2024 | 07/16/2024 |  | 05/30/2024 |  |  | 05/02/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complaint from the literature review reported 57 impella patients over 2 years reviewed. from this population, 1 patient endured a pump displacement. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. complaint device not returned for investigation. device serial number unknown to review the data logs. insufficient clinical data was provided. the root cause of the positioning issue cannot be determined. d4-corrected model number. g4-pma number.
25962 | 1220648-2024-11615 | 19424127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19424127&pc=OZD | 08/31/2019 | 07/15/2024 |  | 05/30/2024 |  |  | 05/02/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported 130 patients supported by 2.5, cp, 5.0 or rp, complications noted during hospital stay, acute kidney injury occurred in 56.7% of the patients; need of renal replacement therapy (rrt) 31.7%; access site-related bleeding 14.3%; life-threatening bleeding 31.5%; acute limb ischemia 14.5%; hemolysis 33.3%. the rate of all-cause mortality at 30 days was 39.7%. | [5 empty] | The investigation for the access site bleed, limb ischemia, hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, limb ischemia, hemolysis could not be determined. d1-corrected brand name. d2-updated common device name. d4-corrected model number. g1-corrected mfr site name and address. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25963 | 1220648-2024-11616 | 19424255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19424255&pc=PYX | 08/31/2019 | 07/30/2024 |  | 05/30/2024 |  |  | 05/02/2024 | 05/30/2024 | 05/30/2024 |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  | IMPELLA RP | IMPELLA RP |  | PYX | PYX |  | IMPELLA RP | IMPELLA RP |  | ABIOMED, INC. | ABIOMED, INC. |  | US | US | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | No | No |  | Returned to Manufacturer | Returned to Manufacturer | Yes | No | Health Professional | Health Professional |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Impairment | [9 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported 130 patients supported by 2.5, cp, 5.0 or rp, complications noted during hospital stay, acute kidney injury occurred in 56.7% of the patients; need of renal replacement therapy (rrt) 31.7%; access site-related bleeding 14.3%; life-threatening bleeding 31.5%; acute limb ischemia 14.5%; hemolysis 33.3%. the rate of all-cause mortality at 30 days was 39.7%. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed, hemolysis and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis, access site bleed, and limb ischemia could not be determined. the instructions for use for the impella rp flex with smartassist instructions for use & clinical reference manual state the following: section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ section: warnings and cautions . ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound . ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 d1-corrected brand name to impella rp, corrected product code to pyx. d2-corrected common device name to temporary non-roller type right heart support blood pump. d4-corrected model number. g1-corrected mfr site name and address. g4-corrected pma number.
25964 | 1220648-2024-11801 | 19424535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19424535&pc=OZD | 05/15/2024 | 07/12/2024 |  | 05/30/2024 |  |  | 05/15/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465359 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 52-year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that intermittent clusters of air bubbles were seen in the patient¿s left ventricle. the medical staff did not know the cause of the bubbles. there were no preparatory issues for the pump and the purge setup. the patient remains on support. |  |  |  |  | The investigation for the reported embolism has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of embolism could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." systemic embolization, serious injury, or death."
25965 | 1220648-2024-11693 | 19424695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19424695&pc=OZD | 07/26/2021 | 07/11/2024 |  | 05/30/2024 |  |  | 05/02/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | IMPELLA RP, VA-ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The abiomed complaints team received a publication entitled: use of impella devices for acute cardiogenic shock in the perioperative period of cardiac surgery" which detailed 11 impella patients. this report refers to the 7 impella cp patients and the 1 cp-rp bipella. the publication noted device related complications included varying degrees of hemolysis in 8 patients and device malfunction in 1 patient. patient 1 (intraoperative) with the device malfunction had both rp and impella cp support as well as veno-arterial extracorporeal membrane oxygenation (va-ecmo). | [5 empty] | The investigation for the reported hemolysis and the pump stop has been completed. the impella was not returned for evaluation. the root cause of the hemolysis cannot since the product was not returned and insufficient clinical details were provided. the root cause of the pump stop cannot since the product was not returned and insufficient clinical details were provided. citation: sicouri, s., shah, v. n., buckley, m., imperato, n., mcgee, j., casanova, e., gnall, e., & plestis, k. a. (2023). use of impella devices for acute cardiogenic shock in the perioperative period of cardiac surgery. brazilian journal of cardiovascular surgery, 38(1). https://doi.org/10.21470/1678-9741-2021-0398 d1-brand name corrected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
25966 | 1220648-2024-11802 | 19424939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19424939&pc=OZD | 12/20/2021 | 03/24/2025 |  | 05/30/2024 |  |  | 12/20/2021 | 05/30/2024 |  |  | 07/31/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107299 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was noted to have a hematoma in the groin; therefore, the impella was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
25967 | 1220648-2024-11803 | 19425168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19425168&pc=OZD | 12/21/2021 | 05/30/2024 |  | 05/30/2024 |  |  | 12/21/2021 | 05/30/2024 |  |  | 09/30/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp for support due to multiple rounds of cardiopulmonary resuscitation and ventricular tachycardia/ventricular fibrillation. while on support bleeding at impella site, suspected gastrointestinal bleed, and impella unable to re-cross valve was noted. impella was inserted after intra-aortic balloon pump was removed and systemic heparin was put on hold due to bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25968 | 1220648-2024-11804 | 19425295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19425295&pc=PYX | 05/17/2024 | 07/15/2024 |  | 05/30/2024 |  |  | 05/17/2024 | 05/30/2024 |  |  | 07/31/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024390083 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. it was reported that impella rp support was ongoing for a patient admitted in with post operative right ventricular dysfunction. the impella rp had been supporting when there was low flows thought to be due to clot ingestion and signs of hemolysis in red urine. the issues were treated by the pump being explanted and replaced. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the hemolysis and low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed suction alarms and pa pressure was rising with upward shift in placement signal trend and low flows after 88 hours. the impella was returned for evaluation. upon visual inspection, there was biomaterial found blocking the outflow cage, in the cannula throughout the impeller, inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion. the root cause of hemolysis was biomaterial ingestion. the instructions for use for the impella rp smart assist system with automated impella controller state the following: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
25969 | 1220648-2024-11805 | 19425468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19425468&pc=OZD | 04/07/2022 | 05/30/2024 |  | 05/30/2024 |  |  | 04/07/2022 | 05/30/2024 |  |  | 01/31/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185294 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there was a slight ooze (bleeding) at the site the device was explanted after a femstop was placed and manual pressure was held. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25970 | 1220648-2024-11806 | 19425592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19425592&pc=OZD | 04/07/2022 | 05/30/2024 |  | 05/30/2024 |  |  | 04/07/2022 | 05/30/2024 |  |  | 10/31/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154769 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. therefore, heparin was on hold. it was also reported that the patient had low pulse. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
25971 | 1220648-2024-11807 | 19425786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19425786&pc=OZD | 05/17/2024 | 07/12/2024 |  | 05/30/2024 |  |  | 05/17/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 50-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced slow flow distal to the impella pump following implant. flow was noted to be sluggish and an external fem-fem bypass was subsequently completed to manage the condition. flow below the knee remained sluggish and the decision was ultimately made to explant the pump. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25972 | 1220648-2024-11808 | 19425955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19425955&pc=OZD | 05/18/2024 | 06/27/2024 |  | 05/30/2024 |  |  | 05/18/2024 | 05/30/2024 |  |  | 06/30/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient implanted with an impella cp exhibited access site bleeding and limb ischemia. heparin was withheld and the patient was transfused with two units of packed red blood cells to obtain hemostasis. the impella cp was removed and the patient was implanted with an impella 5.5 to address the loss of distal pulses and toes turning black. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia ¿ reversible and the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
25973 | 1220648-2024-11809 | 19426168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19426168&pc=PYX | 05/22/2024 | 09/22/2025 |  | 05/30/2024 |  |  | 05/23/2024 | 05/30/2024 |  |  | 12/31/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024320086 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 69-year-old male patient postoperative valve replacement was implanted with an impella rp device for mechanical circulatory support. it was reported that the patient experienced access site bleed that was treated with fresh frozen plasma (ffp) and protamine. the patient also was no longer producing urine and there were concerns or hemolysis, so continuous renal replacement therapy was started. the pump was explanted, and a right ventricular assist device was placed after 27 hours of support. the patient was reported as stable. | The complainant also reported that the placement signal had been lost but did not require intervention. |  |  |  |  | The investigation for the access site bleed and the hemolysis has been completed. the impella was not returned for evaluation. data log analysis confirmed the presence of suction alarms before a steep decrease in snr and contrast and increase in lamp. however, clinical details is sufficient to determine the root cause of the access site bleed. the most likely cause of the major access site bleeding was patient condition related. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. added- h6 component code 925 f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for a visual inspection. data log analysis confirmed the presence of suction alarms before a steep decrease in signal-to-noise ratio (snr) and contrast and increase in lamp. the cause of the optical signal issue was not determined because no product was returned and not enough clinical information was given. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11809. g1 reporting contact email updated.
25974 | 1220648-2024-11810 | 19426221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19426221&pc=PYX | 05/18/2024 | 07/18/2024 |  | 05/30/2024 |  |  | 05/18/2024 | 05/30/2024 |  |  | 10/31/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024414193 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 36-year-old male with cardiomyopathy was on an impella 5.5 for support for approximately one month when it was decided to add impella rp flex pump support. after three days on the impella rp flex pump, high purge pressure alarms started and the decision was made to explant the impella rp flex due to ongoing purge pressure high alarms. extracorporeal membrane oxygenation support was placed when troubleshooting with tissue plasminogen activator failed. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Data logs show that 57 hours into support there is a sudden 3min rise in high pressure purge (hpp) with drop in purge flow. at the time it begins to rise, there is a mc spike, upwards shift in ps, drop in flow, and slight speed deviation (without p-level change); this is indicative of biomaterial ingestion. there are hpp and purge system blocked alarms. pp remains elevated for 8hours before decreasing over 8 hours (likely due to tissue plasminogen activator (tpa) being run). there is a second rise in pp over 3 min with a mc spike and drop in flows. it remains high but with numerous dips and spikes in pp retriggering the alarms. the pump is removed without the hpp resolving. biomaterial was found on the returned product wrapped around the impeller and purge gap. when run in a bucket in the lab there was a moment of low pump flows with higher pp (no alarms), but was quickly resolved as the pump ran and some biomaterial was expelled and caught in the outflow cage. hpp was not reproduced in a bucket in the lab at p-9. there was still a significant amount of biomaterial on the pump after it was run. the root cause of the high purge pressure was likely biomaterial. d9 added date device returned to manufacturer g1 reporting contact fax number missing
25975 | 1220648-2024-11811 | 19426244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19426244&pc=OZD | 05/23/2024 | 07/26/2024 |  | 05/30/2024 |  |  | 05/23/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479272 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | Additional information received that the decision was made to withdraw support and the patient expired. per medical safety officer review there is not enough information to associate use of device with outcome. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during patient support, the sterile sleeve partially tore near the touhy borst valve. the site was cleaned and the nurse used tegaderm to close the area. no additional complications were encountered. there was no patient harm. the patient remains on impella 5.5 support on the date of this report. |  |  |  |  | The investigation for sterile sleeve damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added. b5 updated information the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11811: f6/f8-should have been left blank . g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25976 | 1220648-2024-11812 | 19426338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19426338&pc=OZD | 05/23/2024 | 07/16/2024 |  | 05/30/2024 |  |  | 05/23/2024 | 05/30/2024 |  |  | 11/30/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438298 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | Material Split, Cut or Torn | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The complainant reported a 65-year-old female patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that after cross clamp was removed from aorta and impella flow resumed, abnormal purge flows, waveforms, and elevated motor currents were noted. a transesophageal echocardiogram (tee) found the cp to be 6 cm too deep from position, so medical staff repositioned the cp. during the reposition, the sleeve, covered with tegaderm, was torn. after repositioning, the inlet was 3.15 cm from the aortic annulus. waveforms were still dissociated. the left ventricle waveform was reading significantly higher than the aortic waveform. purge pressure was over 900 mmhg and purge flows were down to just over 2 ml/hr. flows were also 3.4l/min on p4. medical staff left the cp at p4 and stuck to their initial decision to explant the cp the next day. the patient only made 200 ml of urine during coronary artery bypass graft (cabg) and had almost no urine post-cabg. medical staff noted bright red urine even after reevaluating the position again by a tee. about an hour after arrival to the cardiovascular intensive care unit, the waveforms almost completely normalized. purge pressure and flows also normalized. the motor current was still slightly elevated and remained elevated until explant. | Correction: patient was 67. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into hemolysis, product damage (sterile sleeve) and saturated flow issues has been completed. the impella cp pump was returned. visually inspected the pump and sterile sleeve was found to be torn and damages. borescope view of cannula obtained and found biomaterial wrapped around impeller, purge gap at outflow. no other kinks, damages or abnormalities were found. data logs were returned, and motor current spike > 1400 ma observed when pump restarted after surgery. suction alarms and saturated flows were seen. only few instances of position alarms were observed later. the cause of the hemolysis issue was due to biomaterial wrapped around impeller and outflow. the cause of the product damage (sterile sleeve) issue was use related due to excessive force. the cause of the saturated flow issue was due to biomaterial. d9 added device returned and date returned b1 product problem should have been checked b5 corrected patient's age h6 health effect code corrected from 2558 hematuria to 1886 hemolysis e1 facility name was omitted.
25977 | 1220648-2024-11813 | 19426535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19426535&pc=OZD | 05/18/2024 | 07/15/2024 |  | 05/30/2024 |  |  | 05/18/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025463304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 72-year-old morbidly obese male patient noted with high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed uncontrollable access site bleeding. the access site was explored in the operating room and it was discovered that access was made at the bifurcation. the impella cp was explanted and the patient was to be escalated to impella 5.5. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient condition due to the insertion angle and free-moving state of the sheath and patient anatomy surrounding the reposition sheath causing significant internal bleeding. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25978 | 1220648-2024-11815 | 19427041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427041&pc=OZD | 05/24/2024 | 08/01/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  | 10/31/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024414204 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported an 86-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that pump diplayed three purge pressure low alarms, there were no loose connectors or fluid leaking from the cassette so, the staff suspected fluid leaking inside the cassette. additionally, the patient's circulation was stable however, there was significant bleeding from areas other than the impella access site, so it was decided to remove the impella on the same day. the doctor stated the patient had a tendency to bleed, and it is unclear whether the higher-than-normal amount of impella purge injected may have accelerated the bleeding. the impella was weaned and removed. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | A purge cassette was returned for evaluation. no pump was returned for evaluation. upon visual inspection, no abnormalities were noted. purge cassette was connected to lab console and lab pump and purged for approximately 12 minutes and no alarms were noted and purge cassette was able to maintain normal purge pressure and purge flow. no data logs were returned for evaluation. clinical details noted no leaks were observed in purge line in the field. additionally, clinical team noted that patient was bleeding from non-impella sites and patient had tendency to bleed but it was not clear if increased purge flow was making bleeding worse. the root cause of the purge cassette leak cannot be determined as insufficient product and data logs was returned. the root cause of the access site bleeding - major is most likely due to patient condition which was unrelated to the impella. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11815. f6/f8-should have been left blank.
25979 | 1220648-2024-11816 | 19427266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427266&pc=OZD | 05/24/2024 | 07/15/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  | 11/30/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443276 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 48-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that on day two of impella cp support a purge system blocked alarmed occurred. no kinks or blockages were observed. the impella at p9 showed concerns for pump saturation. purge cassette changed with no significant improvement to purge flows and purge pressure. pump stopped with successful restart. the pump stopped again shortly after with an unsuccessful restart. the pump was explanted, and the patient was reported as hemodynamically stable. | [5 empty] | Pump was returned for evaluation. upon visual inspection, no abnormalities were noted. electrical testing was performed and all motor cable lines were found to be within range. blood was present in the purge gap and blood ingress was seen in purge lumen. pump was placed in a bucket of water and purged for approximately 5 minutes before high purge pressure was observed. pump was attempted to be run but an immediate pump stop occurred. while purging the pump, the catheter was cut at the 18cm mark and purge pressure resolved almost immediately. no data logs were returned for evaluation. the root cause of the pump stop was most likely biomaterial in the motor. d9 added date device returned to manufacturer f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings, cautions, and precautions to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: troubleshooting the purge system unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
25980 | 1220648-2024-11817 | 19427287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427287&pc=OZD | 05/20/2024 | 07/18/2024 |  | 05/30/2024 |  |  | 05/20/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469310 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and during support, low pump flow was encountered. the patient had a successful coronary artery bypass graft surgery. the impella cp pump was turned to p-2, and there was a loud "whining" sound. once changed to p-0, the sound went away. when coming off bypass and attempting to increase the pump flow, the pump was unable to go above p-1 without suction alarms. on transesophageal echocardiogram (tee), the pump placement looked shallow at two centimeters from the aortic valve. the physician attempted to reposition and advance the pump three centimeters with little change and still would not go above p-1 with suction alarms persistent when attempted. it was noted the volume was adequate and right ventricle appropriately contracting on tee. the physician stated the patient was hemodynamically stable off pressors, ejection fraction improved, and the decision was made to remove the impella cp pump which was explanted one hour later in the intensive care unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs confirmed the failure mode and clinical narrative. logs showed after about 118 hours, p-level was down to zero for about 1.75 hours & restarted on p-2. no motor current spike was found. after pump restarted, low flow and suction alarms observed that triggered per auto suction response. several suction alarms were seen during most of the case and also placement signal was down trending since start of the pump support. product was not returned for investigation. the root cause of the low pump flow was most likely patient condition related per log and clinical review. added- h6 impact code 4628.
25981 | 1220648-2024-11818 | 19427689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427689&pc=OZD | 05/20/2024 | 07/22/2024 |  | 05/30/2024 |  |  | 05/20/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted with high-risk percutaneous coronary intervention to the left anterior descending, left circumflex and left main arteries was implanted with an impella cp device for mechanical circulatory support. following implant, the impella cp pump could not achieve flows greater than one liter per minute. continuous suction was observed, and the decision was ultimately made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella was not returned for evaluation. data log was reviewed find that a suction alarm occurred. no motor current spikes were observed. no positioning issues occurred. no indication to confirm the reported low pump flow issue. the cause of the low pump flow cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided.
25982 | 1220648-2024-11819 | 19427746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427746&pc=OZD | 12/15/2023 | 07/09/2024 |  | 05/30/2024 |  |  | 05/09/2024 | 05/30/2024 |  |  | 10/31/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412246 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured large right groin hematoma formation with a "probable" relationship to the impella device used: ¿ecmo (extracorporeal membrane oxygenation), cannulation was complicated by large right groin hematoma formation. " "(b)(6) 2023: hematoma: patient with bilateral groin hematoma, with r>l post ecmo both are currently soft and compressible. plan: -regular assessments with complete blood count twice per day" "right external iliac/sfa repair was performed on (b)(6) 2023: overnight: hemoglobin was 6.4, received 1 unit of blood with improvement to 7.4" baseline hb prior to implantation was 9.9 pt received a total of 4 units of rbcs 1 on each of (b)(6)." the patient recovered with no sequelae. | [5 empty] | The investigation into the access site bleeding ¿ major issue been completed. the device was not returned for investigation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from manufacturer report number 1220648-2024-11819. h.4 revised as previously entered incorrectly on manufacturer report number 1220648-2024-11819. h.6 code 1888 was reported incorrectly on manufacturer report number 1220648-2024-11819. a new code has been added to health effect - clinical codes. h.6 code 4114 was reported incorrectly on manufacturer report number 1220648-2024-11819. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from manufacturer report number 1220648-2024-11819 in accordance with updated procedures. additional manufacturer narrative on manufacturer report number 1220648-2024-11819 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25983 | 1220648-2024-11814 | 19427832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427832&pc=OZD | 05/19/2024 | 07/17/2024 |  | 05/30/2024 |  |  | 05/19/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451676 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, alarms of high purge pressure and purge system blocked occurred. due to the noted issue, the decision was made replace the impella cp with another. there was no patient harm. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned for evaluation. data logs from case show a sudden spike in purge pressure to 1000mmhg, with a drop in purge flow to around 0. the purge flow ticks flatline at the same time as the purge pressure spike. motor current and pump flow are not affected. clinical details noted that the purge pressure spiked suddenly and "high purge pressure" alarms were triggered. the root cause of the high purge pressure is most likely due to a kinked purge line. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25984 | 1220648-2024-11820 | 19427862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427862&pc=OZD | 05/18/2024 | 07/15/2024 |  | 05/30/2024 |  |  | 05/20/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455062 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory report. the patient arrived at the intensive care unit with adequate flows. the chest was left open with a plan to return to the operating room for closure when ready. overnight, it was noted the patient had bleeding from the chest. the chest tube and wound vac had significant output. four units of packed red blood cells were required and heparin was withheld. additionally, after turning the patient, since then, there was access site bleeding. the dressing was removed and sutures, still in place, were tightened which stopped the bleeding. the patient was noted to be in stable condition. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed has been completed. the impella device was returned for evaluation. no damages or abnormalities were found on the returned pump. however, clinical details were sufficient to determined the root cause. the root cause of the access site bleeding major was most likely patient movement related as the site was found to be oozing after the patient was turned.
25985 | 1220648-2024-11821 | 19427907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427907&pc=OZD | 05/24/2024 | 06/28/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1560022 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had a battery showing 100% when plugged in and not connected with a patient. however, when connected with an unknown aged patient and still plugged in, the aic battery level continually decreased. the aic is a loaner device which will be returned for service and replaced with another loaner. there was no report of patient harm. | [5 empty] | The investigation into the aic - battery charging/other issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the reported complaint was confirmed. the batttest utility was run, and the results showed that the console was not detecting when ac power was connected, and therefore the console would not charge the batteries from ac power. the reported battery issue was determined to be caused by insufficient power input from the power supply to the pbm. return date was not provided. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25986 | 1220648-2024-11822 | 19427973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19427973&pc=OZD | 05/24/2024 | 09/11/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1162032 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility biomedical department reported that the automated impella controller (aic) had a battery failure during patient use. no patient harm was reported. | [5 empty] | D.9 updated as the console was returned for investigation. the investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. the console was interrogated in the lab. logs from the failure investigation confirmed battery level low alarm being issued. it was observed that the console was at 49 percent battery on boot up. the console was confirmed to not be charging the batteries despite being supplied with alternating current (ac). it was confirmed that ac inline fuses had blown. replacing the fuses did not resolve the issue as they were immediately blown. interrogation of power supply confirmed that it was not delivering output of 24 volt direct current from its output. replacing the power supply resolved the issue. the root cause for the battery not charging was a defective power supply unit. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11822 and should not have been. | The aic device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25987 | 1220648-2024-11823 | 19428018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428018&pc=OZD | 05/20/2024 | 07/19/2024 |  | 05/30/2024 |  |  | 05/20/2024 | 05/30/2024 |  |  | 11/30/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430940 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The user facility reported a patient undergoing coronary artery bypass (cabg) was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient experienced an access site ooze. the physician tightened the sutures around the blue plug and heparin was withheld. additionally, the doctor was unable to get a good seal around the catheter for proper seal during aortic clamp. the doctor took off the aortic clamp and re-adjusted a second time. around the proper seal, it was noted on second attempt as the doctor stated the left ventricle was decompressed and soft. the aortic cannulation needed to be repositioned so the impella was pulled back with intention to re-advance. there was resistance when advancing so the pump was explanted. | [5 empty] | The investigation for the access site bleed and the positioning issue has been completed. per the clinical information provided, the access site was not sutured tightly which led to bleed of the access site. the oozing from the site was resolved by tightening the sutures. therefore, the root cause of the access site bleeding issue was use related due to improper technique. as per clinical the pump was pulled out of lv with intention to re-advance post procedure, but at the end of the case resistance was met while trying to advance and device explanted. therefore, the root cause of the positioning issue was most likely use related. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25988 | 1220648-2024-11824 | 19428316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428316&pc=OZD | 05/24/2024 | 07/22/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490208 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | Device-Device Incompatibility | Unintended Movement |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump was placed but the team made decision to place venoarterial extracorporeal membrane oxygenation and with that, the impella cp came out of position and needed to be replaced there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. data review: data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position aorta. no other issues were found on the logs. product analysis: visual inspection found no issues on the pump. conclusion: the root cause of positioning issue was most likely use related. added d9 device return and date f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25989 | 1220648-2024-11825 | 19428550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428550&pc=OZD | 05/15/2024 | 07/10/2024 |  | 05/30/2024 |  |  | 05/16/2024 | 05/30/2024 |  |  | 07/31/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443831 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced significant bleeding from their groin site during impella cp support. it was noted that multiple factors contributed to the bleed including the 14fr introducer that had been left in place and the active movement by the patient. one unit of blood was transfused to counteract the bleed and the pump was subsequently explanted. the patient remained stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the peel away sheath was left in place after sending to icu. f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
25990 | 1220648-2024-11826 | 19428571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428571&pc=OZD | 05/21/2024 | 07/12/2024 |  | 05/30/2024 |  |  | 05/21/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448624 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The complainant reported a 33-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient had no pulse present in their right lower extremity below the popliteal artery after the implant. medical staff then escalated the patient to the impella 5.5 for ongoing support. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing h.6 codes 4627, 4642, and 4625 were reported incorrectly on manufacturer device report 1220648-2024-11826. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
25991 | 1220648-2024-11828 | 19428610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428610&pc=OZD | 05/24/2024 | 07/18/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443361 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Injury | Yes | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 pump was being placed for mechanical circulatory support of a patient admitted with st-segment elevation myocardial infarction. the catheter kinked with repeated delivery attempts, and a left ventricle perforation was noted requiring covered stent and/or balloon tamponade. the team removed and replaced the impella 5.5 pump. upon removal kink proximal to motor housing noted. a new impella 5.5 was successfully placed supporting the patient. the patient's outcome condition was noted as stable. | [5 empty] | The investigation for the ventricle perforation and delivery issues has been completed. the impella device was not returned for evaluation. perforation was covered stent and/or balloon tamponade. impella was unable to restart and noticed kink proximal to motor housing after removal. they were able to place a new 5.5 later. the root cause of delivery issue is determined to be use issue because catheter perforated lv during implant. hence, the root cause was determined to be use issue. the root cause of ventricle perforation was determined to be guidewire puncture because the lv was perforated during catheter placement into lv. d4-corrected model number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
25992 | 1220648-2024-11827 | 19428635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428635&pc=OZD | 05/22/2024 | 07/18/2024 |  | 05/30/2024 |  |  | 05/22/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458002 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | Implant Pain | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a 33-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient was in shock and wearing a lifevest with a history of post-partum cardiomyopathy. the 5.5 was being delivered via the axillary artery through a 23fr sheath. the medical staff found the sheath to have an inability to flush. the staff then made the decision to replace the sheath with another, and then support started. the patient remained on impella support with plans for a mitraclip and assessment of the patient¿s pain at the insertion site. the 5.5 was weaned then explanted, and the patient survived the event. | [5 empty] | The investigation into the introducer damage-mechanical has been completed. according to the clinical review, while prepping the 23fr peel away sheath for insertion the tech could not flush the sheath. the decision was then made to use a new sheath from another impella 5.5 kit. the following day, the patient began to feel a lot of pain and swelling around the insertion site. six days later the patient was weaned off of support and the pump was removed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the cause of the mechanical introducer damage was not determined because no product was returned and not enough clinical information was provided. the cause of the pain at insertion site was not determined due to insufficient clinical details. note pain at insertion site is not a reportable event. h6 code patient effect will be corrected in h11. f6 and f8 should have been blank h6 patient effect-remove 4561 as this is not a reportable event. replace with 4582. h6 patient impact-remove 4642 and replace with 2199. 4642 was not related to the impella device but rather a procedure the patient was going to have. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
25993 | 1220648-2024-11829 | 19428720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428720&pc=OZD | 05/23/2024 | 07/18/2024 |  | 05/30/2024 |  |  | 05/23/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482192 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The complainant reported a 65-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed hemolysis after the implant. medical staff repositioned the cp and held off on additional intervention until the patient could be reassessed. the patient died later that morning from cardiac arrest. it is unknown if the hemolysis contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. product was not returned for investigation. data logs were returned and at the time of reported hemolysis, there was no motor current spike or suction observed. there were no positioning issues as confirmed with an echo. the patient's urine remained red until explant even after several repositioning. therefore, per the clinical information provided and data log review, the root cause of the hemolysis issue was not determined. the pump was tried to be repositioned but was not successful since echo showed that the pump was sitting deep and pushing towards apex. therefore, per the clinical information provided, the root cause of the positioning issue was most likely use related to improper technique. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ added-h6 med dev prob code 2616 f6/f8 should have been left blank in the initial report.
25994 | 1220648-2024-11832 | 19428781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428781&pc=OZD | 05/24/2024 | 07/25/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474279 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported difficulty in placing an impella cp pump for mechanical circulatory support in a patient undergoing high-risk percutaneous coronary intervention (pci). the impella cp pump had issues crossing the valve and was advanced, removed and the wire became kinked. the .018 wire was replaced with a v18 compatible wire and therapy initiated. 7fr companion sheath was used without issue for single access and pci completed. at the conclusion of the case, the companion sheath and impella cp pump were removed, and the team observed a right common femoral artery dissection which required vascular repair. restoration of appropriate flow was confirmed, and the patient's outcome condition was noted as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage and product damage has been completed. impella pump and guidewire was returned. visual inspection was performed and the guidewire was found to be kinked. no damages seen with the pump. there are no other abnormalities or damages with guidewire or pump. companion sheath was not returned. the cause of the vascular damage issue was not determined due to insufficient clinical information. the cause of the product damage (guidewire) issue was use related per return product investigation and clinical review. added-d9 device return date, h6 impact code 4627, d4-corrected model number.
25995 | 1220648-2024-11834 | 19428877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428877&pc=OZD | 05/25/2024 | 07/18/2024 |  | 05/30/2024 |  |  | 05/25/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed access site bleeding. heparin was stopped to address the oozing at the site. once the team realized they were at an act of 140, the purge pressure high alarms were alarming. staff attempted to start tissue plasminogen activator however, the pump stopped. they attempted to restart the pump a couple times to no avail. the pump was replaced. there was no patient harm related to the pump exchange. | [5 empty] | The investigation for the pump stop, high purge pressure, and the access site bleed has been completed. no product was returned for investigation. data logs show that about 7 hours into support there is an increase in purge pressure (pp) and drop in purge flow triggering high purge pressure (hpp) and purge system blocked alarms. 2 hours into the case the pp begins to slowly increase over about 5 hours before triggering these purge alarms. less than 1 hour later there are pump stop alarm 13. the pump is unable to be restarted and is explanted. clinical states that heparin was turned off prior to this due to bleeding. the root cause of the pump stop was likely due to high purge pressure. the root cause of the high purge pressure was unable to be determined because product was not returned. the root cause of the access site bleeding major issue was not determined since product was not returned and there is a lack of clinical details. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist for use during cardiogenic shock and high risk pci section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25996 | 1220648-2024-11833 | 19428915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428915&pc=OZD | 05/24/2024 | 07/22/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025474723 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility reported a 62-year-old female patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the pump kept coming out of position. the pump position was adjusted, and flows were appropriate at the time of pump re-adjustments. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issues has been completed. complaint device was not returned for investigation. data logs were reviewed which found positioning alarms in aorta occurred at the time of the reported event. also drop in motor current and flow observed that returned back to normal range after repositioning process. the cause of the low pump flow issue was improper positioning due to improper technique and after repositioning the device, flow returned to required range. added- d6b f6/f8 should have been left blank in the initial report.
25997 | 1220648-2024-11835 | 19428936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428936&pc=OZD | 05/25/2024 | 08/02/2024 |  | 05/30/2024 |  |  | 05/25/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469235 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was also supported with an intra-aortic balloon pump. the patient had arrhythmia issues at the start of support. the patient had small runs of ectopy and rate controlled atrial fibrillation. the impella was found deep and it was repositioned. no further ectopy was noted. additionally, suction alarms occurred. the impella was unable to flow despite good position and volume. the impella was explanted. staff noted there was biomaterial on the impella post explant. the team determine the patient was not a candidate for escalated treatment as the patient is jehovah's witness. | [5 empty] | The investigation for the atrial fibrillation and low pump flow has been completed. the impella has been returned for evaluation. upon visual inspection, biomaterial was found on the returned product. data logs confirm suction alarms. data logs also show motor current (mc) spike with drops in flow. clinical states that biomaterial was observed on the pump at explant. the root cause of the low pump flow issue was likely biomaterial. the root cause of the atrial fibrillation could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
25998 | 1220648-2024-11831 | 19428942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428942&pc=OZD | 05/24/2024 | 07/25/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025476442 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The complainant reported a 64-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient¿s heart failed to recover and the patient died of multiple organ failure due to infection and poor peripheral control. no impella operation issues were identified. the relationship between the infection and the impella device remains unknown. | [5 empty] | The investigation for infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11831: f6/f8-should have been left blank. g1 reporting contact fax number missing. g2 report source: foreign missing. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
25999 | 1220648-2024-11830 | 19428946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19428946&pc=OZD | 05/24/2024 | 08/05/2024 |  | 05/30/2024 |  |  | 05/24/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1423376 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/24/2024 |  |  |  |  | The user facility biomedical department reported that the automated impella controller (aic) had a controller error. the console was removed from service. no patient harm was reported. | [5 empty] | D.9 updated as the console was returned for investigation. the investigation was completed. the console was returned for investigation. the logs confirmed the reported complaint. a controller error occurred due to the optical bench ambient pressure being out of range. real time logs revealed that the controller error was triggered when the ambient pressure dropped below approximately 550 mmhg, and that the controller error was automatically cleared when the ambient pressure rose back above approximately 550 mmhg. the console¿s optical system is designed to perform ambient pressure compensation for ambient pressure between approximately 550 mhg and 825 mmhg, and the observed behavior for the reported complaint matches this expected behavior. the triggered controller error was expected based on the ambient pressure at which the console was being used. there was no evidence of a device malfunction. the consoled was received and investigated. the optical system was tested and confirmed to be working correctly. there was no evidence of a device malfunction. there was no issue found during the case. the device functioned/operated to specification. a.3 revised as information was inadvertently omitted from the submission of the initial manufacturer device report number 1220648-2024-11830. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11830 and should not have been. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26000 | 1220648-2024-11838 | 19429013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19429013&pc=OZD | 05/27/2024 | 07/19/2024 |  | 05/30/2024 |  |  | 05/27/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025483474 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention (pci). the impella cp had a sudden decrease in motor current and drop in flows. the patient also experienced a gastrointestinal bleed. systemic anticoagulation was held. the team lowered heparin in purge for extended time. bedside echo showed correct placement with possible thrombus near inlet. the pci was successfully performed. the impella cp was explanted. | [5 empty] | The investigation for low pump flow and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow was most likely due to biomaterial based on a spike in the data logs. the root cause of the vascular damage - peripheral vessel cannot be determined as there was limited information on pump insertion. instructions for use: section: warnings and cautions: cautions physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26001 | 1220648-2024-11837 | 19429027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19429027&pc=PYX | 05/27/2024 | 07/29/2024 |  | 05/30/2024 |  |  | 05/27/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025437999 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The user facility reported a 57-year-old male patient of unknown race post operative cardiothoracic surgery was implanted with an impella rp device for mechanical circulatory support. it was reported that the flows of an implanted impella rp pump dropped after adjustments were made to an existing swan during patient support. there was no heparin available at the time of the noted issue and it was believed that a clot had entered the pump based on the displayed waveforms. due to the noted issue, the decision was made to replace the pump with a protek duo for ongoing support. there was no patient harm resulting from the noted issue. | [5 empty] | The investigation for the low pump flow has been completed. the impella was returned for evaluation and was visually inspected. no cannula kink observed. no broken blades found. biomaterial observed around impeller. the cause of the low pump flow was biomaterial ingestion due to improper anticoagulation. added- d9 device return date, h6 impact code 4629 f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26002 | 1220648-2024-11836 | 19429219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19429219&pc=OZD | 05/14/2024 | 08/01/2024 |  | 05/30/2024 |  |  | 05/27/2024 | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025452428 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2024 |  |  |  |  | The user facility reported a 74-year-old female patient was implanted with an impella cp device for mechanical circulatory support. it was reported that during a scheduled percutaneous coronary intervention, a rupture occurred that was noted to be unrelated to the impella device. however, due to the resulting oozing, heparin was removed from the purge line. following a period of support without the administration of heparin, a clot was found from the abdomen to the impella access site. a fogarty catheter was subsequently utilized to successfully remove the clot. the patient was reported as stable. | [5 empty] | The investigation thrombus and vascular damage - great vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11836: f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
26003 | 1220648-2024-10980 | 19430880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19430880&pc=OZD | 02/20/2022 | 03/24/2025 |  | 05/30/2024 |  |  | 02/20/2022 | 05/30/2024 |  |  | 12/31/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2022 |  |  |  |  | The user facility reported a 73-year-old male on impella cp support for acute myocardial infarction / cardiogenic shock. while on support, the patient experienced access site bleeding that was resolved with manual pressure, femostop, and suture. the patient was reported to be in stable condition. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26004 | 1220648-2024-11296 | 19430885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19430885&pc=OZD | 02/25/2022 | 05/30/2024 |  | 05/30/2024 |  |  | 02/25/2022 | 05/30/2024 |  |  | 05/31/2023 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124059 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient lost pulses to the left lower extremity on the impella side. additionally, hematuria was noted and labs were hemolyzed. the decision was made to wean and explant the implant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26005 | 1220648-2024-07863 | 19431072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19431072&pc=OZD | 08/15/2022 | 05/30/2024 |  | 05/30/2024 |  |  | 08/15/2022 | 05/30/2024 |  |  | 02/29/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 79-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support and experienced significant oozing, so a mattress suture was placed after repositioning sheath. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
26006 | 1220648-2024-07917 | 19431082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19431082&pc=OZD | 08/15/2022 | 05/30/2024 |  | 05/30/2024 |  |  | 08/15/2022 | 05/30/2024 |  |  | 04/30/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220589 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 75-year-old male patient on an impella cp for mechanical circulatory. the patient developed a hematoma at the impella insertion site. manual pressure was held at bedside. it was recommended that due to the growing hematoma that bicarb be used and heparin be reduced in the purge. systematic heparin was withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
26007 | 1220648-2024-08081 | 19431088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19431088&pc=OZD | 09/07/2022 | 03/17/2025 |  | 05/30/2024 |  |  | 09/07/2022 | 05/30/2024 |  |  | 03/31/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204119 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 57-year-old male patient was implanted with an impella cp for mechanical circulatory support as a pre percutaneous coronary intervention. the patient developed a hematoma from the bleeding that occurred at the impella cp insertion site. the bleeding stopped and the impella was removed due to the hematoma at the site. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26008 | 1220648-2024-06483 | 19431089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19431089&pc=OZD | 03/08/2023 | 05/30/2024 |  | 05/30/2024 |  |  | 05/11/2023 | 05/30/2024 |  |  | 11/30/2024 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303059 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 94 KG |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | [9 empty] | The user facility reported 65-year-old male patient of unknown race and ethnicity was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a rip in the device's anticontamination sleeve during support. there was no infection, harm, or additional intervention reported. | [5 empty] | The investigation into the reported product damage has been completed. the returned product was visually inspected, and the sterile sleeve was confirmed to have been pulled out from the repo hub. there were no other damages or abnormalities observed. there was no evidence of tearing was observed but stretching of the sleeve was observed. the outer diameter of the hub was measured to be 10.411mm and the inner diameter of the locking cylinder was measured to be 10.211mm indicating an interference fit. the evaluation revealed that the cause of the damaged sterile sleeve was use related due to excessive force. stretching of the sterile sleeve was observed and the contamination sleeve was confirmed to be an interference fit. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
26009 | 1220648-2024-11839 | 19432039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19432039&pc=OZD | 04/02/2024 | 10/01/2025 |  | 05/30/2024 |  |  | 04/02/2024 | 05/30/2024 |  |  | 11/30/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024434118 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | Perforation of Vessels | [11 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/02/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia (vt) that was resolved with shock. the patient also experienced peripheral vascular damage that required addition of packed red blood cells. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant also reported that the pump lost placement/optical signal after two days of support. the pump remained on for support despite the loss. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: arrhythmia impella 5.5 system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the vascular damage and ventricular tachycardia has been completed. clinical details state that the patient experienced ventricular tachycardia (vt) while on support that was resolved with shock. the patient also experienced peripheral vascular damage that required addition of packed red blood cells. no product was returned. data logs show continuous impella position unknown alarms. the root cause of the vascular damage - peripheral vessel and vt was not determined due to limited clinical information. added h6 component code 925 corrections to the initial mfr report: b2-selected death and added date of death. f6/f8 should have been left blank. g1 mdr reporting contact email. h1-type of reportable event changed to death. h6 impact code 1802 added. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. data logs show continuous impella position unknown alarms. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated b5 updated to include placement/optical signal issue description. h6 updated to include placement/optical signal issue coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11839. d10 concomitant medical products and therapy dates updated.
26010 | 1220648-2024-11840 | 19432123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19432123&pc=OZD | 11/27/2023 | 09/25/2025 |  | 05/30/2024 |  |  | 11/27/2023 | 05/30/2024 |  |  | 09/30/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410716 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/27/2023 |  |  |  |  | The complainant reported the patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient had access site oozing that required adjustment of heparin administration and the site was redressed. due to the patient's condition and health, the family withdrew care and the impella device was removed. | The patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The complainant also reported that the waveforms on the associated automated impella controller were showing the left ventricle was inverted and above the aorta placement signal at time of the event, and several hours later the waveforms returned back to normal. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the access site bleed has been completed. pump was not returned for investigation. data logs were investigated. the 5s parameters were observed to be in specification. no other issues were observed related to placement signal and optical parameters. there was no issue found during the case. the device functioned/operated to specification. clinical mentions that there was slight oozing observed and therefore heparin was changed to sodium bicarb solution. but the reason is unknown. therefore, the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. corrections to the initial mfr report: b5-mso review added as it was inadvertently left off the initial report. f6/f8 should have been left blank. g1 mdr reporting contact email. h6 component code 925. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were investigated. the 5s parameters were observed to be in specification. no other issues were observed related to placement signal and optical parameters. there was no issue found during the case. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem. b2 updated to death. b5 updated to include the optical signal issue description. h6 updated to include death and optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11840. d10 concomitant medical products and therapy dates updated.
26011 | 1220648-2024-11841 | 19432231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19432231&pc=OZD | 10/19/2023 | 07/17/2024 |  | 05/30/2024 |  |  | 10/19/2023 | 05/30/2024 |  |  | 06/30/2025 |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370802 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The complainant reported a patient was implanted with an impella cp due to acute myocardial infarction/ cardiogenic shock. while on support, the patient experienced ventricular fibrillation that was treated with lidocaine. additionally, the patient experienced access site bleeding that required addition of blood products. the patient was successfully weaned from the impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Section h6 has been updated, as the investigation has been completed. according to clinical, the patient was placed onto impella cp due to acute myocardial infarction / cardiogenic shock. while on support, the patient experienced ventricular fibrillation that was treated with lidocaine. additionally, the patient experienced access site bleeding that required addition of blood products. the impella device was not return, and therefore an evaluation of the device was not possible. product history review was completed, and the pump passed all sterile inspection checks. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the ventricular fibrillation, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the ventricular fibrillation was not determined. pertaining to the access site bleeding- the root cause of this adverse event cannot be determined as product was not returned and sufficient information is not provided in the clinical description. correction was made to the following section(s): f6: date captured to the initial report should be disregarded; section not applicable.
26012 | MW5155642 | 19435983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19435983&pc=OZD | 05/24/2024 | 05/28/2024 |  | 05/30/2024 |  |  |  | 05/30/2024 |  |  |  |  | 05/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED INC. | [6 empty] | 0042-0040-US |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 48 YR |  | Male | 121 KG | Non Hispanic | White | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Impella alarmed "purge pressure high" verified no kinks or obstructions in line. contacted 1-800 number and spoke with impella representative. new alarm: "purge system blocked" appeared while on phone with impella rep. instructed by rept to change cassette and purge fluid in impella. fluids and cassette changed and alarms did not clear. tpa ordered from pharmacy and pharmacist came to bedside with tpa. instructed to change purge fluid with tpa. purge fluid changed to tpa. new alarm: "impella stopped" appeared. instructed to attempt restart by rep. "restart failed" appeared. attempted this an additional three times and restart failed. instructed by impella rep to withdraw the device out of the ventricle. bedside ultrasound confirmed complete retraction out of the left ventricle and into the descending aorta. pt remained stable throughout the procedure.
26013 | 1220648-2024-11741 | 19412510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19412510&pc=OZD | 04/08/2022 | 05/29/2024 |  | 05/29/2024 |  |  | 04/08/2022 | 05/29/2024 |  |  | 12/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177688 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The user facility reported an 80-year-old male patient on impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock (ami/csg). while on support, the patient had limb ischemia that required the patient to have an external bi-fem/sfa (superficial femoral artery) bypass performed. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26014 | 1220648-2024-11742 | 19412927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19412927&pc=OZD | 05/16/2024 | 07/10/2024 |  | 05/29/2024 |  |  | 05/16/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 64-year-old male of unknown race in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced oozing and developed a hematoma at their access site during impella cp support. a figure eight suture was initially placed to manage the oozing following implant, but the condition returned. the site was then managed with dressing changes and angle match dressings for the remainder of support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, a figure eight suture and angle matching was done to prevent the bleeding/oozing as per the clinical description.
26015 | 1220648-2024-11743 | 19413018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19413018&pc=OZD | 05/16/2024 | 07/15/2024 |  | 05/29/2024 |  |  | 05/16/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481273 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 66-year-old male with a right coronary artery dissection was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient had limb ischemia. the patient's foot was cold without pulses. a fem-fem bypass was used to successfully restore flow and pulse. additionally, the patient had hemolysis. the pigtail was thought to be caught on the chordae as significant movement was seen on echocardiogram and pump was jumping back when retraction was attempted. the staff attempted to reposition the pump but were unsuccessful. care was withdrawn after 30.30 hours of support and the patient expired. the cause of death is unknown. the relationship between the cause of death and the impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and positioning issue has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause of the positioning issue. the cause of the positioning issue most likely use related due to improper technique. the root cause of the limb ischemia could not be determine without sufficient clinical details. f6/f8 should have been left blank in the initial report.
26016 | 1220648-2024-11744 | 19413056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19413056&pc=OZD | 04/30/2024 | 06/24/2024 |  | 05/29/2024 |  |  | 04/30/2024 | 05/29/2024 |  |  | 05/31/2025 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351891 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient had an impella 5.5 implant and low pump flow was observed after the implant. after the implant, the staff were unable to maintain flows. the staff wanted to escalate to extracorporeal membrane oxygenation, but the doctor declined. the patient expired on the table after 0.33 hours of support. there is not enough information to associate the use of device with the patient's outcome. | [5 empty] | The impella device was not received from the customer and thereforean evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. product was not returned for investigation. data logs were not returned as well. therefore, the root cause of the low pump flow issue was not determined since product was not returned and insufficient clinical information was provided.
26017 | 1220648-2024-11745 | 19413791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19413791&pc=OZD | 09/18/2021 | 03/25/2025 |  | 05/29/2024 |  |  | 09/18/2021 | 05/29/2024 |  |  | 05/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed oozing at the impella access site and it was noted that the nurse was addressing it; however, it was not reported how the bleeding was being addressed. additionally, the patient experienced episodes of ventricular tachycardia, which resulted in swan removal. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26018 | 1220648-2024-11746 | 19413847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19413847&pc=OZD | 09/23/2021 | 03/25/2025 |  | 05/29/2024 |  |  | 09/23/2021 | 05/29/2024 |  |  | 05/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083754 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2021 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced access site bleeding and hematoma. the patient was administered blood products. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26019 | 1220648-2024-11747 | 19413917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19413917&pc=OZD | 09/24/2021 | 03/25/2025 |  | 05/29/2024 |  |  | 09/24/2021 | 05/29/2024 |  |  | 01/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022045947 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2021 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced hemolysis, ventricular tachycardia (vt), and thrombocytopenia, all of which resolved with no further details specified. the patient remained on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26020 | 1220648-2024-11748 | 19413984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19413984&pc=OZD | 09/25/2021 | 03/25/2025 |  | 05/29/2024 |  |  | 09/25/2021 | 05/29/2024 |  |  | 06/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022120528 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported the patient had bleeding at the groin site. the patient was administered arterenol and the bleeding discontinued. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. g1 manufacturing site fax number.
26021 | 1220648-2024-11749 | 19414030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414030&pc=OZD | 09/30/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 09/30/2021 | 05/29/2024 |  |  | 06/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104204 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/30/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed oozing at the impella access site. manual pressure was held to maintain hemostasis. the patient was also administered two units of packed red blood cells. the patient survived the procedure. no further details were provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26022 | 1220648-2024-11750 | 19414143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414143&pc=OZD | 09/30/2021 | 04/16/2025 |  | 05/29/2024 |  |  | 09/30/2021 | 05/29/2024 |  |  | 07/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112078 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/30/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was attempted to be implanted with an impella device for mechanical circulatory support. it was reported that the device motor would not pass from the iliac into the aorta. additional efforts to place the impella were futile. it was reported that the medical staff believed the catheter kinked affecting placement. implantation was aborted and another impella cp device was implanted. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26023 | 1220648-2024-11751 | 19414301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414301&pc=OZD | 09/30/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 09/30/2021 | 05/29/2024 |  |  | 02/28/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057742 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/30/2021 |  |  |  |  | The user facility reported a 58-year-old male patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during placement, the physician was repositioning, the pump, there was significant resistance. the patient also experienced a rapidly growing hematoma. the groin site was closed with perclose (two were placed prior to the case, but one failed) and manual pressure was held for 25 minutes. the decision was made to remove the impella. the device was explanted, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26024 | 1220648-2024-11752 | 19414332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414332&pc=OZD | 10/19/2023 | 06/19/2024 |  | 05/29/2024 |  |  | 04/30/2024 | 05/29/2024 |  |  | 08/31/2025 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393852 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured right groin pseudoaneurysm with a "definite" relationship to the impella device used. the patient recovered with no sequelae. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation. due to limited clinical details and lack of product returned, the root cause of the vascular damage - peripheral vessels cannot be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
26025 | 1220648-2024-11753 | 19414557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414557&pc=OZD | 09/30/2021 | 03/25/2025 |  | 05/29/2024 |  |  | 09/30/2021 | 05/29/2024 |  |  | 04/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076484 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced several instances of ventricular tachycardia requiring multiple defibrillations. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26026 | 1220648-2024-11613 | 19414588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414588&pc=PYX | 03/27/2020 | 07/03/2024 |  | 05/29/2024 |  |  | 05/02/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Ischemia Stroke | Thrombosis/Thrombus | [6 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 the complaints team forwarded publications. the entitled: "deploying mechanical circulatory support via the axillary artery in cardiogenic shock and high-risk percutaneous coronary intervention" details retrospective review from january 2016 to september 2018 - 48patients in which 4 had upper limb ischemia, 14 had acute kidney injury requiring renal replacement therapy, 1 had care escalated to ecmo (extracorporeal membrane oxygenation) , 1 had ischemic stroke, 6 had bleeding barc >2, hemolysis in 4, and finally 1 had axillary artery thrombosis. the devices were 2.5, cp, rp. unspecified allegations for the one patient care escalated to ecmo. | [5 empty] | The impella devices were not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported access site bleed, hemolysis, limb ischemia, renal failure, stroke, thrombus, and ventricular arrhythmia has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the reported issues could not be determined. citation: kajy, m., laktineh, a., blank, n., tayal, r., tanveer, s., mohamad, t., elder, m., schreiber, t., & kaki, a. (2020). deploying mechanical circulatory support via the axillary artery in cardiogenic shock and high-risk percutaneous coronary intervention. the american journal of cardiology, 128, 127¿133. https://doi.org/10.1016/j.amjcard.2020.04.039 f6 and f8 should have been blank g1 reporting contact email address corrected.
26027 | 1220648-2024-11754 | 19414663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414663&pc=OZD | 10/01/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 10/01/2021 | 05/29/2024 |  |  | 06/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092817 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding including hematoma was noted and a blood transfusion was required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26028 | 1220648-2024-11755 | 19414795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414795&pc=OZD | 10/01/2021 | 05/20/2025 |  | 05/29/2024 |  |  | 10/01/2021 | 05/29/2024 |  |  | 07/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022117136 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a slightly pink tinged urinary output was noted. the patient was weaned successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26029 | 1220648-2024-11756 | 19414956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19414956&pc=OZD | 10/01/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 10/01/2021 | 05/29/2024 |  |  | 06/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094441 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2021 |  |  |  |  | The complainant reported a patient was presented in cardiac arrest. the patient was intubated in the emergency room and a stent was placed in the right coronary artery. the patient became pulseless in the catheterization laboratory, cardiopulmonary resuscitation was performed and the impella was placed, and the patient regained pulse but experienced an excessive arrhythmia during ventricular fibrillation. the medical staff made the decision to remove the impella, disconnect the patient¿s ventilator and call time of death. without warning the patient regained pulse independently, the support devices were reconnected, and the patient was transported to the intensive care unit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26030 | 1220648-2024-11757 | 19415175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415175&pc=OZD | 10/03/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 10/03/2021 | 05/29/2024 |  |  | 07/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118792 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/03/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the patient was having bleeding issues and platelets have dropped. they were suspecting heparin induced thrombocytopenia and they removed heparin in purge. in addition, the patient was also septic. ejection fraction was 40% but pressures labile so they didn't want to remove, however the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26031 | 1220648-2024-11758 | 19415271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415271&pc=OZD | 02/21/2021 | 06/25/2024 |  | 05/29/2024 |  |  | 05/02/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHARMACIST | Literature | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A publication received indicated a retrospective observational cohort study of patients that required an impella device support for at least 12 hours and received 25-u/ml concentration of heparin purge solution experienced bleeding and thrombosis complications. conclusion and relevance noted of the 2% of patients experienced a thrombotic event involving the pump motor. of the 63% bleeding events, 35% patients had clinically relevant bleeding barc 3a and 3b (bleeding academic research consortium) and no patients met criteria for barc types 3c, 4, or 5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. note(s): patient-specific information sections a2/a3 represents average (median) patient age and sex of the population majority respectively as cited in the attached article. device-specific information other than brand name was not provided and therefore sections d4/h4 are respectively indicated as unknown. journal article reference: christine jiang, pharmd, bcps et. al (2021). safety and efficacy of a percutaneously inserted ventricular support device purge solution heparin 25 u/ml. annals of pharmacotherapy vol. 55(2) 174¿180. | The investigation into the access site bleeding ¿ major and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
26032 | 1220648-2024-11759 | 19415388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415388&pc=OZD | 05/10/2024 | 07/05/2024 |  | 05/29/2024 |  |  | 05/10/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463962 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient they had a hematoma at the access site. the patient was taken to the operating room to have a washout. case wrap up: patient has successful left ventricle assist device implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock. added h6 component code 925 f6/f8-should have been left blank.
26033 | 1220648-2024-11760 | 19415433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415433&pc=OZD | 10/10/2021 | 03/25/2025 |  | 05/29/2024 |  |  | 10/10/2021 | 05/29/2024 |  |  | 07/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding issues. femostop was placed over the impella site to control bleeding, and the patient received three (3) units of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26034 | 1220648-2024-11761 | 19415436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415436&pc=OZD | 05/10/2024 | 06/24/2024 |  | 05/29/2024 |  |  | 05/10/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494772 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of a patient admitted in/transferred to the tertiary center for care escalation. the pump was placed femorally and limb ischemia developed. the limb was treated by perfusion placed via antegrade access. the flow was restored, and pump remained on for 4 days wean/explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
26035 | 1220648-2024-11762 | 19415477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415477&pc=OZD | 05/15/2024 | 07/17/2024 |  | 05/29/2024 |  |  | 05/15/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025459866 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 50-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that during peel-away sheath removal, cp was inadvertently pulled back into the aorta, and it could not be re-advanced into proper position. next, the team wired the re-access port and removed the pump over the wire, followed by insertion of a new 14 french peel-away sheath. in the process of inserting a new sheath, the sterile integrity of the impella was compromised and had to be discarded. the team then opened a new cp to be inserted in the existing access. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue and product damage has been completed. pump was not returned for investigation. the cause of the positioning issue was use related since pump was accidentally pulled back while removing peel-away sheath. as per the clinical details, sterile integrity of the impella was compromised in the process of inserting a new sheath. the cause of the sterile sleeve damage issue was most likely use related. f6/f8 should have been left blank in the initial report.
26036 | 1220648-2024-11763 | 19415589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415589&pc=OZD | 12/30/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 12/30/2021 | 05/29/2024 |  |  | 10/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150411 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the surgeon stated that he did not find any active bleeding, but a large amount of clot was removed from patient's chest. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
26037 | 1220648-2024-11764 | 19415711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415711&pc=OZD | 05/21/2024 | 07/24/2024 |  | 05/29/2024 |  |  | 05/21/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433345 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported delivery issues were encountered during an attempt to implant an impella cp pump in a patient in need for a high-risk percutaneous coronary intervention (hrpci) for the left main artery. the team had the 14fr sheath placed, and despite all troubleshooting measures, the anatomy precluded the placement of the impella cp pump. the delivery was unsuccessful, and when the percutaneous coronary intervention completed without support, the 14fr was removed. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be patient condition because the pump was unable to advance past the illiac artery despite multiple attempts. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as information was inadvertently left off the previously submitted manufacturer device report number 1220648-2024-11764 and was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-11764 in accordance with updated procedures.
26038 | 1220648-2024-11765 | 19415924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19415924&pc=OZD | 12/30/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 12/30/2021 | 05/29/2024 |  |  | 10/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s pulse pressure became due to small arteries, most likely limb ischemia. the physician decided to only treat the right coronary artery ostial and mid and leave cir (small artery and territory) to bring pt back and do cir through the arm, | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26039 | 1220648-2024-11766 | 19416133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416133&pc=OZD | 12/30/2021 | 05/20/2025 |  | 05/29/2024 |  |  | 12/30/2021 | 05/29/2024 |  |  | 08/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122605 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and there was an increasing purge pressure over night. the option was discussed via phone if purge pressure is still increasing to call back when cardiologist gets involved. the cardiologist decided to remove the pump due to intermittent high-pressure alarm situation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26040 | 1220648-2024-11767 | 19416521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416521&pc=OZD | 12/30/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 12/30/2021 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s pulses were palpable. the urine was red in color and prompting a discussion about concerns of hemolysis. upon arrival, the patient suffered a stroke and was transferred to interventional radiology where an observed thrombus was removed. heparin was held due to the questionable cerebral vascular accident and the medical staff was considering tissue plasminogen activator. it was stated that the patient would be transitioned to left ventricular assist device destination therapy. the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26041 | 1220648-2024-11768 | 19416607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416607&pc=OZD | 05/21/2024 | 07/03/2024 |  | 05/29/2024 |  |  | 05/21/2024 | 05/29/2024 |  |  | 11/30/2025 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430949 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 was malposition and maps decreasing without adequate imaging to reposition appropriately. the surgeon pulled the impella across the valve and return to the operation room for removal. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in section: use of echocardiography for positioning of the impella catheter states: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ added- e4-no d4-corrected model number. f6/f8 should have been left blank. g1 reporting contact email.
26042 | 1220648-2024-11769 | 19416672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416672&pc=OZD | 05/22/2024 | 07/18/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468626 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support, and a high purge pressure alarm triggered during the support. tissue plasminogen activator was run through the purge line to manage the low purge flows and high purge pressure. there was no patient harm, and support with the implanted impella 5.5 pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for high purge pressure has been completed. the product was not returned for investigation. data logs were evaluated. there was a gradual increase in purge pressure observed leading to "purge system blocked" alarms. there was no motor current spikes observed. clinical mentions that issue plasminogen activator (tpa) was added to the purge which resolved the high purge pressure as seen in the data logs. therefore, as per the clinical information and data log review, the root cause of the high purge pressure issue was biomaterial. d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11769: a5 missing. d4 model number. h.6 code 4641 was reported incorrectly. new code was added to health effect - impact code.
26043 | 1220648-2024-11770 | 19416761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416761&pc=OZD | 05/22/2024 | 07/10/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025493209 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | Hypovolemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The user facility reported a female patient of unknown age ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced an increase in pulse and arterial pressures, which prompted the physician to start weaning the patient off inotropic/vasopressor support. neosynephrine drip was ordered at 0.4 micrograms/kilogram/minute. the circulating nurse unintentionally increased the neosynephrine to 20 micrograms/kg/minute. this resulted in a rapid increase in blood pressure with values as high as 300/160. during this time the physician removed the 14 french impella sheath and introduced the repositioning sheath into the arteriotomy. once hypertension was identified the patient was treated with 30 milligrams of hydralazine intravenously. the patient¿s hemodynamics quickly began to trend downward. an angiogram of the superficial femoral artery (sfa) was performed which showed perforation of the sfa at the arteriotomy. the patient exhibited massive blood loss and required rapid transfusion of blood products as well as fluids to treat hypovolemia. the decision was ultimately made to remove the impella cp and stent the arteriotomy site to stop bleeding. the patient expired in the intensive care unit. the physician believes that the patient's vasculature was calcified and believes that will the extensive hypertension the patient was exhibiting as well as introduction of the repositioning sheath into the arteriotomy that the vascular state became compromised. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings and cautions: physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation into the vascular damage - peripheral vessel issue has been completed. the device was not returned for investigation. based on the provided clinical details, the root cause of the vascular damage - peripheral vessels is most likely due to the patient¿s condition. as noted in the impella instructions for use: impella cp with smart assist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11770 and should not have been. h.6 code 4114 was reported incorrectly on manufacturer device report number 1220648-2024-11770. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report number 1220648-2024-11770 in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
26044 | 1220648-2024-11771 | 19416839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416839&pc=OZD | 05/22/2024 | 07/18/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460523 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there was concern over impella cp position being up against anterior wall. it was adjusted and now impella is too shallow and had to come down from p8 to p6 post adjustment. no alarms currently. impella looks shallow on echo but during adjustment sterile sleeve had been torn. the impella cannot be considered sterile for advancement and catheter needs to be cleaned with betadine, the whole sleeve and tegaderm catheter up to the cathlock. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for sterile sleeve damage has been completed. the impella device was not returned for evaluation. the root cause of the sterile sleeve damage issue was most likely use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11771. f6/f8-should have been left blank.
26045 | 1220648-2024-11772 | 19416924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19416924&pc=OZD | 05/22/2024 | 07/12/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  | 11/30/2025 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp was inserted successfully with no complications but during the pci (percutaneous coronary intervention) procedure patient developed increasing right limb pain on impella leg, starting from groin, and migrated down her thigh. the md decreased impella from p6 to p2 and initiated weaning in a quick manner since the physician thought the patient¿s leg was becoming ischemic. after impella was removed the md confirmed distal pulses with doppler. the patient said the leg started feeling better at this time. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation and data logs were not provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26046 | 1220648-2024-11773 | 19417480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19417480&pc=PYX | 02/21/2021 | 06/13/2024 |  | 05/29/2024 |  |  | 05/02/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | Abiomed complaints department received a publication entitled : "safety and efficacy of a percutaneously inserted ventricular support device purge solution heparin 25 u/ml" which was inclusive of 161 screened patients supported by 2.5, cp. 5.0, and impella rp. complications noted as bleeding and pump thrombosis: within our study, 2% of patients experienced a thrombotic event involving the pump motor. of the 63 bleeding events, 35 patients had clinically relevant bleeding (bleeding academic research consortium (barc) 3a and 3b). this report is for the impella rp. | [5 empty] | The impella devices were not received from the customer and therefore, an evaluation of the devices was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smart assist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ literature volume and authors: christine jiang, pharmd, bcps1 annals of pharmacotherapy 2021, vol. 55(2) 174 ¿180 © the author(s) 2020 article reuse guidelines: sagepub.com/journals-permissions doi: 10.1177/1060028020944831 journals.sagepub.com/home/aop https://doi.org/10.1177/1060028020944831 , misa stuart, pharmd, bcps1, charles makowski, pharmd, bcps1, douglas l. jennings, pharmd, facc, faha, fccp, fhfsa, bcps2,3 and long to, pharmd, bcps1. | The investigation into the access site bleeding ¿ major issue and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
26047 | 1220648-2024-11774 | 19417788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19417788&pc=OZD | 02/21/2021 | 06/27/2024 |  | 05/29/2024 |  |  | 05/02/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The abiomed complaints department received a publication entitled "safety and efficacy of percutaneously inserted ventricular support device purge solution heparin 25 u/ml" which was inclusive of 161 screened patients supported by either impella 2.5, impella cp, impella 5.0, or impella rp. complications noted as bleeding and pump thrombosis: within our study, 2% of patients experienced a thrombotic event involving the pump motor. of the 63 bleeding events, 35 patients had clinically relevant bleeding (bleeding academic research consortium 3a and 3b.) this report is for the impella cp. | [5 empty] | The impella devices were not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." literature authors: christine jiang, pharmd, bcps1 annals of pharmacotherapy 2021, vol. 55(2) 174 ¿180 © the author(s) 2020 article reuse guidelines: sagepub.com/journals-permissions doi: 10.1177/1060028020944831 journals.sagepub.com/home/aop https://doi.org/10.1177/1060028020944831 , misa stuart, pharmd, bcps1, charles makowski, pharmd, bcps1, douglas l. jennings, pharmd, facc, faha, fccp, fhfsa, bcps2,3 and long to, pharmd, bcps1. | The investigation for the reported access site bleed and thrombus has been completed. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 component code was reported on the previously submitted report. no component code is necessary for this complaint. code 4114 was reported incorrectly on the previously submitted report. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
26048 | 1220648-2024-11775 | 19417833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19417833&pc=OZD | 05/14/2024 | 06/24/2024 |  | 05/29/2024 |  |  | 05/14/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025441077 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 24-year-old female patient of unknown ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient had a thrombus burden observed after 20 days of support in the left ventricle (lv). the pump remained on for another 3 days and was then explanted. the patient was reported as stable. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section warnings and cautions: in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis . unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
26049 | 1220648-2024-11776 | 19417848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19417848&pc=OZD | 05/20/2024 | 06/24/2024 |  | 05/29/2024 |  |  | 05/20/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025482178 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The user facility reported a 75-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was believed that due to the change in waveforms and the lack of systemic anticoagulation that the impella cp had ingested a thrombus. the decision was made to explant the cp and implant a new device for continue support. the patient is stable. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
26050 | 1220648-2024-11777 | 19417900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19417900&pc=OZD | 05/21/2024 | 07/03/2024 |  | 05/29/2024 |  |  | 05/21/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025492797 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 65-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient's impella limb (left) was noticeably hypo-perfused, with mottling present and no doppler pulse. decision was made to perfuse the impella limb with external fem-fem bypass. flow was successfully restored to the leg with the external bypass. additionally, upon explanting the impella to escalate to an impella 5.5, the doctor did a direct approach double barrel technique. the cp caused a left ventricle perforation when manipulating the heart for 5.5 insertion. the perforation was repaired with a surgical suture. the patient was reported as stable. | [5 empty] | The investigation for the reported limb ischemia and ventricle perforation has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. based on the provided clinical details, the root cause of the ventricle perforation was most likely due to use as the perforation occurred while manipulating the heart for 5.5 insertion. the root cause of the limb ischemia could not be determined without additional clinical details. added-e4-no f6/f8-should have been left blank. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
26051 | 1220648-2024-11778 | 19418042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19418042&pc=OZD | 05/21/2024 | 11/22/2024 |  | 05/29/2024 |  |  | 05/21/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493287 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported delivery issues were encountered during implant of an impella cp pump for mechanical circulatory support to a patient experiencing acute myocardial infarction/cardiogenic shock. it was stated that the impella cp pump was unable to pass the iliac bifurcation and subsequently kinked during the delivery. a new impella cp pump was successfully implanted for planned support. there was no report of patient harm resulting from the delivery issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the pump was unable to pass the iliac bifurcation and subsequently kinked during the delivery as per the clinical description. the instructions for use for the impella cp states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ added-h6 impact code 4629 corrections to the initial mfr report: d4 model number.
26052 | 1220648-2024-11780 | 19418085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19418085&pc=OZD | 04/08/2024 | 07/22/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465193 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | Us complainant reported the patient had an attempted impella cp implant and it was aborted. when advancing the impella cp into impella peel away sheath physician was not able to advance. the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ | The investigation for difficulty inserting pump through introducer has been completed. the root cause of the difficulty inserting pump through introducer issue was not determined since product was not returned and insufficient clinical information was provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11780: d4 model number. b2 should have been blank. d4 model number. g1 reporting contact fax number missing. f6/f8-should have been left blank. h.6 codes 4641 and 4627 were reported incorrectly. new codes have been added to health effect - impact code.
26053 | 1220648-2024-11781 | 19418096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19418096&pc=OZD | 05/22/2024 | 07/31/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1590813 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | Battery Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/22/2024 |  |  |  |  | The complainant reported a patient undergoing coronary artery bypass was implanted with an impella 5.5 device for mechanical circulatory support. during support, the automated impella controller (aic) seemed to shut off without any incident. the aic restarted seconds later. the patient was slightly hypotensive but did not require any intervention and was stable after restart. the aic was exchanged. | [5 empty] | The investigation for the console (aic) battery failure alarm and the pump stop has been completed. data logs were reviewed which confirm that there was unexpected controller shutdown of the console on the day of the complaint. leading to the shutdown, there were numerous sau_powermod_1 errors as sau_motor_1 errors. the console was visually inspected in the lab with no revelation of an anomaly to the internal circuitry of the console. the console was run at maximum p-level for more than 4 hours in unsuccessful attempt to recreate the complaint failure. the root cause for the unexpected controller shutdown could not be determined due to the failure not being able to be reproduced. due to the shutdown, the pump was momentarily stopped but then automatically restarted back to its prior p-level (speed). there was no evidence in the logs for a battery failure alarm. f6/f8 should have been left blank in the initial report. h6 med dev prob code 2880 changed to 2885. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
26054 | 1220648-2024-11782 | 19418097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19418097&pc=OZD | 05/22/2024 | 07/16/2024 |  | 05/29/2024 |  |  | 05/22/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465288 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed an access site complication. the impella was emergently placed during cardiopulmonary resuscitation. oozing from all access sites, venous and arterial, and the impella access site developed a hematoma. manual pressure held to establish hemostasis and one unit of blood products was given. the patient was transported to the unit in stable condition. in unit update noted no oozing and site area soft to touch. patient noted to be recovering well with plans to wean and explant the impella cp pump as planned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding/hematoma has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related, oozing from all access sites, venous and arterial. this includes impella site which is developing hematoma as per the clinical description. h6 added component code 925 d4 model number corrected.
26055 | 1220648-2024-11783 | 19418107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19418107&pc=OZD | 05/23/2024 | 07/23/2024 |  | 05/29/2024 |  |  | 05/23/2024 | 05/29/2024 |  |  |  |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485015 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem |  |  |  | Hemolysis | Ischemia | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [34 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/23/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support when the patient thrashed about and broke the purge sidearm in half. the physician was advised to explant the pump. the motor showed signs of saturation, pulling 3.3/3.2 at p3. the pump then stopped. the pump was explanted. | [5 empty] | D.9 updated as the pump and purge cassette were returned for investigation. the investigation has been completed. according to the clinical description, on the third day of impella cp support, the patient was being weaned they became agitated and broke the purge sidearm joint. after purge system open alarms and signs of saturation were recorded, a pump stop occurred while the patient was being prepped for explant. product evaluation confirmed damage to the sidearm joint, and the presence of biomaterial in the purge gap and on the impeller. data log analysis confirmed the presence of a motor current spike and purge related trend. additionally, flow saturation was observed after the low purge pressure event and prior to the pump stop. the cause of the pump stop was damage to the reservoir to filter joint. the saturated flow and fluid leak were intermediate failures that caused the pump to stop. the root cause of the intermediate failures was damage to the reservoir to filter joint cause by patient movement. a5. revised ethnicity was previously entered incorrectly on manufacturer report number 1220648-2024-11783. b.7 revised as this information was inadvertently omitted from manufacturer report number 1220648-2024-11783. d.4 revised model number from manufacturer report number 1220648-2024-11783 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer report number 1220648-2024-11783 and should not have been. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer report number 1220648-2024-11783. h6 health effect-clinical code: 1942 ischemia, 1886 hemolysis and 4440 thrombosis entered in error on the initial report. replaced with 4582 no clinical signs or symptoms h6 health effect impact code-4641 unexpected medical intervention entered on the initial report in error. h6 medical device problem code: 2993 was entered in the initial report in error. replacing with correct failure modes, 1503 pumping stopped, 2464 infusion or flow problem and 1250 fluid leak. h6 added component code. | As noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." additional manufacturer narrative instructions for use were revised from the initial manufacture device report number 1220648-2024-11783 in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26056 | 1220648-2024-11784 | 19418114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19418114&pc=OZD | 05/17/2024 | 07/23/2024 |  | 05/29/2024 |  |  | 05/23/2024 | 05/29/2024 |  |  | 08/31/2024 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023260315 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support when the patient's condition suddenly changed. the physician placed an impella cp and extracorporeal membrane oxygenation. the impella was removed the next day because it was suspected that the lower extremity in which the impella was inserted was ischemic. the patient is stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11784: f6/f8-should have been left blank. g1 reporting contact fax number missing.
26057 | 1220648-2024-11785 | 19419307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19419307&pc=OZD | 05/23/2024 | 07/25/2024 |  | 05/29/2024 |  |  | 05/23/2024 | 05/29/2024 |  |  | 10/31/2025 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABOIMED, INC |  |  | US |  | 2024426444 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 58-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had a bleeding tendency and the patient's circulation was improved. the impella was removed. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11785: f6/f8-should have been left blank.
26058 | 1220648-2024-11786 | 19419371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19419371&pc=OZD | 04/12/2022 | 03/25/2025 |  | 05/29/2024 |  |  | 04/12/2022 | 05/29/2024 |  |  | 12/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164199 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2022 |  |  |  |  | The user facility reported a 46-year-old male on impella cp for mechanical circulatory support. it was reported the patient experienced a blood clot in the apex of the heart caused by myocarditis moved to the superior mesenteric artery (sma) as the impella cp was removed. additionally, the patient had slight hemolysis, but the hemolysis has subsided due to the administration of haptoglobin and the addition of ecmo (extracorporeal membrane oxygenation) at night. the patient was successfully weaned. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26059 | 1220648-2024-11787 | 19419963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19419963&pc=OZD | 04/20/2022 | 03/25/2025 |  | 05/29/2024 |  |  | 04/20/2022 | 05/29/2024 |  |  | 12/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183360 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported a 72-year-old female on impella cp for mechanical circulatory support. it was reported the patient possibly had heparin induced thrombocytopenia. the patient went on bypass without complications. after grafts were completed, significant aortic bleeding was noted and aortic graft was used to reinforce the aorta. manual pressure was applied and the patient was given 14 blood products and 4600ml of cell saver. the impella cp device had suction that would not resolve. echo was performed and the impella cp device was advanced, but suction and position were never corrected. due to almost 12 days of use and patient being vitally stable the impella cp was pulled back into the aorta by cardiovascular surgeon. the patient was successfully weaned. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. section h6, medical device problem code 1248 should have been reported in the initial mfr 1220648-2024-11787.
26060 | 1220648-2024-11788 | 19420786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19420786&pc=OZD | 08/16/2022 | 03/25/2025 |  | 05/29/2024 |  |  | 08/16/2022 | 05/29/2024 |  |  | 04/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081109 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 72-year-old male patient implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock. while on support, the patient experienced access site bleeding that was resolved by blood transfusion and enhanced fixation. the patient was successfully weaned. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26061 | 1220648-2024-11789 | 19421012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19421012&pc=OZD | 09/27/2022 | 03/24/2025 |  | 05/29/2024 |  |  | 09/27/2022 | 05/29/2024 |  |  | 06/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094868 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a 75-year-old female was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced access site bleeding with hematoma and received blood products. no further information was provided. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26062 | 1220648-2024-11790 | 19421270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19421270&pc=OZD | 12/30/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 12/30/2021 | 05/29/2024 |  |  | 04/30/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070608 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a 58 year old female patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient was observed to have a clot in their leg. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
26063 | 1220648-2024-11791 | 19421302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19421302&pc=OZD | 12/31/2021 | 05/29/2024 |  | 05/29/2024 |  |  | 12/31/2021 | 05/29/2024 |  |  | 10/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150976 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2021 |  |  |  |  | The complainant reported the 42 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support occasional suction alarms were noted. an observed hematoma was resolved by applying pressure. the patient was successfully weaned from support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26064 | 1220648-2024-11792 | 19421323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19421323&pc=OZD | 01/01/2022 | 05/29/2024 |  | 05/29/2024 |  |  | 01/01/2022 | 05/29/2024 |  |  | 10/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151886 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/01/2022 |  |  |  |  | The complainant reported a 75 year old female patient was placed on bilevel positive airway pressure (bipap) at the medical facility. while in the catheterization laboratory, an angiogram revealed significant disease to the proximal left anterior descending artery (plad) and the decision was made to implant an impella cp for mechanical circulatory support. pre-close technique was utilized on the right femoral artery (rfa) and the impella was implanted. the patient tolerated the procedure well despite experiencing groin site bleeding through the process. the patient was transfused with two (2) units of blood. the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26065 | 1220648-2024-11793 | 19421405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19421405&pc=OZD | 04/07/2022 | 03/24/2025 |  | 05/29/2024 |  |  | 04/07/2022 | 05/29/2024 |  |  | 01/31/2024 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187105 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding and hemostasis was achieved with perclose sutures and manual pressure. although standard care was presented to the patient for the access site bleeding, the device was also removed. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26066 | 1220648-2024-11794 | 19421558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19421558&pc=OZD | 04/08/2022 | 05/29/2024 |  | 05/29/2024 |  |  | 04/08/2022 | 05/29/2024 |  |  | 12/31/2023 |  | 05/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174322 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported that the 74 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support, hemolysis and access site bleeding occurred. the patient went into cardiac arrest in the evening and cardiopulmonary resuscitation (cpr) was performed. pulseless electrical activity (pea) was noted on the monitor and return of spontaneous circulation (rosc) was obtained; however, the patient remained unresponsive. the patient was transported to the catheterization laboratory for impella placement and high risk percutaneous coronary intervention (hrpci). it was reported that the impella device was left in place and the patient received one (1) unit of blood prior to transfer to the intensive care unit (icu) to receive hypothermia protocol. oozing and hematoma were observed at the impella access site. manual pressure, angle matching, femstop and forward tension suturing were utilized to resolve this issue. it was reported that the patient¿s laboratory results returned hemolyzed. echocardiogram (echo) imaging identified the tip of the pump was resting on the mitral valve. the medical professional repositioned and torqued the impella into a better location at the center of the apex. flows improved and the pump speed level was reduced to p4 from p7. weaning was successful, the impella device was explanted, and the procedural outcome is the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
26067 | 1220648-2024-11682 | 19403697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19403697&pc=OZD | 10/14/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/14/2021 | 05/28/2024 |  |  | 02/28/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022054330 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding around repo sheath after peel-away exchange. tightening perclose sutures and manual pressure around sheath did not help; therefore, the decision was made to explant the pump. perclose sutures failed on pump removal. 14f sheath was inserted with bleeding around the sheath. hemostasis was achieved with unplanned surgical cutdown. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26068 | 1220648-2024-11683 | 19403728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19403728&pc=OZD | 10/10/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/15/2021 | 05/28/2024 |  |  | 02/28/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022049446 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient became septic, though it is unknown how this issue was addressed. the patient later experienced atrial fibrillation, which was addressed with cardioversion. the patient remained on impella support until the device was successfully weaned and escalated to an impella 5.5. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26069 | 1220648-2024-11684 | 19403858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19403858&pc=OZD | 10/16/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/16/2021 | 05/28/2024 |  |  | 06/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding at the access site. therefore, the pump was removed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26070 | 1220648-2024-11685 | 19403870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19403870&pc=OZD | 10/16/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/16/2021 | 05/28/2024 |  |  | 07/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116254 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was noted, resulting in explant of the impella. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26071 | 1220648-2024-11686 | 19403903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19403903&pc=OZD | 10/18/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/18/2021 | 05/28/2024 |  |  | 06/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094939 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that the patient had gastrointestinal bleeding (gib) with melena. systemic heparin was stopped until after an esophagogastroduodenoscopy (egd) and colonoscopy were performed, but heparin was kept in the purge flow. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26072 | 1220648-2024-11687 | 19403998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19403998&pc=OZD | 10/29/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/29/2021 | 05/28/2024 |  |  | 07/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022113965 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2021 |  |  |  |  | The complainant reported a patient was attempted to be implanted with an impella 5.0 device for hemodynamic support. the implant procedure was aborted; the device could not be implanted due to the patient's vascular anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26073 | 1220648-2024-11688 | 19404044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404044&pc=PYX | 04/29/2024 | 06/12/2024 |  | 05/28/2024 |  |  | 04/29/2024 | 05/28/2024 |  |  | 11/30/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025439634 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient post cardiothoracic surgery in a critically ill state was implanted with an impella rp device for bipella mechanical circulatory support and experienced access site bleeding. during sheath exchange oozing/bleeding was observed at the impella rp access site. two mattress sutures and a 15-minute manual pressure were applied to resolution. it was noted there was no hematoma; however, the patient required two units of packed red blood cells (prbcs) post impella rp placement based on hemoglobin less than 6 lab result. the patient had returned back to the unit in the critically ill state with a grim prognosis. in the unit, groin access site was checked with no oozing or hematoma noted. the outcome noted the patient subsequently expired after 14 hours of support after the decision was made to withdraw care. medical safety review outcome noted there is not enough information to associate use of the device with the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related since the issue was resolved by applying mattress suture x 2 placed and 15 min manual pressure. d4-updated model number.
26074 | 1220648-2024-11689 | 19404069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404069&pc=OZD | 12/30/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 12/30/2021 | 05/28/2024 |  |  | 05/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022086982 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella flow was gradually decreased in auto mode, and when the team switched to performance level, they were unable to increase the flow over p3 without suction events. in addition, the impella performance decreased to p2. there were continuous suction alarms that did not break when dropping to p1. a fluid bolus and blood products were given. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26075 | 1220648-2024-11690 | 19404079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404079&pc=OZD | 12/30/2021 | 03/17/2025 |  | 05/28/2024 |  |  | 12/30/2021 | 05/28/2024 |  |  | 10/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152646 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. following explant of the impella, 2 perclose units were deployed, but there was still oozing at the insertion site, and this continued after a 3rd perclose was used. a 6fr sheath was placed over the wire, creating hemostasis, and then a 6fr angioseal was deployed with no further oozing at the site. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26076 | 1220648-2024-11692 | 19404255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404255&pc=PYX | 02/09/2020 | 07/03/2024 |  | 05/28/2024 |  |  | 05/02/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The abiomed complaints team received a publication from american society for artificial organs (asaio) 10 impella rp devices used between 2015-2018. entitled: "contemporary outcomes of temporary percutaneous right ventricular support in a heterogeneous right ventricular failure cohort". there were 5 rp explanted, 3 died, and 1 required device exchange. the malfunction that prompted device exchange was not noted. no further information is available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported pump stop and death has been completed. neither the impella nor the data logs were returned for evaluation. therefore, without additional clinical details, the root cause of the pump stop or the death could be determined. f6/f8-should have been left blank in the initial report.
26077 | 1220648-2024-11694 | 19404305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404305&pc=PYX | 07/26/2021 | 06/28/2024 |  | 05/28/2024 |  |  | 05/02/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 complaints forwarded publication entitled: "use of impella devices for acute cardiogenic shock in the perioperative period of cardiac surgery" which detailed 11 impella patients, 7cp, 1 rp, 1 cp-rp bipella, and 2 5.0 pumps. the publication noted "device related complications included varying degrees of hemolysis in 8 patients (73%) and device malfunction in 1 patient (9%). patient 1 (intraoperative) with the device malfunction had both rp and impella cp support as well as va-ecmo." unspecify allegation for the device malfunction. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into the hemolysis and the pump stop issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the hemolysis and the pump stop issues cannot be determined since the product was not returned and insufficient clinical details were provided.
26078 | 1220648-2024-11696 | 19404394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404394&pc=OZD | 10/06/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/06/2021 | 05/28/2024 |  |  | 06/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104845 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had hemolysis. urine started to turn red. echo was quickly received where it appeared too deep, and device was pulled back two (2) cm and locked. impella was brought to p6 to break up some of the hemolysis. after an hour the impella had been repositioned and p level dropped urine is still bright red and then started to clear up. the following day, the clinical staff decreased p-level and removed all heparin from purge and systemically due to hemolysis. nurse stated that hemolysis has cleared up and the patient will be started back on heparin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26079 | 1220648-2024-11695 | 19404397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404397&pc=PYX | 04/01/2023 | 07/16/2024 |  | 05/28/2024 |  |  | 05/02/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP EU |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On (b)(6)2024 complaints forwarded abstract from journal heart and lung transplantation entitled: "¿long-term¿ use of impella - safe to do?" in which impella supports from 5/2017 to 5/2021 were analyzed. the complications noted during impella rp support were device migration, pump thrombosis, and general device malfunction. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ zeschky, c., ahmed, h., alayyar, m., jothidasan, a., husain, m., stock, u., & smail, h. (2022). ¿long-term¿ use of impella - safe to do? the journal of heart and lung transplantation, 41(4). https://doi.org/10.1016/j.healun.2022.01.1603 | The investigation for the positioning issues has been completed. complaint device not returned for investigation. device serial number unknown to reviewed the data logs. insufficient clinical data was provided. the root cause of the positioning issue cannot be determined.
26080 | 1220648-2024-11697 | 19404460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404460&pc=PYX | 01/23/2023 | 06/25/2024 |  | 05/28/2024 |  |  | 05/02/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A complaint from the literature review 'safety and efficacy of impella rp support for acute right ventricular failure complicated by cardiogenic shock: post market approval subanalysis of the cvad registry" was received. the review monitored 110 patients over 3 years who were implanted with an impella rp pump. symptoms encountered by an unknown quantity of those patients included major bleeding and hemolysis. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation into the access site bleeding ¿ major and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue and the hemolysis issue was not determined since product was not returned and insufficient clinical information provided, and data logs were not available.
26081 | 1220648-2024-11698 | 19404493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404493&pc=OZD | 10/08/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/08/2021 | 05/28/2024 |  |  | 07/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119542 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had gi bleeding. therefore, they switched to bicarbonate. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26082 | 1220648-2024-11699 | 19404623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404623&pc=OZD | 05/17/2024 | 07/12/2024 |  | 05/28/2024 |  |  | 05/17/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479169 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. the implant was aborted due to the patient's anatomy. the device was unable to pass through the patient's subclavian calcification. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. f6/f8 should have been left blank in the initial report.
26083 | 1220648-2024-11700 | 19404632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404632&pc=OZD | 10/08/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/08/2021 | 05/28/2024 |  |  | 07/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119603 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s urine was dark red in color and laboratory results identified hemolysis. echocardiogram observed that the impella was caught in the papillary muscle and was unable to be repositioned. the impella cp was removed due to positioning and hemolysis was unresolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist. section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26084 | 1220648-2024-11701 | 19404654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404654&pc=PYX | 03/28/2022 | 07/03/2024 |  | 05/28/2024 |  |  | 05/02/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Respiratory Failure | [8 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complaints team was forwarded a publication titled, "temporary right ventricular assist device support for acute right heart failure: a single-center experience". which outlines 43 patients implanted with right ventricular assist devices (rvad): 32 protek duo, 6 impella rp, and 4 surgical rvad. the publication noted that 23 patients suffered a major bleed which required blood transfusions or intervention, 3 patients had a stroke, 18 patients had a new onset renal failure requiring at least temporary renal replacement therapy, 2 patients were inotrope dependent at discharge, and 25 patients experienced prolonged respiratory failure requiring tracheostomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported access site bleed, renal failure, stroke, vascular damage, and ventricular arrhythmia has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the reported issues could not be determined. the instructions for use referenced in the initial report is for the impella rp system with automated impella controller in the section: potential adverse events (united states). d4-corrected model number. e1 initial reporter fist and last name . f6/f8-should have been left blank in the initial report. g1 reporting contact email.
26085 | 1220648-2024-11702 | 19404664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404664&pc=PYX | 01/23/2023 | 06/20/2024 |  | 05/28/2024 |  |  | 05/02/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Literature | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A complaint from the literature review 'safety and efficacy of impella rp support for acute right ventricular failure complicated by cardiogenic shock: post market approval subanalysis of the cvad registry" was received. the review monitored 110 patients over 3 years who were implanted with an impella rp pump. during that time, an unknown quantity of those patients experienced unspecified malfunctions with their impella device. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the reported pump stop has been completed. the pump nor the data logs were returned for evaluation. the cause of the pump stop issue was not determined since product logs nor sufficient clinical information was provided. the instructions for use for the impella rp smart assist system with automated impella controller and rp flex with smart assist system with automated impella controller state: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h5-labeled for single yes corrected for yes. h6-corrected med dev prob code to 1503. h8-usage of device selected initial use of device.
26086 | 1220648-2024-11703 | 19404668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404668&pc=OZD | 05/17/2024 | 07/10/2024 |  | 05/28/2024 |  |  | 05/17/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489559 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VA EXTRA-CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp pump and began to experience low flows during venous-arterial extra-corporeal membrane oxygenation-impella support. as a result of the issue, the impella cp pump was explanted and new one placed for ongoing support. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the pump was not returned. data logs were reviewed and low flows and suction were confirmed with a downward trend in placement signal. based on the data log analysis and the clinical details provided, the root cause of the low pump flow is most likely due to the patients condition. h.6 added code 4629 as it was inadvertently left off manufacturer device report number 1220648-2024-11703.
26087 | 1220648-2024-11704 | 19404714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404714&pc=OZD | 05/07/2024 | 06/13/2024 |  | 05/28/2024 |  |  | 05/07/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453609 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The complainant reported a 76-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that after the cp placement, the patient had no pulse in the right lower extremity. medical staff placed 9f and 5f distal reperfusion sheaths fem-fem, which restored the patient¿s blood flow. the patient eventually expired. there is not enough information to associated the patient's outcome with the use of the impella device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26088 | 1220648-2024-11705 | 19404725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404725&pc=OZD | 11/16/2022 | 07/25/2024 |  | 05/28/2024 |  |  | 05/23/2024 | 05/28/2024 |  |  | 08/31/2024 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC |  |  | US |  | 2023251037 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Death | Yes | No | Failure to Advance | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2024 |  |  |  |  | The user facility reported a 71-year-old male patient of unknown ethnicity post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that multiple complications were encountered during impella cp support. it was initially noted that during insertion, access could not be obtained through the aortic valve and direct access was required. however, following the complication, bleeding from the "hiatus" that had formed at the root was difficult to manage and the cardiopulmonary bypass was terminated. after several hours, multiple cerebral infarctions were observed, and the patient passed away shortly thereafter. | [5 empty] | The investigation for the delivery issue, vascular damage, and stroke has been completed. the impella device was not returned for evaluation. clinical details were sufficient to determine the root cause of the delivery issues. the root cause of the delivery issue was patient condition related to tavr interaction. however, without additional clinical details, the root cause of the stroke and vascular damage could not be determined. added-h6 med dev prod code 2682 d4-corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
26089 | 1220648-2024-11706 | 19404746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404746&pc=OZD | 05/13/2024 | 06/28/2024 |  | 05/28/2024 |  |  | 05/13/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025453885 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support when they had a middle cerebral artery stroke. heparin was withheld due to jp (jackson-pratt) drain output. heparin was restarted and support continued. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: us complainant reported the patient was on impella 5.5 support when they had a middle cerebral artery stroke. heparin was withheld due to jp (jackson-pratt) drain output. heparin was restarted and support continued. | The investigation into the stroke/cva and the access site bleeding ¿ major issues been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as clinical information related to stroke/cva was not provided, the root cause could not be determined.
26090 | 1220648-2024-11707 | 19404756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404756&pc=OZD | 05/18/2024 | 07/10/2024 |  | 05/28/2024 |  |  | 05/18/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient admitted in ami/cgs (acute myocardial infarction/cardiogenic shock). the support had signs of hemolysis and kidney function issues. no urine was being made and the pump position was unknown so the team explanted the pump after 2 days of support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs showed about 28 hours into support, the suction alarms were triggered corresponded with the time hemolysis was reported but resolved within 1.5 hours. there were no other issues on the logs. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
26091 | 1220648-2024-11717 | 19404856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404856&pc=OZD | 10/10/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/10/2021 | 05/28/2024 |  |  | 06/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022120524 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that required addition of blood products to resolve. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26092 | 1220648-2024-11724 | 19404885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404885&pc=OZD | 05/20/2024 | 07/17/2024 |  | 05/28/2024 |  |  | 05/20/2024 | 05/28/2024 |  |  | 09/30/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410380 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | Bradycardia | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The complainant reported a 54-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported after the cp implant and during transport, the cp was pulled back into the aorta. the physician made multiple unsuccessful attempts to advance the cp back across the aortic valve and then decided to explant the cp. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issues has been completed. the product was not returned and data logs are not available. per the clinical review, the root cause of the positioning issue was use related . the pump was pulled out inadvertently by the patient. h6 health effect codes 1751 and 4868 should be removed. these were patient's symptoms prior to implantation of our device. they were entered in error. h6 health impact code: 4624 and 4625 should be removed. these codes are not related to the malfunction. they were entered in error. f6 and f8 should have been blank.
26093 | 1220648-2024-11708 | 19404959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19404959&pc=OZD | 05/20/2024 | 07/17/2024 |  | 05/28/2024 |  |  | 05/20/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478375 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient was weaned off heart-lung bypass and onto impella support with no issues. while closing the chest, a red "impella stopped" alarm presented and restarting the impella was unsuccessful. the automated impella controller (aic) was changed and displayed the same alarm, with high motor current alarm. the 5.5 was explanted. the impella was examined by the surgeon and a calcium piece was lodged in the motor. the decision was made not to replace the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Impella 5.5 was returned for evaluation. visual inspection was performed and no abnormalities were seen. a small blue fiber was present on the impeller blade. no other mechanical abnormalities were noted. the pump was run in a bucket of water for approximately 10 minutes and no pump stop occurred. data logs from field show a pump stop with a "impella stopped - motor current high" alarm and a motor current spike. clinical details noted that before pump was implanted, patient underwent coronary arterial bypass graft (cabgx4), mitral valve and atrial valve replacement. once pump was explanted from patient, physician noted that a piece of calcium was present on the motor and was removed. the root cause of the pump stop was due to ingested biomaterial. d9 date device returned to manufacturer added f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
26094 | 1220648-2024-11715 | 19405072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405072&pc=OZD | 05/18/2024 | 07/09/2024 |  | 05/28/2024 |  |  | 05/18/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The complainant reported a 79-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported an access site ooze/bleed. the patient was given a forward tension suture and two units of blood. additionally, during placement of the .018 wire and pump, the patient experienced ventricular tachycardia and was defibrillated twice. the patient remains on impella support. | It was further reported that the patient expired on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported ventricular tachycardia (vt) and access site bleed has been completed. the investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt and access site bleed could not be determined. added- b5 mso review, h5 clinical code 2095, h6 impact code 4647 added- d6b f6/f8 should have been left blank in the initial report.
26095 | 1220648-2024-11714 | 19405109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405109&pc=OZD | 10/09/2021 | 05/28/2024 |  | 05/28/2024 |  |  | 10/09/2021 | 05/28/2024 |  |  | 06/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108057 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, when the dressing change was performed, the site looked to be oozing. however, upon changing the dressing, the site was intact and no active oozing had been noted. later, oozing at the groin was noted and was controlled by administration of lido w/ epinephrine iv heparin added overnight for sub-therapeutic activated clotting time. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26096 | 1220648-2024-11722 | 19405149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405149&pc=OZD | 05/15/2024 | 07/26/2024 |  | 05/28/2024 |  |  | 05/15/2024 | 05/28/2024 |  |  | 11/30/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430603 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and experienced ventricular tachycardia. the patient was defibrillated and return of spontaneous circulation was achieved. additionally, while transporting the patient to the intensive care unit (icu), access site bleeding was found. hemostasis was achieved by withholding heparin, and sutures were replaced in the icu. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and the ventricular tachycardia (vt) has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, hemostasis was achieved by withholding heparin and adding sutures as per the clinical description. without additional clinical details, the root cause of the vt could not be determined. added- h6 component code 925.
26097 | 1220648-2024-11719 | 19405293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405293&pc=OZD | 04/03/2023 | 08/05/2024 |  | 05/28/2024 |  |  | 05/21/2024 | 05/28/2024 |  |  | 11/30/2024 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2024314058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | Abiomed received information from the stemi dtu study which reported a patient was implanted with an impella cp in (b)(6) 2023 which was successful for one day. later, the study team relayed a "probable" relationship of a hematoma/anemia to the use of the impella cp. the hematoma was treated with manual hold/pressure and was not a reportable event. in (b)(6) 2024, the study team reported additional events of access site bleeding and "torsion" of the pump for this patient. the kink/torsion and blood loss prompted pump explant with hypotension from the kinked and underperforming pump. the patient received two units of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and product damage has been completed. the impella has not been returned for evaluation. clinical details were not sufficient to determine the root cause. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of product damage was not determined as insufficient clinical details were provided and no product and the data logs were returned for analysis. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11719 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11719 was submitted. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report 1220648-2024-11719. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the previously submitted manufacturer device report 1220648-2024-11719 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
26098 | 1220648-2024-11709 | 19405324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405324&pc=OZD | 10/08/2021 | 05/28/2024 |  | 05/28/2024 |  |  | 10/08/2021 | 05/28/2024 |  |  | 06/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093976 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/08/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding; therefore, the pump was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26099 | 1220648-2024-11726 | 19405465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405465&pc=OZD | 05/15/2024 | 06/21/2024 |  | 05/28/2024 |  |  | 05/15/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493197 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 64-year-old male patient of unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after successful implantation of an impella cp, the patient experienced ventricular fibrillation. the patient was shocked and return of spontaneous circulation was achieved. the physician decided to remove the impella due to poor distal perfusion and hematoma. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation for the reported limb ischemia, ventricular tachycardia, and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia, ventricular tachycardia, and access site bleed could not be determined as the device was not returned nor sufficient clinical details.
26100 | 1220648-2024-11721 | 19405568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405568&pc=OZD | 05/21/2024 | 06/06/2024 |  | 05/28/2024 |  |  | 05/21/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported at the conclusion of the support for a patient experiencing acute myocardial infarction/cardiogenic shock implanted with an impella cp device, the automated impella controller blue clip broke off when the team removed the purge cassette. there was no report of harm to the patient. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. the issue of broken purge disk retainer was confirmed. the cause of the issue was a broken purge retainer. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26101 | 1220648-2024-11716 | 19405726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405726&pc=OZD | 05/19/2024 | 07/11/2024 |  | 05/28/2024 |  |  | 05/19/2024 | 05/28/2024 |  |  | 09/30/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025445381 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported a 52-year-old male patient undergoing high risk coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp was removed at request of surgeon because of unseen, suspected aortic (valve) insufficiency (ai) caused by impella. additional information received that transesophageal echocardiography (tee) was done in the operating room to reposition the impella cp to correct position because the surgeon wanted the impella shallowed out during the surgery to manipulate the heart. no ai was seen under tee, once back to the icu the cardiologist that assisted in repositioning the impella requested a tte at beside and picture of depth of impella was not very good due to the patient¿s size, but there was no ai seen there either. the cardiologist took the patient back to the catheterization lab to confirm impella placement and position and under fluoroscopy there were no signs of the suspected ai that the surgeon was concerned about. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarm if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported heart valve damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the heart valve damage could not be determined. f6/f8-should have been left blank. h6-impact code changed from 4627 to 4628.
26102 | 1220648-2024-11725 | 19405932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405932&pc=OZD | 05/21/2024 | 07/16/2024 |  | 05/28/2024 |  |  | 05/21/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466364 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 81-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having a percutaneous coronary intervention procedure. during explant, the impella remained stuck within the peel away sheath as the pump was being removed. the doctor tried to regain single access entry into the arteriotomy with 0.018 wire, unable to advance. therefore, the doctor decided to remove sheath and impella device together as one unit. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ | The investigation into the removal issues has been completed. pump, guidewire and peel-away sheath were returned. the pump was returned stuck in the sheath along with guidewire which was used as a troubleshooting for removal of pump. per the clinical details, patient had severe stenosis and therefore a long sheath was used. the pump outlet was stuck at the kink of the peel-away sheath. therefore, per the clinical details provided about patients vasculature and the product investigation, the root cause of the difficulty removing pump through introducer issue was a introducer sheath kink caused due to tortuous vessels. the guidewire was used to gain single access to assist in pump removal which was stuck in the peel away sheath during explant. this is not suggested in the impella ifu as a troubleshooting step. therefore, the root cause of the product damage (guidewire damage - peripheral vessel) issue was use related to improper technique. added d9 device returned and date h6 corrected health impact code. removed 4641 and added 4627. d4 corrected model number. g1 corrected reporting contact email address. f6 and f8 should have been blank.
26103 | 1220648-2024-11713 | 19405937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19405937&pc=OZD | 10/09/2021 | 05/22/2025 |  | 05/28/2024 |  |  | 10/09/2021 | 05/28/2024 |  |  | 04/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078693 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed bleeding from an unknown source. despite blood products being given (prbc/plasma/platelets/cryo), the patient was taken to the operating room to attempt to stop the bleeding. heparin was not given due to bleeding from an unknown source, unexplained coagulopathy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26104 | 1220648-2024-11720 | 19406169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19406169&pc=OZD | 10/10/2021 | 03/25/2025 |  | 05/28/2024 |  |  | 10/10/2021 | 05/28/2024 |  |  | 10/31/2022 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021004859 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2021 |  |  |  |  | The complainant reported a patient in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the impella continued to suction on p1, and massive stroke was noted. femostop was applied to the left groin. the patient¿s condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (stroke/cva) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26105 | 1220648-2024-11718 | 19406203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19406203&pc=OZD | 05/19/2024 | 07/19/2024 |  | 05/28/2024 |  |  | 05/19/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025488000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported a 45-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the bedside echocardiogram showed the impella cp to be very shallow and cannula flipped. the physician attempted to advance the impella into the left ventricle (lv) unsuccessfully. several attempts were made using ultrasound guidance, however, was unsuccessful. the decision was made to exchange the impella cp. the was no patient harm reported. | [5 empty] | The investigation into the positioning issues has been completed. the product was not returned for investigation. data log was not provided for this case. according to clinical notes, the pump was found to be shallow and flipped. during repositioning, the pump was fully inserted into the aorta. after several attempts to reposition the device, the doctor replaced the pump. the cause of the positioning issue was most likely use related. f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26106 | 1220648-2024-11723 | 19406231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19406231&pc=OZD | 05/15/2024 | 07/16/2024 |  | 05/28/2024 |  |  | 05/15/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 93-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the long peel away sheath became kinked during insertion. the impella implant was aborted after multiple attempts to guide the impella into the left ventricle. there was no patient harm reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues and the introducer damage has been completed. the product was not returned for investigation. during insertion, the log peel-away sheath was kinked, and the doctor was unable to place the impella. patient reported with tortuous vessels. the cause of the introducer damage issue was related to the patient's tortuous vessel condition, which also caused an issue delivering the pump. added-h6 med dev prob code 2889 f6/f8 should have been left blank in the initial report.
26107 | 1220648-2024-11727 | 19406438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19406438&pc=OZD | 05/15/2024 | 07/15/2024 |  | 05/28/2024 |  |  | 05/15/2024 | 05/28/2024 |  |  | 08/31/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had ectopy. the staff dropped the p-level to p5, but the patient had more frequent ectopy that resolved when increased to p 7. additionally, the patient later had ventricular tachycardia and required cardiopulmonary resuscitation. furthermore, the impella had positional suction alarms that were occurring causing hemolysis. the impella appeared to be against the lateral wall and the doctor could not successfully reposition. the pump was removed and replaced with a new impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported arrhythmia, low pump flow, and positioning issue has been completed. complaint device was not returned for investigation. data logs were reviewed finding no motor current spike observed. many suction alarms occurred. many positioning alarms in aorta and ventricle occurred. pump flow dropped to p-2. the cause of the positioning issue most likely was use related due to improper technique that led to hemolysis issue and low pump flow issue as impella appeared to be against lateral wall and against mitral apparatus. the cause of both hemolysis issue and low pump flow most likely were improper positioning due to improper technique. the instructions for use for the impella cp with smartassist system in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation) in accordance with updated procedures, section h10 has been revised from the initial submission. f6/f8 should have been left blank in the initial report.
26108 | 1220648-2024-11711 | 19406663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19406663&pc=OZD | 04/01/2024 | 07/05/2024 |  | 05/28/2024 |  |  | 05/06/2024 | 05/28/2024 |  |  | 07/31/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443640 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating an allegation of hemolysis: ¿very profoundly coagulopathic in the or with significant product transfusion with 6u prbc's, 4u plt's, 8 ffp, and 3 cryo. also loaded with brilinta. 2u prbc's since arriving to icu.¿ the patient is unknown at this time. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were returned and reviewed. there was no motor current spike seen. there was one suction alarm seen around the time of the alleged hemolysis event. the placement signal was seen to be spread throughout the case with no positioning issue. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided and data log analysis was inconclusive.
26109 | 1220648-2024-11731 | 19406844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19406844&pc=OZD | 05/21/2024 | 07/18/2024 |  | 05/28/2024 |  |  | 05/21/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478418 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. after the impella implant, doctor was struggling to advance the catheter. once the impella was in the left ventricle, the automated impella controller alerted the impella was in the aorta. after multiple attempts to ensure proper position, the doctor opted to replace the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical issue has been completed. product was not returned for review. data logs were reviewed and no psnr observed. all 5s optical parameters were stable and no motor current spike was found. false position in aorta alarm was present for 6 minutes. device was explanted due to this issue. the cause of the false position in aorta alarm was most likely patient condition related. f6/f8 should have been left blank in the initial report.
26110 | 1220648-2024-11732 | 19407063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19407063&pc=OZD | 05/21/2024 | 06/28/2024 |  | 05/28/2024 |  |  | 05/21/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492737 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and after the implant, there was access site bleeding. the site was redressed and the bleeding resolved. additionally, the pump came into the aorta and the staff was unable to advance the impella. the impella was replaced. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues and the access site bleeding ¿ minor issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding - minor is most likely due to use as the bleeding resolved when the site was redressed. the root cause of the positioning issue is most likely due to use as the pump was pulled back into the aorta during cpr.
26111 | 1220648-2024-11730 | 19407835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19407835&pc=OZD | 04/29/2024 | 06/04/2024 |  | 05/28/2024 |  |  | 04/29/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486142 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced limb ischemia signs in the impella-accessed limb. the color changed in the right hand/fingers digits two and three. the patient suffered anoxic brain injury resulting in declaration of brain death. life support was withdrawn and the patient expired. per abiomed's medical safety officer review, there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. h.6 codes 3233 and 11 were entered incorrectly on the previously submitted report. the investigation was complete at the time the original report was submitted.h.10 additional manufacturer narrative on the previously submitted report stated that the investigation was not completed yet. the investigation was complete at the time the original report was submitted.
26112 | 1220648-2024-11728 | 19408395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19408395&pc=OZD | 05/15/2024 | 05/16/2025 |  | 05/28/2024 |  |  | 05/15/2024 | 05/28/2024 |  |  | 08/31/2025 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390889 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Shock | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 05/15/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. complications were encountered during implant of the impella 5.5 pump. after passing over the wire, the pump met resistance at the innominate and was unable to advance. the axillary implant was aborted and the pump was subsequently implanted via direct access. there was no patient harm associated with the delivery issue. | Additional information was received from the clinical site that following the implant, the patient experienced bleeding from their chest tube and suffered vasodilatory shock, both of which were considered to have a possible relationship to the impella procedure. the patient was put on inotropic support in response to the vasodilatory shock. the dressing was reinforced at the ct sites and a blood transfusion was performed to treat the bleed. it was later reported that the patient also showed signs of hemolysis during support. following a scheduled procedure on impella support, the patients pfhb values were greater than 40. four units of packed red blood cells, one unit of fresh frozen plasma, and one unit of platelets were transfused to treat the condition. the hemolysis was considered to have a possible relationship to the impella device and procedure. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the pump met resistance at the innominate and was unable to advance as per the clinical description. added-h6 med dev prob code 2682 f6/f8 should have been left blank. g1 reporting contact fax number missing. | B.1 revised to reflect both adverse event and product problem based on the additional information received. b.2 revised to reflect requited intervention due to additional information received from the clinical site. b.5 added information received from the clinical site. d.9 device date returned/information was added. g.2 added study due to additional information received from the clinical site. h.1 revised to serious injury based on the additional information received. h.3 revised if the device was not evaluated, to device evaluation anticipated, but not yet begun due to device being returned. h.6 added codes due to additional information received from the clinical site. h.10 added additional instructions for use due to additional information received from the clinical site. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-11728 was submitted in accordance with updated procedures. | The investigation of delivery issues, hemolysis, vascular damage - great vessel, and vf/vt/af/at has been completed. the root cause of the delivery issue was most likely patient condition related as evidenced by the clinical detail of the pump meeting resistance at the innominate and was unable to advance. the root cause of the access site bleeding - major was not determined as no product was returned for analysis and limited clinical information related to failure was provided. the root cause of the hemolysis issue was not determined as there were no abnormalities in the data logs and insufficient clinical information was provided related to failure. as the necessary information was not provided, the root cause of the vt/vf was not determined. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-11728. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2024-11728. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-11728.
26113 | 1220648-2024-11729 | 19408546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19408546&pc=OZD | 05/17/2024 | 07/15/2024 |  | 05/28/2024 |  |  | 05/17/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449998 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The complainant reported an 82-year-old male patient with pre-operative placement was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported after the coronary artery bypass graft, the 5.5 was directed towards the mitral valve. the doctor tried to reposition the 5.5 in the intensive care unit but stated it felt like there was torque near the blue suture hub just past the locking mechanism, so they were unable to torque the pump back into position. the patient developed hemolysis but did not develop mitral regurgitation or ectopy. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported positioning issue and the hemolysis has been completed. data logs were reviewed which showed the pump ran without issue. there were no motor current spikes and suction alarms were triggered but resolved quickly. logs also showed some alarms impella position in aorta on near the end of the case, but unrelated to hemolysis event. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. b1-selected product problem. f6/f8 should have been left blank in the initial report.
26114 | 1220648-2024-11710 | 19408869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19408869&pc=OZD | 05/18/2024 | 06/24/2024 |  | 05/28/2024 |  |  | 05/18/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493202 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Upon follow up from the site, it was reported that two days following explant of the impella cp pump, pulses remained absent and a fasciotomy was performed. it was subsequently determined that the lower leg was not viable and a below knee amputation was completed two days later. it was suggested that the condition was potentially related to heparin induced thrombocytopenia, but this was not verified. the patient survived the operation. | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support, the patient lost flow in the leg, and in response to the condition, the patient was taken to the operating room for a washout and escalation to impella 5.5 pump support. the outcome of the patient was noted as unknown. |  |  |  |  | Additional narrative/information was added to b.2 and b.5. additional coding added to h.6 to address subsequent medical treatment given to the patient to manage the condition. d.4 revised model number from the initial submission in accordance with updated procedures. e.1 facility name revised as previously entered incorrectly. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system-section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ irreversible issue has been completed since the original report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
26115 | 1220648-2024-11712 | 19408955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19408955&pc=OZD | 04/04/2024 | 03/17/2025 |  | 05/28/2024 |  |  | 05/06/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453325 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/06/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating an allegation of ventricular tachycardia ¿cp, insertion of direct impella 5.5 on (b)(6) 2024. ventricular tachycardia in the or prior to start of procedure. very profoundly coagulopathic in the or with significant product transfusion with 6u prbc's, 4u plt's, 8 ffp, and 3 cryo¿. the patient is unknown at this time. | Additional information received via abiomed¿s long-term outcome and quality indicator impella (loqi) registry indicating an allegation of hemolysis. the loqi states: very profoundly coagulopathic in the or with significant product transfusion with 6u prbc's, 4u plt's, 8 ffp, and 3 cryo. also loaded with brilinta. 2u prbc's since arriving to icu. the relatedness to the pump use is "possibly". |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The root cause of the hemolysis was not determined since the product was not returned and insufficient clinical details were provided. b5 revised with the additional information received from the clinical site. h6 health effect - clinical code revised to include the additional information received. f6/f8-should have been left blank on the initial manufacturer device report 1220648-2024-11712 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-11712 in accordance with updated procedures. h6 component code revised from the initial manufacturer device report 1220648-2024-11712 in accordance with updated procedures on. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. d4-updated model number.
26116 | 1220648-2024-11733 | 19409281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19409281&pc=OZD | 05/21/2024 | 07/26/2024 |  | 05/28/2024 |  |  | 05/21/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025459139 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  |  | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2024 |  |  |  |  | The user facility reported a 64-year-old female patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during the procedure flows reduced to 0.0 and they were not able to achieve flow to the patient and the mean arterial pressure (map) dropped in to the 40s. position was confirmed, however there were suction alarms and the physician attempted to reposition the pump. while repositioning, the patient coded, and cardiopulmonary resuscitation was performed. the patient expired after 3.75 hours of support. the cause of death is unknown. | Per medical safety officer review there is not enough information to associate use of device with outcome of death. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for low pump flow has been completed. data logs confirmed the failure mode and clinical narrative. logs showed about 2.5 hours into the case, mc spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11733: b1 product problem only. b2 death was inadvertently selected. b5 updated with corrected information. f6/f8-should have been left blank. h1 type of reportable event. h.6 code 1802 was reported incorrectly.
26117 | 1220648-2024-11734 | 19411608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19411608&pc=OZD | 04/30/2024 | 06/25/2024 |  | 05/28/2024 |  |  | 04/30/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465297 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed along with an impella rp. the impella cp was displaying low flows and suctions events. the device was seen to be kinked and was explanted and replaced with a new impella cp. the patient went into pulseless electrical activity within an hour of placement and being sent to the intensive care unit. the family made the patient a do not resuscitate and the patient expired. per abiomed medical safety office review: there is not enough information to associate use of device with outcome. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. visual inspection found a kink on cannula about 14 mm from outflow cage and cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was kinked cannula.
26118 | 1220648-2024-11735 | 19411894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19411894&pc=OZD | 05/09/2024 | 02/13/2025 |  | 05/28/2024 |  |  | 05/13/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025455140 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient who was implanted with an impella 5.5 device for mechanical circulatory support experienced cerebellar infarction after explant of the device. it is unclear if it is related to impella. regarding the blood vessel properties at the time of impella insertion, there was no calcification. it was determined that there is a high possibility that it occurred during support or removal. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Manufacturer device report is a duplicate of manufacturer device report 1220648-2024-09746. the investigation results for stroke are found in manufacturer device report 1220648-2024-09746. no further information will be submitted for manufacturer device report.
26119 | 1220648-2024-11736 | 19411943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19411943&pc=OZD | 05/16/2024 | 07/10/2024 |  | 05/28/2024 |  |  | 05/16/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478417 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported an 71-year-old patient with cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. it was reported the impella 5.5¿s position was too deep. when the blue locking button was pressed, the catheter moved back approximately five centimeters as if tension had been built up. the doctor tried to advance the pump but met resistance. the doctor then had to reopen the pocket on the right chest to evaluate further. the doctor found and removed a clot on the axillary graft and then repositioned the impella 5.5. the anti-contamination sleeve broke when repositioning. the patient survived the event. | [5 empty] | The investigation for the product damage, positioning issues, and thrombus has been completed. the pump was not returned for investigation. data logs show several impella position unknown alarms. as per the clinical, pump position was reported as deep. the catheter was jumped back after pressing the blue locking button. the md felt resistance during repositioning. later, the md found a clot encompassing the axillary graft, which caused the pump to be advanced. the pump was placed successfully after the removal of the clot. during repositioning, the sterile sleeve was found broken. the cause of the first positioning issue was most likely cath anchor use issue, based on clinical details. the cause of the second positioning issue was a patient condition related to thrombus buildup in the graft, which prevented the pump from being properly repositioned. the cause of the sterile sleeve damage issue was most likely use related, based on clinical details. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
26120 | 1220648-2024-11737 | 19411956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19411956&pc=OZD | 05/16/2024 | 07/11/2024 |  | 05/28/2024 |  |  | 05/16/2024 | 05/28/2024 |  |  |  |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471276 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 60-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. positioning issues and suction alarms were reported. the position was deep across the aortic valve, and the doctor attempted to pull the pump back and it was reported to have jumped. the aorta alarm sounded. the doctor was not able to re-advance the pump, and the pump was replaced. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed | The investigation for the low pump flow and positioning issues has been completed. the impella device was not returned for investigation. data logs were reviewed finding suction alarms occurred. many positioning issues occurred during the impella support. the cause of the positioning issue most likely was use related due to improper technique as while repositioning they were accidently moved it all the way back across the aortic valve. the cause of the initial low flow was most likely improper positioning as it was identified that the pump was not in an optimal position during the suction alarms. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: use of impella with transcatheter aortic valves states: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ f6/f8 should have been left blank in the initial report. g4-added pma number. h8-usage of device added initial use of device.
26121 | 1220648-2024-11738 | 19412217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19412217&pc=OZD | 01/06/2022 | 03/25/2025 |  | 05/28/2024 |  |  | 01/06/2022 | 05/28/2024 |  |  | 09/30/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149806 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2022 |  |  |  |  | The user facility reported a 52-year-old male on impella 5.5 for mechanical circulatory support. while on support, the patient had stroke code and was sent to or (operating room) where thrombus was removed. the patient's outcome is unknown as the facility did not provide any further information. | [5 empty] | The impella 5.5 device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | As this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
26122 | 1220648-2024-11739 | 19412421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19412421&pc=OZD | 01/16/2022 | 05/27/2025 |  | 05/28/2024 |  |  | 01/16/2022 | 05/28/2024 |  |  | 10/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153876 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/16/2022 |  |  |  |  | The complainant reported a patient was escalated to the impella 5.5 due to ami/csg (acute myocardial infarction/cardiogenic shock). a trans-esophageal echo (tte) was performed to check the position due to multiple ventricular tachycardia (vt) episodes treated with an external defibrillator. the impella 5.5 device was pulled back to 46 cm. the patient coughed and went into vt at 288 beats per minute for approximately 5 minutes, and was treated with shock. the patient was then transitioned to an external blood pump, centrimag. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26123 | 1220648-2024-11740 | 19412491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19412491&pc=OZD | 04/08/2022 | 05/28/2024 |  | 05/28/2024 |  |  | 04/08/2022 | 05/28/2024 |  |  | 12/31/2023 |  | 05/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179938 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The user facility reported a 63-year-old male patient implanted with an impella cp for mechanical circulatory support. it was reported that ventricular access was achieved and the impella cp was backloaded and implanted without issue as placement was verified under fluoroscopy. bleeding was noted at the impella cp groin site. access site bleeding subsided, and an additional suture was placed to resolve this issue. the patient was successfully weaned and the impella cp explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26124 | 1220648-2024-11673 | 19398530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398530&pc=OZD | 10/14/2021 | 03/25/2025 |  | 05/25/2024 |  |  | 10/14/2021 | 05/25/2024 |  |  | 05/31/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087801 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the complainant reported the patient had access site bleeding that required being closed via double perclose and standard care for 5 minutes of manual pressure for the oozing. the impella was weaned and explanted without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26125 | 1220648-2024-11674 | 19398567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398567&pc=OZD | 10/14/2021 | 03/25/2025 |  | 05/25/2024 |  |  | 10/14/2021 | 05/25/2024 |  |  | 04/30/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082057 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump flow issue. the device appeared to be kinked near the motor. with repositioning the kink appeared to be resolved by angiographic appearance. the cp flow never was able to get much above 1 l/min without suction alarms. patient baseline left ventricular end-diastolic pressure (lvedp) high 20's. the physician did not feel it was enough volume; even thou the position was good. the physician removed the impella cp, flushed the device, and replanted as support was still necessary. the impella cp still had poor flows after being re-implanted. there were suction alarms and the pump was unable to get above 1 l/min despite being in the correct position. the impella was explanted and another device was used. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: low pump flow impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message - could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26126 | 1220648-2024-11675 | 19398591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398591&pc=OZD | 10/15/2021 | 03/25/2025 |  | 05/25/2024 |  |  | 10/15/2021 | 05/25/2024 |  |  | 08/31/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had an episode of ventricular tachycardia. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: arrhythmia impella cp section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26127 | 1220648-2024-11676 | 19398608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398608&pc=OZD | 10/15/2021 | 03/25/2025 |  | 05/25/2024 |  |  | 10/15/2021 | 05/25/2024 |  |  | 05/05/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089673 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed bleeding at the access site. two (2) percloses were deployed in the axillary access site post-removal of the sheath, but bleeding continued internally and externally. the patient received blood products, and the bleeding was successfully managed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26128 | 1220648-2024-11677 | 19398615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398615&pc=OZD | 10/22/2021 | 05/25/2024 |  | 05/25/2024 |  |  | 10/22/2021 | 05/25/2024 |  |  | 08/31/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124593 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/22/2021 |  |  |  |  | The user facility reported that a patient was presented for a high-risk percutaneous coronary intervention (hrpci) with coronary stent implantation (csi). it was noted that the patient had severe aortic stenosis and calcification. an attempt was made to implant the impella; however, the device was unable to advance across the valve. the impella cp procedure was aborted and a percutaneous coronary intervention (pci) with orbital csi was performed without impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26129 | 1220648-2024-11678 | 19398623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398623&pc=OZD | 10/24/2021 | 05/25/2024 |  | 05/25/2024 |  |  | 10/24/2021 | 05/25/2024 |  |  | 07/31/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111034 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/24/2021 |  |  |  |  | The user facility reported an 82-year-old patient was implanted with an impella cp for mechanical circulatory support. it was reported, that upon removal of the impella, there was also an attempt to place a 11 fr sheath over a .035 wire, however, the wire dissected a vessel. the wire was removed, and vessel closed successfully. a clot was found in the superficial femoral artery, used penumbra to remove clot with success. additionally, patient experienced oozing at groin, flow resolved. the patient was later successfully weaned from the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
26130 | 1220648-2024-11679 | 19398634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398634&pc=OZD | 10/24/2021 | 05/20/2025 |  | 05/25/2024 |  |  | 10/24/2021 | 05/25/2024 |  |  | 04/30/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075452 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/24/2021 |  |  |  |  | The user facility reported a patient needing an impella cp for mechanical circulatory support due to a non-st elevation myocardial infarction (nstemi). upon insertion of the foley catheter, the patient coded. the impella device was placed. it was reported that the patient¿s laboratory results were hemolyzed despite the pump appearing to be in good position via echocardiogram (echo). the decision was made to reposition the impella to resolve hemolysis. the device was observed to be colliding with the papillary muscle; however, additional repositioning was not confirmed. it was stated that dialysis was started, no signs of hemolysis were identified. weaning was successful, the impella was explanted, and the procedural outcome was that the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.